## Summary
Raw Data Export | 4.14 | 2022-07-21 15:58 | 1.9.2.6
Sample ID | Sample Type | Sample Specimen Type | Test Specimen Type | Specimen Tube Type | Assay Name | Result Code | RP,Assay Version | Start Date/Time | End Date/Time | Status | Overall Result | Flags | Replicate Number | Extraction Number | Elution Number | Yellow Target Ct | Yellow Target EP | Yellow Target Log Conc. | Yellow Target Flag | Red Target Ct | Red Target EP | Red Target Log Conc. | Red Target Flag | Green Target Ct | Green Target EP | Green Target Log Conc. | Green Target Flag | Orange Target Ct | Orange Target EP | Orange Target Log Conc. | Orange Target Flag | Far Red Target Ct | Far Red Target EP | Far Red Target Log Conc. | Far Red Target Flag | Test Operator | Comment | Run Type | Sample Origin | Sample Guid | Test Guid | Test Comment | Control Name | Replicates | Extractions | Elutions | Delta Ct | EPR Ratio | EPR Difference | Patient ID | Dilution Factor | Instrument Serial Number | Control Prefix | Control Lot | Control Serial | Assay Id
BKV 10m4 U R1 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:32 | 2022-07-21 12:59 | Completed | Indeterminate |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.3 | 29213 |  | 1012 (Error, Outlier Removal Failed) | [8 empty] | ebixler |  | Normal | ManuallyScanned | 0b5c9e96-1209-ed11-85f4-5cf370a20867 | 0d5c9e96-1209-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R2 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:32 | 2022-07-21 12:59 | Completed | Indeterminate |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.24 | 29527 |  | 1012 (Error, Outlier Removal Failed) | [8 empty] | ebixler |  | Normal | ManuallyScanned | 125c9e96-1209-ed11-85f4-5cf370a20867 | 145c9e96-1209-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R3 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:32 | 2022-07-21 12:59 | Completed | Indeterminate |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.2 | 28852 |  | 1012 (Error, Outlier Removal Failed) | [8 empty] | ebixler |  | Normal | ManuallyScanned | 195c9e96-1209-ed11-85f4-5cf370a20867 | 1b5c9e96-1209-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R4 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:32 | 2022-07-21 12:59 | Completed | Indeterminate |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.22 | 28421 |  | 1012 (Error, Outlier Removal Failed) | [8 empty] | ebixler |  | Normal | ManuallyScanned | 205c9e96-1209-ed11-85f4-5cf370a20867 | 225c9e96-1209-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R5 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:32 | 2022-07-21 12:59 | Completed | Indeterminate |  | 1 | 1 | 1 | 16.52 | 468 |  | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) |  |  |  |  | 16.25 | 29297 |  | 1012 (Error, Outlier Removal Failed) | [8 empty] | ebixler |  | Normal | ManuallyScanned | 275c9e96-1209-ed11-85f4-5cf370a20867 | 295c9e96-1209-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R6 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:32 | 2022-07-21 12:45 | Completed | Positive |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.65 | 23255 | 6.73 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 2e5c9e96-1209-ed11-85f4-5cf370a20867 | 305c9e96-1209-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R7 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:32 | 2022-07-21 12:45 | Completed | Positive |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.42 | 23336 | 6.8 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 355c9e96-1209-ed11-85f4-5cf370a20867 | 375c9e96-1209-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R8 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:32 | 2022-07-21 12:46 | Completed | Positive |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.3 | 22748 | 6.84 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 3c5c9e96-1209-ed11-85f4-5cf370a20867 | 3e5c9e96-1209-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R1 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:34 | 2022-07-21 13:19 | Completed | Positive |  | 1 | 1 | 1 | 16.52 | 471 |  | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) |  |  |  |  | 16.34 | 27960 | 6.83 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 0b5c9e96-1209-ed11-85f4-5cf370a20867 | 6042e1f8-1209-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R2 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:34 | 2022-07-21 13:09 | Completed | Positive |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.35 | 28768 | 6.82 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 125c9e96-1209-ed11-85f4-5cf370a20867 | 6642e1f8-1209-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R3 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:34 | 2022-07-21 13:09 | Completed | Positive |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.32 | 29845 | 6.83 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 195c9e96-1209-ed11-85f4-5cf370a20867 | 6c42e1f8-1209-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R4 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:34 | 2022-07-21 13:09 | Completed | Indeterminate |  | 1 | 1 | 1 | 16.37 | 745 |  | 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold), 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed) |  |  |  |  | 16.37 | 30635 |  | 1012 (Error, Outlier Removal Failed) | [8 empty] | ebixler |  | Normal | ManuallyScanned | 205c9e96-1209-ed11-85f4-5cf370a20867 | 7242e1f8-1209-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R5 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:34 | 2022-07-21 13:09 | Completed | Positive |  | 1 | 1 | 1 | 16.27 | 874 |  | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) |  |  |  |  | 16.46 | 30118 | 6.79 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 275c9e96-1209-ed11-85f4-5cf370a20867 | 7842e1f8-1209-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R6 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:34 | 2022-07-21 13:01 | Completed | Positive |  | 1 | 1 | 1 |  |  |  | 2610 (Informational, Delays in sample processing), 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.12 | 16922 | 6.89 | 2610 (Informational, Delays in sample processing) | [8 empty] | ebixler |  | Normal | ManuallyScanned | 2e5c9e96-1209-ed11-85f4-5cf370a20867 | 7e42e1f8-1209-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R7 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:34 | 2022-07-21 12:59 | Completed | Indeterminate |  | 1 | 1 | 1 | 16.18 | 699 |  | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) |  |  |  |  | 16.29 | 30075 |  | 1012 (Error, Outlier Removal Failed) | [8 empty] | ebixler |  | Normal | ManuallyScanned | 355c9e96-1209-ed11-85f4-5cf370a20867 | 8442e1f8-1209-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R8 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:34 | 2022-07-21 12:59 | Completed | Positive |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.13 | 26676 | 6.89 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 3c5c9e96-1209-ed11-85f4-5cf370a20867 | 8a42e1f8-1209-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R1 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:37 | 2022-07-21 13:19 | Completed | Positive |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.28 | 27149 | 6.85 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 0b5c9e96-1209-ed11-85f4-5cf370a20867 | 5a7b815a-1309-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R2 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:37 | 2022-07-21 13:19 | Completed | Positive |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.37 | 25184 | 6.82 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 125c9e96-1209-ed11-85f4-5cf370a20867 | 607b815a-1309-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R3 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:37 | 2022-07-21 13:19 | Completed | Positive |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.45 | 25142 | 6.79 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 195c9e96-1209-ed11-85f4-5cf370a20867 | 667b815a-1309-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R4 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:37 | 2022-07-21 13:19 | Completed | Positive |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.38 | 25415 | 6.82 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 205c9e96-1209-ed11-85f4-5cf370a20867 | 6c7b815a-1309-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R5 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:37 | 2022-07-21 13:19 | Completed | Positive |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.44 | 25080 | 6.8 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 275c9e96-1209-ed11-85f4-5cf370a20867 | 727b815a-1309-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R6 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:37 | 2022-07-21 13:27 | Completed | Positive |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.43 | 25496 | 6.8 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 2e5c9e96-1209-ed11-85f4-5cf370a20867 | 787b815a-1309-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R7 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:37 | 2022-07-21 13:19 | Completed | Positive |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.4 | 25271 | 6.81 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 355c9e96-1209-ed11-85f4-5cf370a20867 | 7e7b815a-1309-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R8 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:37 | 2022-07-21 13:19 | Completed | Positive |  | 1 | 1 | 1 | 16.23 | 315 |  | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) |  |  |  |  | 16.18 | 24908 | 6.88 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 3c5c9e96-1209-ed11-85f4-5cf370a20867 | 847b815a-1309-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R1 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:41 | 2022-07-21 13:27 | Completed | Positive |  | 1 | 1 | 1 | 16.74 | 541 |  | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) |  |  |  |  | 16.51 | 22703 | 6.78 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 0b5c9e96-1209-ed11-85f4-5cf370a20867 | b3de19db-1309-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R2 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:41 | 2022-07-21 13:27 | Completed | Positive |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.61 | 29669 | 6.75 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 125c9e96-1209-ed11-85f4-5cf370a20867 | b9de19db-1309-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R3 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:41 | 2022-07-21 13:27 | Completed | Indeterminate |  | 1 | 1 | 1 | 16.26 | 510 |  | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) |  |  |  |  | 16.29 | 29533 |  | 1012 (Error, Outlier Removal Failed) | [8 empty] | ebixler |  | Normal | ManuallyScanned | 195c9e96-1209-ed11-85f4-5cf370a20867 | bfde19db-1309-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R4 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:41 | 2022-07-21 13:27 | Completed | Indeterminate |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.47 | 30993 |  | 1012 (Error, Outlier Removal Failed) | [8 empty] | ebixler |  | Normal | ManuallyScanned | 205c9e96-1209-ed11-85f4-5cf370a20867 | c5de19db-1309-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R5 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:41 | 2022-07-21 13:27 | Completed | Positive |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.49 | 25803 | 6.78 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 275c9e96-1209-ed11-85f4-5cf370a20867 | cbde19db-1309-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R6 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:41 | 2022-07-21 13:27 | Completed | Positive |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.29 | 24689 | 6.84 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 2e5c9e96-1209-ed11-85f4-5cf370a20867 | d1de19db-1309-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R7 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:41 | 2022-07-21 13:34 | Completed | Indeterminate |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.27 | 29909 |  | 1012 (Error, Outlier Removal Failed) | [8 empty] | ebixler |  | Normal | ManuallyScanned | 355c9e96-1209-ed11-85f4-5cf370a20867 | d7de19db-1309-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R8 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:41 | 2022-07-21 13:27 | Completed | Indeterminate |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.06 | 29515 |  | 1012 (Error, Outlier Removal Failed) | [8 empty] | ebixler |  | Normal | ManuallyScanned | 3c5c9e96-1209-ed11-85f4-5cf370a20867 | ddde19db-1309-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R9 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:45 | 2022-07-21 13:34 | Completed | Positive |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.27 | 26286 | 6.85 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 63fdae7a-1409-ed11-85f4-5cf370a20867 | 65fdae7a-1409-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R9 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:45 | 2022-07-21 13:34 | Completed | Positive |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.02 | 30622 | 6.92 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 63fdae7a-1409-ed11-85f4-5cf370a20867 | 6bfdae7a-1409-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R9 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:45 | 2022-07-21 13:34 | Completed | Positive |  | 1 | 1 | 1 | 15.89 | 667 |  | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) |  |  |  |  | 16.14 | 28700 | 6.89 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 63fdae7a-1409-ed11-85f4-5cf370a20867 | 71fdae7a-1409-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R9 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 11:45 | 2022-07-21 13:34 | Completed | Indeterminate |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.03 | 29973 |  | 1012 (Error, Outlier Removal Failed) | [8 empty] | ebixler |  | Normal | ManuallyScanned | 63fdae7a-1409-ed11-85f4-5cf370a20867 | 77fdae7a-1409-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R10 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 12:01 | 2022-07-21 13:56 | Completed | Indeterminate |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.14 | 29636 |  | 1012 (Error, Outlier Removal Failed) | [8 empty] | ebixler |  | Normal | ManuallyScanned | f27fc1ae-1609-ed11-85f4-5cf370a20867 | f47fc1ae-1609-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R10 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 12:01 | 2022-07-21 13:56 | Completed | Indeterminate |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.12 | 32203 |  | 1012 (Error, Outlier Removal Failed) | [8 empty] | ebixler |  | Normal | ManuallyScanned | f27fc1ae-1609-ed11-85f4-5cf370a20867 | fa7fc1ae-1609-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R10 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 12:01 | 2022-07-21 13:56 | Completed | Indeterminate |  | 1 | 1 | 1 | 16 | 402 |  | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) |  |  |  |  | 16.13 | 32904 |  | 1012 (Error, Outlier Removal Failed) | [8 empty] | ebixler |  | Normal | ManuallyScanned | f27fc1ae-1609-ed11-85f4-5cf370a20867 | 0080c1ae-1609-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R1 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 13:03 | 2022-07-21 13:06 | AbortedBySystem | NoResult |  | 1 | 1 | 1 |  |  |  | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) | [7 empty] | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) | [8 empty] | ebixler |  | Repeat | ManuallyScanned | 0b5c9e96-1209-ed11-85f4-5cf370a20867 | a79dad68-1f09-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R2 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 13:03 | 2022-07-21 13:06 | AbortedBySystem | NoResult |  | 1 | 1 | 1 |  |  |  | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) | [7 empty] | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) | [8 empty] | ebixler |  | Repeat | ManuallyScanned | 125c9e96-1209-ed11-85f4-5cf370a20867 | ad9dad68-1f09-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R3 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 13:03 | 2022-07-21 13:06 | AbortedBySystem | NoResult |  | 1 | 1 | 1 |  |  |  | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) | [7 empty] | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) | [8 empty] | ebixler |  | Repeat | ManuallyScanned | 195c9e96-1209-ed11-85f4-5cf370a20867 | b39dad68-1f09-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R4 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 13:03 | 2022-07-21 13:06 | AbortedBySystem | NoResult |  | 1 | 1 | 1 |  |  |  | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) | [7 empty] | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) | [8 empty] | ebixler |  | Repeat | ManuallyScanned | 205c9e96-1209-ed11-85f4-5cf370a20867 | b99dad68-1f09-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R5 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 13:03 | 2022-07-21 13:06 | AbortedBySystem | NoResult |  | 1 | 1 | 1 |  |  |  | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) | [7 empty] | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) | [8 empty] | ebixler |  | Repeat | ManuallyScanned | 275c9e96-1209-ed11-85f4-5cf370a20867 | bf9dad68-1f09-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R7 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 13:03 | 2022-07-21 13:06 | AbortedBySystem | NoResult |  | 1 | 1 | 1 |  |  |  | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) | [7 empty] | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) | [8 empty] | ebixler |  | Repeat | ManuallyScanned | 355c9e96-1209-ed11-85f4-5cf370a20867 | c59dad68-1f09-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R10 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 13:03 | 2022-07-21 14:06 | Completed | Positive |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 17.03 | 25393 | 6.62 | [9 empty] | ebixler |  | Normal | ManuallyScanned | f27fc1ae-1609-ed11-85f4-5cf370a20867 | cb9dad68-1f09-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R11 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 13:03 | 2022-07-21 14:06 | Completed | Positive |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.59 | 25183 | 6.75 | [9 empty] | ebixler |  | Normal | ManuallyScanned | d09dad68-1f09-ed11-85f4-5cf370a20867 | d29dad68-1f09-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R11 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 13:09 | 2022-07-21 14:23 | Completed | Positive |  | 1 | 1 | 1 | 16.45 | 703 |  | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) |  |  |  |  | 16.64 | 25353 | 6.74 | [9 empty] | ebixler |  | Normal | ManuallyScanned | d09dad68-1f09-ed11-85f4-5cf370a20867 | 1f42d920-2009-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R12 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 13:09 | 2022-07-21 14:23 | Completed | Positive |  | 1 | 1 | 1 | 16.28 | 438 |  | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) |  |  |  |  | 16.08 | 24419 | 6.91 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 2442d920-2009-ed11-85f4-5cf370a20867 | 2642d920-2009-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R11 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 13:09 | 2022-07-21 14:23 | Completed | Positive |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.63 | 24694 | 6.74 | [9 empty] | ebixler |  | Normal | ManuallyScanned | d09dad68-1f09-ed11-85f4-5cf370a20867 | 2c42d920-2009-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R12 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 13:09 | 2022-07-21 14:23 | Completed | Positive |  | 1 | 1 | 1 |  |  |  | 1020 (Informational, Peak Not Detected) |  |  |  |  | 16.32 | 24142 | 6.83 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 2442d920-2009-ed11-85f4-5cf370a20867 | 3242d920-2009-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R11 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 13:09 | 2022-07-21 14:23 | Completed | Positive |  | 1 | 1 | 1 | 16.78 | 502 |  | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) |  |  |  |  | 16.73 | 24578 | 6.71 | [9 empty] | ebixler |  | Normal | ManuallyScanned | d09dad68-1f09-ed11-85f4-5cf370a20867 | 3842d920-2009-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R12 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 13:09 | 2022-07-21 14:23 | Completed | Positive |  | 1 | 1 | 1 | 16.61 | 706 |  | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) |  |  |  |  | 16.5 | 24232 | 6.78 | [9 empty] | ebixler |  | Normal | ManuallyScanned | 2442d920-2009-ed11-85f4-5cf370a20867 | 3e42d920-2009-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R12 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 13:09 | 2022-07-21 13:15 | AbortedBySystem | NoResult |  | 1 | 1 | 1 |  |  |  | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) | [7 empty] | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) | [8 empty] | ebixler |  | Normal | ManuallyScanned | 2442d920-2009-ed11-85f4-5cf370a20867 | 4442d920-2009-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R3 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 15:50 | 2022-07-21 15:51 | AbortedBySystem | NoResult |  | 1 | 1 | 1 |  |  |  | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) | [7 empty] | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) | [8 empty] | ebixler |  | Repeat | ManuallyScanned | 195c9e96-1209-ed11-85f4-5cf370a20867 | b7477ca2-3609-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R4 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 15:51 | 2022-07-21 15:53 | AbortedBySystem | NoResult |  | 1 | 1 | 1 |  |  |  | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) | [7 empty] | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) | [8 empty] | ebixler |  | Repeat | ManuallyScanned | 205c9e96-1209-ed11-85f4-5cf370a20867 | f5c2a3da-3609-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R7 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 15:53 | 2022-07-21 15:54 | AbortedBySystem | NoResult |  | 1 | 1 | 1 |  |  |  | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) | [7 empty] | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) | [8 empty] | ebixler |  | Repeat | ManuallyScanned | 355c9e96-1209-ed11-85f4-5cf370a20867 | 90b7eb13-3709-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R8 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 15:54 | 2022-07-21 15:56 | AbortedBySystem | NoResult |  | 1 | 1 | 1 |  |  |  | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) | [7 empty] | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) | [8 empty] | ebixler |  | Repeat | ManuallyScanned | 3c5c9e96-1209-ed11-85f4-5cf370a20867 | f2bbce4d-3709-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867
BKV 10m4 U R9 | Patient Sample | Urine | Urine | SecondaryTube13x75 | BKV URINE 10 | L999 | 2.1.9.5,0.0.40 | 2022-07-21 15:56 | 2022-07-21 15:58 | AbortedByUser | Aborted | 4 (Error, User Aborted) | 1 | 1 | 1 |  |  |  | 7162 (Warning, Quantity not sufficient), 1000 (Error, User Aborted Test) | [7 empty] | 7162 (Warning, Quantity not sufficient), 1000 (Error, User Aborted Test) | [8 empty] | ebixler |  | Repeat | ManuallyScanned | 63fdae7a-1409-ed11-85f4-5cf370a20867 | 0e592f87-3709-ed11-85f4-5cf370a20867 |  |  | 1 | 1 | 1 | 0 | 0 | 0 |  | None | 12000112 |  |  |  | f394bea5-0509-ed11-85f4-5cf370a20867

## Chain of Custody
Sample ID | Patient ID | Test Guid | Test Comment | Replicate Number | Extraction Number | Elution Number | Primer/Probe | Sample Tube Rack | Sample Tube Position | Capture Plate Barcode | Capture Plate Well | Heater Module Index | Heater Module Serial | Buffer Barcode | Test Strip NeuMoDx Barcode | Test Strip NeuMoDx Carrier | Test Strip NeuMoDx Carrier Position | Test Strip NeuMoDx Well | Test Strip LDT Primer Probe Barcode | Test Strip LDT Primer Probe Well | Test Strip LDT Primer Carrier | Test Strip LDT Primer Carrier Position | Test Strip LDT Master Mix Barcode | Test Strip LDT Master Mix Carrier | Test Strip LDT Master Mix Carrier Position | Test Strip LDT Master Mix Well | Extraction Cartridge Barcode | Extraction Cartridge Lane | Extraction Module Index | Extraction Module Serial | Pcr Cartridge Barcode | Pcr Cartridge Lane | PCR Module Index | PCR Module Serial | Sample Large Tip Rack Barcode | Sample Large Tip Position | Filter Large Tip Rack Barcode | Filter Large Tip Position | Buffer Large Tip Rack Barcode | Buffer Large Tip Position | LHPB Large Tip Rack Barcode | LHPB Large Tip Position | Small Tip Rack Barcode | Small Tip Position | Filter Barcode | Assay Name | Assay Version | Software Version | Hamilton Software Version | Hamilton Serial Number | Hamilton Firmware Version | N500 Serial Number | Hamilton Method Version | Heater Firmware Version | XPCR Firmware Version | XPCR Actuator Version | Release Reagent Barcode | Wash Reagent Barcode | Bulk Reagent Drawer | Start Heater Ambient Temperature C | End Heater Ambient Temperature C | Start Pcr Ambient Temperature C | End Pcr Ambient Temperature C | Standard Curve GUID | Standard Curve Lot | LHPA ADP Position | LHPA Start Date Time | LHPB ADP Position | LHPB Start Date Time | LHPC ADP Position | LHPC Start Date Time | PCR Start Date Time | Standard Curve Master Mix Lot | Standard Curve Primer Probe Lot | LDT Test Strip Primer Probe Lot | LDT Test Strip Master Mix Lot | Green Crosstalk Into Yellow | Yellow Crosstalk Into Orange | Orange Crosstalk Into Red | Red Crosstalk Into Far Red
BKV 10m4 U R1 |  | 0d5c9e96-1209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 1 | 010081427802028410111464 2100822 17230505 | 5 | 3 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100849 17221209 | 8 | TestStrips3 | 5 | 010081427802019210108908 2100289 17220725 | TestStrips1 | 5 | 8 | 010081427802027710118705 2105880 17240328 | 8 | 3 | V790 | 010081427802027710118705 2105880 17240328 | 8 | 3 | V790 | 235905074779227267 | 49 |  |  | 235905074779255266 | 49 | 235905074779255266 | 49 | 235903060064761242 | 60 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.2 | 34.2 | 35.5 | 36.5 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 1 | 7/21/2022 11:31:39 AM -05:00 | 1 | 7/21/2022 11:50:15 AM -05:00 | 4 | 7/21/2022 12:17:49 PM -05:00 | 7/21/2022 12:28:15 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R2 |  | 145c9e96-1209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 2 | 010081427802028410111464 2100822 17230505 | 6 | 3 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100849 17221209 | 7 | TestStrips3 | 5 | 010081427802019210108908 2100289 17220725 | TestStrips1 | 5 | 7 | 010081427802027710118705 2105880 17240328 | 9 | 3 | V790 | 010081427802027710118705 2105880 17240328 | 9 | 3 | V790 | 235905074779227267 | 50 |  |  | 235905074779255266 | 50 | 235905074779255266 | 50 | 235903060064761242 | 61 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.2 | 34.2 | 35.5 | 36.5 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 2 | 7/21/2022 11:31:39 AM -05:00 | 2 | 7/21/2022 11:50:15 AM -05:00 | 5 | 7/21/2022 12:17:49 PM -05:00 | 7/21/2022 12:28:15 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R3 |  | 1b5c9e96-1209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 3 | 010081427802028410111464 2100822 17230505 | 7 | 3 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100849 17221209 | 6 | TestStrips3 | 5 | 010081427802019210108908 2100289 17220725 | TestStrips1 | 5 | 6 | 010081427802027710118705 2105880 17240328 | 10 | 3 | V790 | 010081427802027710118705 2105880 17240328 | 10 | 3 | V790 | 235905074779227267 | 51 |  |  | 235905074779255266 | 51 | 235905074779255266 | 51 | 235903060064761242 | 62 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.2 | 34.2 | 35.5 | 36.5 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 3 | 7/21/2022 11:31:39 AM -05:00 | 3 | 7/21/2022 11:50:15 AM -05:00 | 6 | 7/21/2022 12:17:49 PM -05:00 | 7/21/2022 12:28:15 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R4 |  | 225c9e96-1209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 4 | 010081427802028410111464 2100822 17230505 | 8 | 3 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100849 17221209 | 5 | TestStrips3 | 5 | 010081427802019210108908 2100289 17220725 | TestStrips1 | 5 | 5 | 010081427802027710118705 2105880 17240328 | 11 | 3 | V790 | 010081427802027710118705 2105880 17240328 | 11 | 3 | V790 | 235905074779227267 | 52 |  |  | 235905074779255266 | 52 | 235905074779255266 | 52 | 235903060064761242 | 63 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.2 | 34.2 | 35.5 | 36.5 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 4 | 7/21/2022 11:31:39 AM -05:00 | 4 | 7/21/2022 11:50:15 AM -05:00 | 7 | 7/21/2022 12:17:49 PM -05:00 | 7/21/2022 12:28:15 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R5 |  | 295c9e96-1209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 5 | 010081427802028410111464 2100888 17230505 | 5 | 4 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100849 17221209 | 4 | TestStrips3 | 5 | 010081427802019210108908 2100289 17220725 | TestStrips1 | 5 | 4 | 010081427802027710118705 2105880 17240328 | 12 | 3 | V790 | 010081427802027710118705 2105880 17240328 | 12 | 3 | V790 | 235905074779227267 | 53 |  |  | 235905074779255266 | 53 | 235905074779255266 | 53 | 235903060064761242 | 64 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.2 | 34.2 | 35.5 | 36.5 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 5 | 7/21/2022 11:31:39 AM -05:00 | 5 | 7/21/2022 11:50:15 AM -05:00 | 8 | 7/21/2022 12:17:49 PM -05:00 | 7/21/2022 12:28:15 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R6 |  | 305c9e96-1209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 6 | 010081427802028410111464 2100888 17230505 | 6 | 4 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100849 17221209 | 3 | TestStrips3 | 5 | 010081427802019210108908 2100289 17220725 | TestStrips1 | 5 | 3 | 010081427802027710118705 2105879 17240328 | 1 | 4 | V1169 | 010081427802027710118705 2105879 17240328 | 1 | 4 | V1169 | 235905074779227267 | 54 |  |  | 235905074779255266 | 54 | 235905074779255266 | 54 | 235903060064761242 | 53 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.2 | 34.2 | 34.2 | 34.5 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 6 | 7/21/2022 11:31:39 AM -05:00 | 6 | 7/21/2022 11:50:15 AM -05:00 | 1 | 7/21/2022 12:10:01 PM -05:00 | 7/21/2022 12:14:35 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R7 |  | 375c9e96-1209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 7 | 010081427802028410111464 2100888 17230505 | 7 | 4 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100849 17221209 | 2 | TestStrips3 | 5 | 010081427802019210108908 2100289 17220725 | TestStrips1 | 5 | 2 | 010081427802027710118705 2105879 17240328 | 2 | 4 | V1169 | 010081427802027710118705 2105879 17240328 | 2 | 4 | V1169 | 235905074779227267 | 55 |  |  | 235905074779255266 | 55 | 235905074779255266 | 55 | 235903060064761242 | 54 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.2 | 34.2 | 34.2 | 34.5 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 7 | 7/21/2022 11:31:39 AM -05:00 | 7 | 7/21/2022 11:50:15 AM -05:00 | 2 | 7/21/2022 12:10:01 PM -05:00 | 7/21/2022 12:14:35 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R8 |  | 3e5c9e96-1209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 8 | 010081427802028410111464 2100888 17230505 | 8 | 4 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100849 17221209 | 1 | TestStrips3 | 5 | 010081427802019210108908 2100289 17220725 | TestStrips1 | 5 | 1 | 010081427802027710118705 2105879 17240328 | 3 | 4 | V1169 | 010081427802027710118705 2105879 17240328 | 3 | 4 | V1169 | 235905074779227267 | 56 |  |  | 235905074779255266 | 56 | 235905074779255266 | 56 | 235903060064761242 | 55 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.2 | 34.2 | 34.2 | 34.5 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 8 | 7/21/2022 11:31:39 AM -05:00 | 8 | 7/21/2022 11:50:15 AM -05:00 | 3 | 7/21/2022 12:10:01 PM -05:00 | 7/21/2022 12:14:35 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R1 |  | 6042e1f8-1209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 1 | 010081427802028410111464 2100822 17230505 | 1 | 3 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100849 17221209 | 16 | TestStrips3 | 5 | 010081427802019210108908 2100289 17220725 | TestStrips1 | 5 | 16 | 010081427802027710118705 2105881 17240328 | 12 | 2 | V834 | 010081427802027710118705 2105881 17240328 | 12 | 2 | V834 | 235905074779227267 | 57 |  |  | 235905074779255266 | 57 | 235905074779255266 | 57 | 235903060064761242 | 80 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.6 | 34.2 | 36.6 | 37.8 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 1 | 7/21/2022 11:34:22 AM -05:00 | 8 | 7/21/2022 11:55:18 AM -05:00 | 8 | 7/21/2022 12:38:37 PM -05:00 | 7/21/2022 12:48:05 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R2 |  | 6642e1f8-1209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 2 | 010081427802028410111464 2100822 17230505 | 2 | 3 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100849 17221209 | 15 | TestStrips3 | 5 | 010081427802019210108908 2100289 17220725 | TestStrips1 | 5 | 15 | 010081427802027710118705 2105880 17240328 | 1 | 3 | V790 | 010081427802027710118705 2105880 17240328 | 1 | 3 | V790 | 235905074779227267 | 58 |  |  | 235905074779255266 | 58 | 235905074779255266 | 58 | 235903060064761242 | 69 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.6 | 34.2 | 36 | 36.3 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 2 | 7/21/2022 11:34:22 AM -05:00 | 1 | 7/21/2022 11:52:47 AM -05:00 | 1 | 7/21/2022 12:33:05 PM -05:00 | 7/21/2022 12:37:52 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R3 |  | 6c42e1f8-1209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 3 | 010081427802028410111464 2100822 17230505 | 3 | 3 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100849 17221209 | 14 | TestStrips3 | 5 | 010081427802019210108908 2100289 17220725 | TestStrips1 | 5 | 14 | 010081427802027710118705 2105880 17240328 | 2 | 3 | V790 | 010081427802027710118705 2105880 17240328 | 2 | 3 | V790 | 235905074779227267 | 59 |  |  | 235905074779255266 | 59 | 235905074779255266 | 59 | 235903060064761242 | 70 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.6 | 34.2 | 36 | 36.3 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 3 | 7/21/2022 11:34:22 AM -05:00 | 2 | 7/21/2022 11:52:47 AM -05:00 | 2 | 7/21/2022 12:33:05 PM -05:00 | 7/21/2022 12:37:52 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R4 |  | 7242e1f8-1209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 4 | 010081427802028410111464 2100822 17230505 | 4 | 3 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100849 17221209 | 13 | TestStrips3 | 5 | 010081427802019210108908 2100289 17220725 | TestStrips1 | 5 | 13 | 010081427802027710118705 2105880 17240328 | 3 | 3 | V790 | 010081427802027710118705 2105880 17240328 | 3 | 3 | V790 | 235905074779227267 | 60 |  |  | 235905074779255266 | 60 | 235905074779255266 | 60 | 235903060064761242 | 71 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.6 | 34.2 | 36 | 36.3 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 4 | 7/21/2022 11:34:22 AM -05:00 | 3 | 7/21/2022 11:52:47 AM -05:00 | 3 | 7/21/2022 12:33:05 PM -05:00 | 7/21/2022 12:37:52 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R5 |  | 7842e1f8-1209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 5 | 010081427802028410111464 2100888 17230505 | 1 | 4 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100849 17221209 | 12 | TestStrips3 | 5 | 010081427802019210108908 2100289 17220725 | TestStrips1 | 5 | 12 | 010081427802027710118705 2105880 17240328 | 4 | 3 | V790 | 010081427802027710118705 2105880 17240328 | 4 | 3 | V790 | 235905074779227267 | 61 |  |  | 235905074779255266 | 61 | 235905074779255266 | 61 | 235903060064761242 | 72 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.6 | 34.2 | 36 | 36.3 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 5 | 7/21/2022 11:34:22 AM -05:00 | 4 | 7/21/2022 11:52:47 AM -05:00 | 4 | 7/21/2022 12:33:05 PM -05:00 | 7/21/2022 12:37:52 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R6 |  | 7e42e1f8-1209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 6 | 010081427802028410111464 2100888 17230505 | 2 | 4 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100849 17221209 | 11 | TestStrips3 | 5 | 010081427802019210108908 2100289 17220725 | TestStrips1 | 5 | 11 | 010081427802027710118705 2105880 17240328 | 5 | 3 | V790 | 010081427802027710118705 2105880 17240328 | 5 | 3 | V790 | 235905074779227267 | 62 |  |  | 235905074779255266 | 62 | 235905074779255266 | 62 | 235903060064761242 | 52 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.6 | 34.2 | 35.6 | 36.4 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 6 | 7/21/2022 11:34:22 AM -05:00 | 5 | 7/21/2022 11:52:47 AM -05:00 | 1, 1 | 7/21/2022 12:28:16 PM -05:00 | 7/21/2022 12:30:23 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R7 |  | 8442e1f8-1209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 7 | 010081427802028410111464 2100888 17230505 | 3 | 4 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100849 17221209 | 10 | TestStrips3 | 5 | 010081427802019210108908 2100289 17220725 | TestStrips1 | 5 | 10 | 010081427802027710118705 2105880 17240328 | 6 | 3 | V790 | 010081427802027710118705 2105880 17240328 | 6 | 3 | V790 | 235905074779227267 | 63 |  |  | 235905074779255266 | 63 | 235905074779255266 | 63 | 235903060064761242 | 58 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.6 | 34.2 | 35.5 | 36.5 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 7 | 7/21/2022 11:34:22 AM -05:00 | 6 | 7/21/2022 11:52:47 AM -05:00 | 2 | 7/21/2022 12:17:49 PM -05:00 | 7/21/2022 12:28:15 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R8 |  | 8a42e1f8-1209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 8 | 010081427802028410111464 2100888 17230505 | 4 | 4 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100849 17221209 | 9 | TestStrips3 | 5 | 010081427802019210108908 2100289 17220725 | TestStrips1 | 5 | 9 | 010081427802027710118705 2105880 17240328 | 7 | 3 | V790 | 010081427802027710118705 2105880 17240328 | 7 | 3 | V790 | 235905074779227267 | 64 |  |  | 235905074779255266 | 64 | 235905074779255266 | 64 | 235903060064761242 | 59 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.6 | 34.2 | 35.5 | 36.5 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 8 | 7/21/2022 11:34:22 AM -05:00 | 7 | 7/21/2022 11:52:47 AM -05:00 | 3 | 7/21/2022 12:17:49 PM -05:00 | 7/21/2022 12:28:15 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R1 |  | 5a7b815a-1309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 1 | 010081427802028410111464 2100884 17230505 | 11 | 2 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100848 17221209 | 5 | TestStrips3 | 4 | 010081427802019210108908 2100293 17220725 | TestStrips1 | 4 | 5 | 010081427802027710118705 2105881 17240328 | 7 | 2 | V834 | 010081427802027710118705 2105881 17240328 | 7 | 2 | V834 | 235905074779227267 | 65 |  |  | 235905074779255266 | 65 | 235905074779255266 | 65 | 235903060064761242 | 75 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.9 | 34.9 | 36.6 | 37.8 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 1 | 7/21/2022 11:37:07 AM -05:00 | 3 | 7/21/2022 11:55:18 AM -05:00 | 3 | 7/21/2022 12:38:37 PM -05:00 | 7/21/2022 12:48:05 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R2 |  | 607b815a-1309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 2 | 010081427802028410111464 2100884 17230505 | 12 | 2 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100848 17221209 | 4 | TestStrips3 | 4 | 010081427802019210108908 2100293 17220725 | TestStrips1 | 4 | 4 | 010081427802027710118705 2105881 17240328 | 8 | 2 | V834 | 010081427802027710118705 2105881 17240328 | 8 | 2 | V834 | 235905074779227267 | 66 |  |  | 235905074779255266 | 66 | 235905074779255266 | 66 | 235903060064761242 | 76 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.9 | 34.9 | 36.6 | 37.8 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 2 | 7/21/2022 11:37:07 AM -05:00 | 4 | 7/21/2022 11:55:18 AM -05:00 | 4 | 7/21/2022 12:38:37 PM -05:00 | 7/21/2022 12:48:05 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R3 |  | 667b815a-1309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 3 | 010081427802028410111464 2100822 17230505 | 9 | 3 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100848 17221209 | 3 | TestStrips3 | 4 | 010081427802019210108908 2100293 17220725 | TestStrips1 | 4 | 3 | 010081427802027710118705 2105881 17240328 | 9 | 2 | V834 | 010081427802027710118705 2105881 17240328 | 9 | 2 | V834 | 235905074779227267 | 67 |  |  | 235905074779255266 | 67 | 235905074779255266 | 67 | 235903060064761242 | 77 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.9 | 34.2 | 36.6 | 37.8 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 3 | 7/21/2022 11:37:07 AM -05:00 | 5 | 7/21/2022 11:55:18 AM -05:00 | 5 | 7/21/2022 12:38:37 PM -05:00 | 7/21/2022 12:48:05 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R4 |  | 6c7b815a-1309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 4 | 010081427802028410111464 2100822 17230505 | 10 | 3 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100848 17221209 | 2 | TestStrips3 | 4 | 010081427802019210108908 2100293 17220725 | TestStrips1 | 4 | 2 | 010081427802027710118705 2105881 17240328 | 10 | 2 | V834 | 010081427802027710118705 2105881 17240328 | 10 | 2 | V834 | 235905074779227267 | 68 |  |  | 235905074779255266 | 68 | 235905074779255266 | 68 | 235903060064761242 | 78 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.9 | 34.2 | 36.6 | 37.8 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 4 | 7/21/2022 11:37:07 AM -05:00 | 6 | 7/21/2022 11:55:18 AM -05:00 | 6 | 7/21/2022 12:38:37 PM -05:00 | 7/21/2022 12:48:05 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R5 |  | 727b815a-1309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 5 | 010081427802028410111464 2100822 17230505 | 11 | 3 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100848 17221209 | 1 | TestStrips3 | 4 | 010081427802019210108908 2100293 17220725 | TestStrips1 | 4 | 1 | 010081427802027710118705 2105881 17240328 | 11 | 2 | V834 | 010081427802027710118705 2105881 17240328 | 11 | 2 | V834 | 235905074779227267 | 69 |  |  | 235905074779255266 | 69 | 235905074779255266 | 69 | 235903060064761242 | 79 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.9 | 34.2 | 36.6 | 37.8 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 5 | 7/21/2022 11:37:07 AM -05:00 | 7 | 7/21/2022 11:55:18 AM -05:00 | 7 | 7/21/2022 12:38:37 PM -05:00 | 7/21/2022 12:48:05 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R6 |  | 787b815a-1309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 6 | 010081427802028410111464 2100884 17230505 | 13 | 2 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100848 17221209 | 8 | TestStrips3 | 4 | 010081427802019210108908 2100293 17220725 | TestStrips1 | 4 | 8 | 010081427802027710118705 2105881 17240328 | 4 | 2 | V834 | 010081427802027710118705 2105881 17240328 | 4 | 2 | V834 | 235905074779227267 | 70 |  |  | 235905074779255266 | 70 | 235905074779255266 | 70 | 235903060064761242 | 88 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.9 | 35 | 37 | 37.6 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 6 | 7/21/2022 11:37:07 AM -05:00 | 8 | 7/21/2022 12:05:09 PM -05:00 | 8 | 7/21/2022 12:48:05 PM -05:00 | 7/21/2022 12:56:11 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R7 |  | 7e7b815a-1309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 7 | 010081427802028410111464 2100884 17230505 | 14 | 2 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100848 17221209 | 7 | TestStrips3 | 4 | 010081427802019210108908 2100293 17220725 | TestStrips1 | 4 | 7 | 010081427802027710118705 2105881 17240328 | 5 | 2 | V834 | 010081427802027710118705 2105881 17240328 | 5 | 2 | V834 | 235905074779227267 | 71 |  |  | 235905074779255266 | 71 | 235905074779255266 | 71 | 235903060064761242 | 73 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.9 | 35 | 36.6 | 37.8 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 7 | 7/21/2022 11:37:07 AM -05:00 | 1 | 7/21/2022 11:55:18 AM -05:00 | 1 | 7/21/2022 12:38:37 PM -05:00 | 7/21/2022 12:48:05 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R8 |  | 847b815a-1309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 8 | 010081427802028410111464 2100884 17230505 | 15 | 2 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100848 17221209 | 6 | TestStrips3 | 4 | 010081427802019210108908 2100293 17220725 | TestStrips1 | 4 | 6 | 010081427802027710118705 2105881 17240328 | 6 | 2 | V834 | 010081427802027710118705 2105881 17240328 | 6 | 2 | V834 | 235905074779227267 | 72 |  |  | 235905074779255266 | 72 | 235905074779255266 | 72 | 235903060064761242 | 74 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.9 | 35 | 36.6 | 37.8 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 8 | 7/21/2022 11:37:07 AM -05:00 | 2 | 7/21/2022 11:55:18 AM -05:00 | 2 | 7/21/2022 12:38:37 PM -05:00 | 7/21/2022 12:48:05 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R1 |  | b3de19db-1309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 1 | 010081427802028410111464 2100884 17230505 | 4 | 2 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100848 17221209 | 9 | TestStrips3 | 4 | 010081427802019210108908 2100293 17220725 | TestStrips1 | 4 | 9 | 010081427802027710118705 2105881 17240328 | 3 | 2 | V834 | 010081427802027710118705 2105881 17240328 | 3 | 2 | V834 | 235905074779227267 | 73 |  |  | 235905074779255266 | 73 | 235905074779255266 | 73 | 235903060064761242 | 87 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 34.3 | 35 | 37 | 37.6 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 1 | 7/21/2022 11:40:41 AM -05:00 | 7 | 7/21/2022 12:05:09 PM -05:00 | 7 | 7/21/2022 12:48:05 PM -05:00 | 7/21/2022 12:56:11 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R2 |  | b9de19db-1309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 2 | 010081427802028410111424 2100836 17230505 | 5 | 1 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100848 17221209 | 14 | TestStrips3 | 4 | 010081427802019210108908 2100293 17220725 | TestStrips1 | 4 | 14 | 010081427802027710118705 2105882 17240328 | 10 | 1 | V639 | 010081427802027710118705 2105882 17240328 | 10 | 1 | V639 | 235905074779227267 | 74 |  |  | 235905074779255266 | 74 | 235905074779255266 | 74 | 235903060064761242 | 82 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 34.4 | 35 | 35.5 | 36.5 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 2 | 7/21/2022 11:40:41 AM -05:00 | 2 | 7/21/2022 12:05:09 PM -05:00 | 2 | 7/21/2022 12:48:05 PM -05:00 | 7/21/2022 12:56:11 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R3 |  | bfde19db-1309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 3 | 010081427802028410111464 2100884 17230505 | 5 | 2 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100848 17221209 | 13 | TestStrips3 | 4 | 010081427802019210108908 2100293 17220725 | TestStrips1 | 4 | 13 | 010081427802027710118705 2105882 17240328 | 11 | 1 | V639 | 010081427802027710118705 2105882 17240328 | 11 | 1 | V639 | 235905074779227267 | 75 |  |  | 235905074779255266 | 75 | 235905074779255266 | 75 | 235903060064761242 | 83 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 34.4 | 35 | 35.5 | 36.5 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 3 | 7/21/2022 11:40:41 AM -05:00 | 3 | 7/21/2022 12:05:09 PM -05:00 | 3 | 7/21/2022 12:48:05 PM -05:00 | 7/21/2022 12:56:11 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R4 |  | c5de19db-1309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 4 | 010081427802028410111464 2100884 17230505 | 6 | 2 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100848 17221209 | 12 | TestStrips3 | 4 | 010081427802019210108908 2100293 17220725 | TestStrips1 | 4 | 12 | 010081427802027710118705 2105882 17240328 | 12 | 1 | V639 | 010081427802027710118705 2105882 17240328 | 12 | 1 | V639 | 235905074779227267 | 76 |  |  | 235905074779255266 | 76 | 235905074779255266 | 76 | 235903060064761242 | 84 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 34.4 | 35 | 35.5 | 36.5 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 4 | 7/21/2022 11:40:41 AM -05:00 | 4 | 7/21/2022 12:05:09 PM -05:00 | 4 | 7/21/2022 12:48:05 PM -05:00 | 7/21/2022 12:56:11 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R5 |  | cbde19db-1309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 5 | 010081427802028410111464 2100884 17230505 | 7 | 2 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100848 17221209 | 11 | TestStrips3 | 4 | 010081427802019210108908 2100293 17220725 | TestStrips1 | 4 | 11 | 010081427802027710118705 2105881 17240328 | 1 | 2 | V834 | 010081427802027710118705 2105881 17240328 | 1 | 2 | V834 | 235905074779227267 | 77 |  |  | 235905074779255266 | 77 | 235905074779255266 | 77 | 235903060064761242 | 85 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 34.4 | 35 | 37 | 37.6 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 5 | 7/21/2022 11:40:41 AM -05:00 | 5 | 7/21/2022 12:05:09 PM -05:00 | 5 | 7/21/2022 12:48:05 PM -05:00 | 7/21/2022 12:56:11 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R6 |  | d1de19db-1309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 6 | 010081427802028410111464 2100884 17230505 | 8 | 2 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100848 17221209 | 10 | TestStrips3 | 4 | 010081427802019210108908 2100293 17220725 | TestStrips1 | 4 | 10 | 010081427802027710118705 2105881 17240328 | 2 | 2 | V834 | 010081427802027710118705 2105881 17240328 | 2 | 2 | V834 | 235905074779227267 | 78 |  |  | 235905074779255266 | 78 | 235905074779255266 | 78 | 235903060064761242 | 86 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 34.4 | 35 | 37 | 37.6 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 6 | 7/21/2022 11:40:41 AM -05:00 | 6 | 7/21/2022 12:05:09 PM -05:00 | 6 | 7/21/2022 12:48:05 PM -05:00 | 7/21/2022 12:56:11 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R7 |  | d7de19db-1309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 7 | 010081427802028410111464 2100884 17230505 | 9 | 2 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100848 17221209 | 16 | TestStrips3 | 4 | 010081427802019210108908 2100293 17220725 | TestStrips1 | 4 | 16 | 010081427802027710118705 2105882 17240328 | 8 | 1 | V639 | 010081427802027710118705 2105882 17240328 | 8 | 1 | V639 | 235905074779227267 | 79 |  |  | 235905074779255266 | 79 | 235905074779255266 | 79 | 235903060064761242 | 96 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 34.4 | 35 | 35.8 | 36.6 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 7 | 7/21/2022 11:40:41 AM -05:00 | 5 | 7/21/2022 12:14:35 PM -05:00 | 5 | 7/21/2022 12:56:11 PM -05:00 | 7/21/2022 1:03:22 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R8 |  | ddde19db-1309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 8 | 010081427802028410111464 2100884 17230505 | 10 | 2 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100848 17221209 | 15 | TestStrips3 | 4 | 010081427802019210108908 2100293 17220725 | TestStrips1 | 4 | 15 | 010081427802027710118705 2105882 17240328 | 9 | 1 | V639 | 010081427802027710118705 2105882 17240328 | 9 | 1 | V639 | 235905074779227267 | 80 |  |  | 235905074779255266 | 80 | 235905074779255266 | 80 | 235903060064761242 | 81 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 34.4 | 35 | 35.5 | 36.5 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 8 | 7/21/2022 11:40:41 AM -05:00 | 1 | 7/21/2022 12:05:09 PM -05:00 | 1 | 7/21/2022 12:48:05 PM -05:00 | 7/21/2022 12:56:11 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R9 |  | 65fdae7a-1409-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 9 | 010081427802028410111424 2100836 17230505 | 1 | 1 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100847 17221209 | 4 | TestStrips3 | 3 | 010081427802019210108908 2100207 17220725 | TestStrips1 | 3 | 4 | 010081427802027710118705 2105882 17240328 | 4 | 1 | V639 | 010081427802027710118705 2105882 17240328 | 4 | 1 | V639 | 235905074779255266 | 85 |  |  | 235905074779227267 | 43 | 235905074779227267 | 43 | 235903060064761242 | 92 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 34.9 | 34.8 | 35.8 | 36.6 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 1 | 7/21/2022 11:45:09 AM -05:00 | 1 | 7/21/2022 12:14:35 PM -05:00 | 1 | 7/21/2022 12:56:11 PM -05:00 | 7/21/2022 1:03:22 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R9 |  | 6bfdae7a-1409-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 9 | 010081427802028410111464 2100884 17230505 | 1 | 2 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100847 17221209 | 3 | TestStrips3 | 3 | 010081427802019210108908 2100207 17220725 | TestStrips1 | 3 | 3 | 010081427802027710118705 2105882 17240328 | 5 | 1 | V639 | 010081427802027710118705 2105882 17240328 | 5 | 1 | V639 | 235905074779255266 | 86 |  |  | 235905074779227267 | 44 | 235905074779227267 | 44 | 235903060064761242 | 93 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 34.9 | 34.8 | 35.8 | 36.6 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 2 | 7/21/2022 11:45:09 AM -05:00 | 2 | 7/21/2022 12:14:35 PM -05:00 | 2 | 7/21/2022 12:56:11 PM -05:00 | 7/21/2022 1:03:22 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R9 |  | 71fdae7a-1409-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 9 | 010081427802028410111464 2100884 17230505 | 2 | 2 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100847 17221209 | 2 | TestStrips3 | 3 | 010081427802019210108908 2100207 17220725 | TestStrips1 | 3 | 2 | 010081427802027710118705 2105882 17240328 | 6 | 1 | V639 | 010081427802027710118705 2105882 17240328 | 6 | 1 | V639 | 235905074779255266 | 87 |  |  | 235905074779227267 | 45 | 235905074779227267 | 45 | 235903060064761242 | 94 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 34.9 | 34.8 | 35.8 | 36.6 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 3 | 7/21/2022 11:45:09 AM -05:00 | 3 | 7/21/2022 12:14:35 PM -05:00 | 3 | 7/21/2022 12:56:11 PM -05:00 | 7/21/2022 1:03:22 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R9 |  | 77fdae7a-1409-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 9 | 010081427802028410111464 2100884 17230505 | 3 | 2 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100847 17221209 | 1 | TestStrips3 | 3 | 010081427802019210108908 2100207 17220725 | TestStrips1 | 3 | 1 | 010081427802027710118705 2105882 17240328 | 7 | 1 | V639 | 010081427802027710118705 2105882 17240328 | 7 | 1 | V639 | 235905074779255266 | 88 |  |  | 235905074779227267 | 46 | 235905074779227267 | 46 | 235903060064761242 | 95 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 34.9 | 34.8 | 35.8 | 36.6 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 4 | 7/21/2022 11:45:09 AM -05:00 | 4 | 7/21/2022 12:14:35 PM -05:00 | 4 | 7/21/2022 12:56:11 PM -05:00 | 7/21/2022 1:03:22 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R10 |  | f47fc1ae-1609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 10 | 010081427802028410111464 2100874 17230505 | 22 | 4 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100847 17221209 | 7 | TestStrips3 | 3 | 010081427802019210108908 2100207 17220725 | TestStrips1 | 3 | 7 | 010081427802027710118705 2105882 17240328 | 1 | 1 | V639 | 010081427802027710118705 2105882 17240328 | 1 | 1 | V639 | 235905074779227267 | 87 |  |  | 235905074779255266 | 82 | 235905074779255266 | 82 | 235903060064761242 | 49 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.9 | 33.9 | 36.5 | 36.4 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 1 | 7/21/2022 12:00:55 PM -05:00 | 1 | 7/21/2022 1:06:06 PM -05:00 | 1 | 7/21/2022 1:20:25 PM -05:00 | 7/21/2022 1:25:19 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R10 |  | fa7fc1ae-1609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 10 | 010081427802028410111464 2100874 17230505 | 23 | 4 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100847 17221209 | 6 | TestStrips3 | 3 | 010081427802019210108908 2100207 17220725 | TestStrips1 | 3 | 6 | 010081427802027710118705 2105882 17240328 | 2 | 1 | V639 | 010081427802027710118705 2105882 17240328 | 2 | 1 | V639 | 235905074779227267 | 88 |  |  | 235905074779255266 | 83 | 235905074779255266 | 83 | 235903060064761242 | 50 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.9 | 33.9 | 36.5 | 36.4 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 2 | 7/21/2022 12:00:55 PM -05:00 | 2 | 7/21/2022 1:06:06 PM -05:00 | 2 | 7/21/2022 1:20:25 PM -05:00 | 7/21/2022 1:25:19 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R10 |  | 0080c1ae-1609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 10 | 010081427802028410111464 2100874 17230505 | 24 | 4 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100847 17221209 | 5 | TestStrips3 | 3 | 010081427802019210108908 2100207 17220725 | TestStrips1 | 3 | 5 | 010081427802027710118705 2105882 17240328 | 3 | 1 | V639 | 010081427802027710118705 2105882 17240328 | 3 | 1 | V639 | 235905074779255266 | 81 |  |  | 235905074779255266 | 84 | 235905074779255266 | 84 | 235903060064761242 | 51 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.9 | 33.9 | 36.5 | 36.4 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 3 | 7/21/2022 12:00:55 PM -05:00 | 3 | 7/21/2022 1:06:06 PM -05:00 | 3 | 7/21/2022 1:20:25 PM -05:00 | 7/21/2022 1:25:19 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R1 |  | a79dad68-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 1 | 010081427802028410111424 2100614 17230505 | 22 | 2 | 0156 | 010081427802024610114545 2100105 17230720 | [20 empty] | 235905074779227267 | 89 |  |  | 235905074779255266 | 89 | [5 empty] | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | [11 empty] | 1 | 7/21/2022 1:03:24 PM -05:00 |  |  |  |  | 7/21/2022 1:06:06 PM -05:00 |  |  |  |  | 0 | 0 | 0 | 0
BKV 10m4 U R2 |  | ad9dad68-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 2 | 010081427802028410111424 2100614 17230505 | 23 | 2 | 0156 | 010081427802024610114545 2100105 17230720 | [20 empty] | 235905074779227267 | 90 |  |  | 235905074779255266 | 90 | [5 empty] | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | [11 empty] | 2 | 7/21/2022 1:03:24 PM -05:00 |  |  |  |  | 7/21/2022 1:06:06 PM -05:00 |  |  |  |  | 0 | 0 | 0 | 0
BKV 10m4 U R3 |  | b39dad68-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 3 | 010081427802028410111424 2100614 17230505 | 24 | 2 | 0156 | 010081427802024610114545 2100105 17230720 | [20 empty] | 235905074779227267 | 91 |  |  | 235905074779255266 | 91 | [5 empty] | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | [11 empty] | 3 | 7/21/2022 1:03:24 PM -05:00 |  |  |  |  | 7/21/2022 1:06:06 PM -05:00 |  |  |  |  | 0 | 0 | 0 | 0
BKV 10m4 U R4 |  | b99dad68-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 4 | 010081427802028410111424 2100622 17230505 | 21 | 3 | 0156 | 010081427802024610114545 2100105 17230720 | [20 empty] | 235905074779227267 | 92 |  |  | 235905074779255266 | 92 | [5 empty] | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | [11 empty] | 4 | 7/21/2022 1:03:24 PM -05:00 |  |  |  |  | 7/21/2022 1:06:06 PM -05:00 |  |  |  |  | 0 | 0 | 0 | 0
BKV 10m4 U R5 |  | bf9dad68-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 5 | 010081427802028410111424 2100622 17230505 | 22 | 3 | 0156 | 010081427802024610114545 2100105 17230720 | [20 empty] | 235905074779227267 | 93 |  |  | 235905074779255266 | 93 | [5 empty] | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | [11 empty] | 5 | 7/21/2022 1:03:24 PM -05:00 |  |  |  |  | 7/21/2022 1:06:06 PM -05:00 |  |  |  |  | 0 | 0 | 0 | 0
BKV 10m4 U R7 |  | c59dad68-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 7 | 010081427802028410111424 2100622 17230505 | 23 | 3 | 0156 | 010081427802024610114545 2100105 17230720 | [20 empty] | 235905074779227267 | 94 |  |  | 235905074779255266 | 94 | [5 empty] | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | [11 empty] | 6 | 7/21/2022 1:03:24 PM -05:00 |  |  |  |  | 7/21/2022 1:06:06 PM -05:00 |  |  |  |  | 0 | 0 | 0 | 0
BKV 10m4 U R10 |  | cb9dad68-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 10 | 010081427802028410111424 2100622 17230505 | 24 | 3 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100847 17221209 | 9 | TestStrips3 | 3 | 010081427802019210108908 2100207 17220725 | TestStrips1 | 3 | 9 | 010081427802027710118705 2105883 17240328 | 11 | 4 | V1169 | 010081427802027710118705 2105883 17240328 | 11 | 4 | V1169 | 235905074779227267 | 95 |  |  | 235905074779255266 | 95 | 235905074779255266 | 95 | 235903060064761242 | 67 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.5 | 33.9 | 35.1 | 35.6 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 7 | 7/21/2022 1:03:24 PM -05:00 | 1 | 7/21/2022 1:16:44 PM -05:00 | 1 | 7/21/2022 1:26:52 PM -05:00 | 7/21/2022 1:35:04 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R11 |  | d29dad68-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 11 | 010081427802028410111464 2100874 17230505 | 21 | 4 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100847 17221209 | 8 | TestStrips3 | 3 | 010081427802019210108908 2100207 17220725 | TestStrips1 | 3 | 8 | 010081427802027710118705 2105883 17240328 | 12 | 4 | V1169 | 010081427802027710118705 2105883 17240328 | 12 | 4 | V1169 | 235905074779227267 | 96 |  |  | 235905074779255266 | 96 | 235905074779255266 | 96 | 235903060064761242 | 68 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.5 | 33.9 | 35.1 | 35.6 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 8 | 7/21/2022 1:03:24 PM -05:00 | 2 | 7/21/2022 1:16:44 PM -05:00 | 2 | 7/21/2022 1:26:52 PM -05:00 | 7/21/2022 1:35:04 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R11 |  | 1f42d920-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 11 | 010081427802028410111424 2100622 17230505 | 18 | 3 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100847 17221209 | 15 | TestStrips3 | 3 | 010081427802019210108908 2100207 17220725 | TestStrips1 | 3 | 15 | 010081427802027710118705 2105883 17240328 | 5 | 4 | V1169 | 010081427802027710118705 2105883 17240328 | 5 | 4 | V1169 | 235905074779237266 | 3 |  |  | 235905074779225267 | 2 | 235905074779225267 | 2 | 235903060064727242 | 3 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.6 | 34 | 35.4 | 35.4 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 1 | 7/21/2022 1:08:34 PM -05:00 | 1 | 7/21/2022 1:35:04 PM -05:00 | 1 | 7/21/2022 1:46:56 PM -05:00 | 7/21/2022 1:52:19 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R12 |  | 2642d920-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 12 | 010081427802028410111424 2100622 17230505 | 19 | 3 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100847 17221209 | 14 | TestStrips3 | 3 | 010081427802019210108908 2100207 17220725 | TestStrips1 | 3 | 14 | 010081427802027710118705 2105883 17240328 | 6 | 4 | V1169 | 010081427802027710118705 2105883 17240328 | 6 | 4 | V1169 | 235905074779237266 | 4 |  |  | 235905074779225267 | 3 | 235905074779225267 | 3 | 235903060064727242 | 4 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.6 | 34 | 35.4 | 35.4 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 2 | 7/21/2022 1:08:34 PM -05:00 | 2 | 7/21/2022 1:35:04 PM -05:00 | 2 | 7/21/2022 1:46:56 PM -05:00 | 7/21/2022 1:52:19 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R11 |  | 2c42d920-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 11 | 010081427802028410111424 2100622 17230505 | 20 | 3 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100847 17221209 | 13 | TestStrips3 | 3 | 010081427802019210108908 2100207 17220725 | TestStrips1 | 3 | 13 | 010081427802027710118705 2105883 17240328 | 7 | 4 | V1169 | 010081427802027710118705 2105883 17240328 | 7 | 4 | V1169 | 235905074779237266 | 5 |  |  | 235905074779225267 | 4 | 235905074779225267 | 4 | 235903060064727242 | 5 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.6 | 34 | 35.4 | 35.4 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 3 | 7/21/2022 1:08:34 PM -05:00 | 3 | 7/21/2022 1:35:04 PM -05:00 | 3 | 7/21/2022 1:46:56 PM -05:00 | 7/21/2022 1:52:19 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R12 |  | 3242d920-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 12 | 010081427802028410111464 2100874 17230505 | 17 | 4 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100847 17221209 | 12 | TestStrips3 | 3 | 010081427802019210108908 2100207 17220725 | TestStrips1 | 3 | 12 | 010081427802027710118705 2105883 17240328 | 8 | 4 | V1169 | 010081427802027710118705 2105883 17240328 | 8 | 4 | V1169 | 235905074779237266 | 6 |  |  | 235905074779225267 | 5 | 235905074779225267 | 5 | 235903060064727242 | 6 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.6 | 34 | 35.4 | 35.4 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 4 | 7/21/2022 1:08:34 PM -05:00 | 4 | 7/21/2022 1:35:04 PM -05:00 | 4 | 7/21/2022 1:46:56 PM -05:00 | 7/21/2022 1:52:19 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R11 |  | 3842d920-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 11 | 010081427802028410111464 2100874 17230505 | 18 | 4 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100847 17221209 | 11 | TestStrips3 | 3 | 010081427802019210108908 2100207 17220725 | TestStrips1 | 3 | 11 | 010081427802027710118705 2105883 17240328 | 9 | 4 | V1169 | 010081427802027710118705 2105883 17240328 | 9 | 4 | V1169 | 235905074779237266 | 7 |  |  | 235905074779225267 | 6 | 235905074779225267 | 6 | 235903060064727242 | 7 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.6 | 34 | 35.4 | 35.4 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 5 | 7/21/2022 1:08:34 PM -05:00 | 5 | 7/21/2022 1:35:04 PM -05:00 | 5 | 7/21/2022 1:46:56 PM -05:00 | 7/21/2022 1:52:19 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R12 |  | 3e42d920-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 12 | 010081427802028410111464 2100874 17230505 | 19 | 4 | 0156 | 010081427802024610114545 2100105 17230720 |  |  |  |  | 010081427802030710108263 2100847 17221209 | 10 | TestStrips3 | 3 | 010081427802019210108908 2100207 17220725 | TestStrips1 | 3 | 10 | 010081427802027710118705 2105883 17240328 | 10 | 4 | V1169 | 010081427802027710118705 2105883 17240328 | 10 | 4 | V1169 | 235905074779237266 | 8 |  |  | 235905074779225267 | 7 | 235905074779225267 | 7 | 235903060064727242 | 8 |  | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | App: 0.35.10, Bolo: 1.9.0 | App: 0.57.2, Bolo: 1.9.0 | 010081427802022210118785 2100082 17230407 | 0100814278020215100M2329 2100198 17240111 | A | 33.6 | 34 | 35.4 | 35.4 | f294bea5-0509-ed11-85f4-5cf370a20867 |  | 6 | 7/21/2022 1:08:34 PM -05:00 | 6 | 7/21/2022 1:35:04 PM -05:00 | 6 | 7/21/2022 1:46:56 PM -05:00 | 7/21/2022 1:52:19 PM -05:00 |  |  | 38281 | 108908 | 1 | 1 | 1 | 1
BKV 10m4 U R12 |  | 4442d920-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 12 | 010081427802028410111464 2100874 17230505 | 20 | 4 | 0156 | 010081427802024610114545 2100105 17230720 | [20 empty] | 235905074779225267 | 1 |  |  | 235905074779225267 | 8 | [5 empty] | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | [11 empty] | 7 | 7/21/2022 1:08:34 PM -05:00 |  |  |  |  | 7/21/2022 1:14:54 PM -05:00 |  |  |  |  | 0 | 0 | 0 | 0
BKV 10m4 U R3 |  | b7477ca2-3609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 3 | 010081427802028410111424 2100614 17230505 | 21 | 2 | 0156 | 010081427802024610114545 2100105 17230720 | [20 empty] | 235905074779237266 | 1 |  |  | 235905074779237266 | 2 | [5 empty] | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | [11 empty] | 1 | 7/21/2022 3:49:39 PM -05:00 |  |  |  |  | 7/21/2022 3:51:13 PM -05:00 |  |  |  |  | 0 | 0 | 0 | 0
BKV 10m4 U R4 |  | f5c2a3da-3609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 4 | 010081427802028410111424 2100896 17230505 | 24 | 1 | 0156 | 010081427802024610114545 2100105 17230720 | [20 empty] | 235905074779225267 | 15 |  |  | 235905074779225267 | 16 | [5 empty] | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | [11 empty] | 1 | 7/21/2022 3:51:13 PM -05:00 |  |  |  |  | 7/21/2022 3:52:49 PM -05:00 |  |  |  |  | 0 | 0 | 0 | 0
BKV 10m4 U R7 |  | 90b7eb13-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 7 | 010081427802028410111424 2100896 17230505 | 23 | 1 | 0156 | 010081427802024610114545 2100105 17230720 | [20 empty] | 235905074779225267 | 13 |  |  | 235905074779225267 | 14 | [5 empty] | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | [11 empty] | 1 | 7/21/2022 3:52:50 PM -05:00 |  |  |  |  | 7/21/2022 3:54:26 PM -05:00 |  |  |  |  | 0 | 0 | 0 | 0
BKV 10m4 U R8 |  | f2bbce4d-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 8 | 010081427802028410111424 2100896 17230505 | 22 | 1 | 0156 | 010081427802024610114545 2100105 17230720 | [20 empty] | 235905074779225267 | 11 |  |  | 235905074779225267 | 12 | [5 empty] | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | [11 empty] | 1 | 7/21/2022 3:54:27 PM -05:00 |  |  |  |  | 7/21/2022 3:56:02 PM -05:00 |  |  |  |  | 0 | 0 | 0 | 0
BKV 10m4 U R9 |  | 0e592f87-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 5 | 9 | 010081427802028410111424 2100896 17230505 | 21 | 1 | 0156 | 010081427802024610114545 2100105 17230720 | [20 empty] | 235905074779225267 | 9 |  |  | 235905074779225267 | 10 | [5 empty] | BKV URINE 10 | 0.0.40 | 1.9.2.6 | 4.5.0.5217 | G636 | 3.4S e 2016-05-02 | 12000112 | 0.9.24 | App: 0.15.4, Bolo: 1.9.0 | [11 empty] | 1 | 7/21/2022 3:56:05 PM -05:00 |  |  |  |  | 7/21/2022 3:57:41 PM -05:00 |  |  |  |  | 0 | 0 | 0 | 0

## Yellow_530_555
Sample ID | Patient ID | Assay Name | Result Code | Assay Version | Test Guid | Target Result Guid | Test Comment | Replicate Number | Extraction Number | Elution Number | Primer/Probe | Calibration Coefficient | Target Name | Target Setting Specimen Type | Replicate Result | Localized Result | Flags | Ct | Max Peak Height | EPR | End Point Fluorescence | Log Conc. | Log Conc. with Unit Conv. | Conc. | Conc. with Unit Conv. | Log Conc. with Dilution Factor | Log Conc. with Dilution Factor Unit Conv. | Log Conc. Precision | Log Conc. Conv. Precision | Conc. with Dilution Factor | Conc. with Dilution Factor Unit Conv. | Conc. Precision | Conc. Conv. Precision | Log Conc. with Quant Adj. | Log Conc. with Quant Adj. Unit Conv. | Conc. with Quant Adj. | Conc. with Quant Adj. Unit Conv. | Log Intermed. Conc. | Log Intermed. Conc. with Unit Conv. | Intermed. Conc. | Intermed. Conc. with Unit Conv. | Cartridge Barcode | Cartridge Lane | XPCR Module Index | XPCR Module Serial | Start Time | Blank Reading | Dark Reading | Baseline Slope | Baseline Y Intercept | Baseline First Cycle | Baseline Last Cycle | Readings 1 | Readings 2 | Readings 3 | Readings 4 | Readings 5 | Readings 6 | Readings 7 | Readings 8 | Readings 9 | Readings 10 | Readings 11 | Readings 12 | Readings 13 | Readings 14 | Readings 15 | Readings 16 | Readings 17 | Readings 18 | Readings 19 | Readings 20 | Readings 21 | Readings 22 | Readings 23 | Readings 24 | Readings 25 | Readings 26 | Readings 27 | Readings 28 | Readings 29 | Readings 30 | Readings 31 | Readings 32 | Readings 33 | Readings 34 | Readings 35 | Readings 36 | Readings 37 | Readings 38 | Readings 39 | Readings 40 | Readings 41 | Readings 42 | Readings 43 | Readings 44 | Readings 45 | Readings 46 | Readings 47 | Readings 48 | Readings 49 | Readings 50 | Readings 51 | Readings 52 | Readings 53 | Readings 54 | Readings 55 | Readings 56 | Readings 57 | Readings 58 | Readings 59 | Readings 60 | Readings 61 | Readings 62 | Readings 63 | Readings 64 | Readings 65 | Readings 66 | Readings 67 | Readings 68 | Readings 69 | Readings 70 | Readings 71 | Readings 72 | Readings 73 | Readings 74 | Readings 75 | Readings 76 | Readings 77 | Readings 78 | Readings 79 | Readings 80 | Readings 81 | Readings 82 | Readings 83 | Readings 84 | Readings 85 | Readings 86 | Readings 87 | Readings 88 | Readings 89 | Readings 90 | Readings 91 | Readings 92 | Readings 93 | Readings 94 | Readings 95 | Readings 96 | Readings 97 | Readings 98 | Readings 99 | Readings 100 | Readings 101 | Readings 102 | Readings 103 | Readings 104 | Readings 105 | Readings 106 | Readings 107 | Readings 108 | Readings 109 | Readings 110 | Readings 111 | Readings 112 | Readings 113 | Readings 114 | Readings 115 | Readings 116 | Readings 117 | Readings 118 | Readings 119 | Readings 120 | Readings 121 | Readings 122 | Readings 123 | Readings 124 | Readings 125 | Readings 126 | Readings 127 | Readings 128 | Readings 129 | Readings 130 | Readings 131 | Readings 132 | Readings 133 | Readings 134 | Readings 135 | Readings 136 | Readings 137 | Readings 138 | Readings 139 | Readings 140 | Readings 141 | Readings 142 | Readings 143 | Readings 144 | Readings 145 | Readings 146 | Readings 147 | Readings 148 | Readings 149 | Readings 150 | Readings 151 | Readings 152 | Readings 153 | Readings 154 | Readings 155 | Readings 156 | Readings 157 | Readings 158 | Readings 159 | Readings 160 | Readings 161 | Readings 162 | Readings 163 | Readings 164 | Readings 165 | Readings 166 | Readings 167 | Readings 168 | Readings 169 | Readings 170 | Readings 171 | Readings 172 | Readings 173 | Readings 174 | Readings 175 | Readings 176 | Readings 177 | Readings 178 | Readings 179 | Readings 180 | Readings 181 | Readings 182 | Readings 183 | Readings 184 | Readings 185 | Readings 186 | Readings 187 | Readings 188 | Readings 189 | Readings 190 | Readings 191 | Readings 192 | Readings 193 | Readings 194 | Readings 195 | Readings 196 | Readings 197 | Readings 198 | Readings 199 | Readings 200 | Readings 201 | Readings 202 | Readings 203 | Readings 204 | Readings 205 | Readings 206 | Readings 207 | Readings 208 | Readings 209 | Readings 210 | Readings 211 | Readings 212 | Readings 213 | Readings 214 | Readings 215 | Readings 216 | Readings 217 | Readings 218 | Readings 219 | Readings 220 | Readings 221 | Readings 222 | Readings 223 | Readings 224 | Readings 225 | Readings 226 | Readings 227 | Readings 228 | Readings 229 | Readings 230 | Readings 231 | Readings 232 | Readings 233 | Readings 234 | Readings 235 | Readings 236 | Readings 237 | Readings 238 | Readings 239 | Readings 240 | Readings 241 | Readings 242 | Readings 243 | Readings 244 | Readings 245 | Readings 246 | Readings 247 | Readings 248 | Readings 249 | Readings 250 | Readings 251 | Readings 252 | Readings 253 | Readings 254 | Readings 255
Raw
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 0d5c9e96-1209-ed11-85f4-5cf370a20867 | 5fc854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105880 17240328 | 8 | 3 | V790 | 2022-07-21 11:32 | 501 | 259 | 18.9314 | 2168.8451 | 5 | 27 | 2353 | 2344 | 2332 | 2339 | 2349 | 2351 | 2345 | 2348 | 2352 | 2358 | 2350 | 2363 | 2367 | 2377 | 2370 | 2382 | 2414 | 2437 | 2503 | 2533 | 2603 | 2626 | 2662 | 2677 | 2669 | 2701 | 2709 | 2720 | 2739 | 2748 | 2738 | 2758 | 2749 | 2767 | 2761 | 2780 | 2781 | 2787 | 2788 | 2785 | 2803 | 2793 | 2794 | 2806 | 2814
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 145c9e96-1209-ed11-85f4-5cf370a20867 | 01c854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.27 | [21 empty] | 010081427802027710118705 2105880 17240328 | 9 | 3 | V790 | 2022-07-21 11:32 | 555 | 259 | 27.7921 | 2118.1265 | 5 | 27 | 2354 | 2369 | 2368 | 2376 | 2385 | 2399 | 2385 | 2382 | 2381 | 2400 | 2398 | 2397 | 2395 | 2402 | 2406 | 2433 | 2468 | 2517 | 2597 | 2675 | 2726 | 2803 | 2826 | 2857 | 2892 | 2909 | 2908 | 2929 | 2933 | 2953 | 2956 | 2968 | 2974 | 2983 | 2969 | 2984 | 2992 | 2990 | 2995 | 3024 | 3006 | 3012 | 3032 | 3029 | 3021
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 1b5c9e96-1209-ed11-85f4-5cf370a20867 | 4aca54cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105880 17240328 | 10 | 3 | V790 | 2022-07-21 11:32 | 524 | 267 | 20.2032 | 2087.4016 | 5 | 27 | 2265 | 2276 | 2262 | 2263 | 2269 | 2274 | 2271 | 2294 | 2294 | 2277 | 2288 | 2278 | 2294 | 2306 | 2309 | 2322 | 2358 | 2379 | 2442 | 2508 | 2530 | 2580 | 2603 | 2624 | 2632 | 2651 | 2665 | 2658 | 2674 | 2686 | 2693 | 2683 | 2722 | 2728 | 2700 | 2726 | 2716 | 2718 | 2712 | 2721 | 2734 | 2728 | 2748 | 2735 | 2727
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 225c9e96-1209-ed11-85f4-5cf370a20867 | e9c954cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105880 17240328 | 11 | 3 | V790 | 2022-07-21 11:32 | 509 | 246 | 19.9302 | 2120.551 | 5 | 27 | 2283 | 2311 | 2294 | 2318 | 2292 | 2307 | 2290 | 2319 | 2328 | 2310 | 2329 | 2322 | 2322 | 2331 | 2327 | 2354 | 2378 | 2408 | 2482 | 2530 | 2559 | 2598 | 2623 | 2647 | 2673 | 2679 | 2693 | 2678 | 2700 | 2723 | 2716 | 2711 | 2717 | 2739 | 2729 | 2741 | 2741 | 2763 | 2748 | 2763 | 2751 | 2752 | 2770 | 2766 | 2782
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 295c9e96-1209-ed11-85f4-5cf370a20867 | 87c954cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.52 | 55.01 | 1.28 | 468 | [20 empty] | 010081427802027710118705 2105880 17240328 | 12 | 3 | V790 | 2022-07-21 11:32 | 499 | 265 | 4.9 | 2369.2333 | 2 | 4 | 2373 | 2374 | 2384 | 2392 | 2395 | 2400 | 2408 | 2396 | 2403 | 2396 | 2402 | 2406 | 2414 | 2431 | 2427 | 2439 | 2477 | 2540 | 2624 | 2707 | 2771 | 2820 | 2863 | 2892 | 2916 | 2941 | 2947 | 2953 | 2970 | 3001 | 2979 | 3015 | 3004 | 3004 | 3001 | 3031 | 3016 | 3030 | 3038 | 3042 | 3044 | 3050 | 3052 | 3047 | 3064
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 305c9e96-1209-ed11-85f4-5cf370a20867 | e454c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105879 17240328 | 1 | 4 | V1169 | 2022-07-21 11:32 | 785 | 294 | 23.1577 | 2506.8071 | 5 | 27 | 2720 | 2708 | 2714 | 2721 | 2725 | 2735 | 2726 | 2735 | 2727 | 2741 | 2744 | 2727 | 2767 | 2733 | 2742 | 2766 | 2799 | 2841 | 2897 | 2973 | 3019 | 3066 | 3114 | 3126 | 3135 | 3152 | 3181 | 3194 | 3193 | 3193 | 3201 | 3222 | 3242 | 3247 | 3229 | 3231 | 3236 | 3260 | 3250 | 3254 | 3255 | 3268 | 3259 | 3284 | 3274
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 375c9e96-1209-ed11-85f4-5cf370a20867 | 8554c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105879 17240328 | 2 | 4 | V1169 | 2022-07-21 11:32 | 729 | 292 | 23.2439 | 2407.3328 | 5 | 27 | 2614 | 2623 | 2601 | 2620 | 2626 | 2622 | 2642 | 2633 | 2619 | 2644 | 2639 | 2650 | 2647 | 2658 | 2656 | 2678 | 2698 | 2746 | 2792 | 2858 | 2911 | 2968 | 3002 | 3024 | 3058 | 3067 | 3078 | 3103 | 3082 | 3120 | 3113 | 3143 | 3151 | 3139 | 3165 | 3160 | 3165 | 3152 | 3164 | 3165 | 3173 | 3180 | 3179 | 3188 | 3180
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 3e5c9e96-1209-ed11-85f4-5cf370a20867 | 4455c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105879 17240328 | 3 | 4 | V1169 | 2022-07-21 11:32 | 739 | 274 | 22.2113 | 2319.6806 | 5 | 27 | 2519 | 2507 | 2528 | 2525 | 2533 | 2515 | 2529 | 2533 | 2536 | 2551 | 2539 | 2550 | 2547 | 2549 | 2538 | 2573 | 2600 | 2644 | 2704 | 2790 | 2814 | 2863 | 2879 | 2911 | 2939 | 2934 | 2942 | 2963 | 2962 | 2988 | 2984 | 2990 | 2994 | 3004 | 3006 | 3023 | 3021 | 3018 | 3024 | 3023 | 3043 | 3031 | 3032 | 3040 | 3035
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 6042e1f8-1209-ed11-85f4-5cf370a20867 | f53c8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.52 | 53.44 | 1.22 | 471 | [20 empty] | 010081427802027710118705 2105881 17240328 | 12 | 2 | V834 | 2022-07-21 11:34 | 614 | 257 | 0.7 | 2263.5 | 2 | 4 | 2256 | 2273 | 2271 | 2269 | 2258 | 2261 | 2269 | 2273 | 2255 | 2276 | 2254 | 2271 | 2285 | 2280 | 2264 | 2290 | 2299 | 2349 | 2428 | 2501 | 2551 | 2586 | 2626 | 2642 | 2689 | 2669 | 2692 | 2704 | 2714 | 2735 | 2723 | 2733 | 2741 | 2745 | 2740 | 2744 | 2758 | 2747 | 2749 | 2752 | 2760 | 2774 | 2763 | 2770 | 2766
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 6642e1f8-1209-ed11-85f4-5cf370a20867 | 508c0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105880 17240328 | 1 | 3 | V790 | 2022-07-21 11:34 | 494 | 260 | 20.1514 | 2002.9866 | 5 | 27 | 2184 | 2192 | 2175 | 2191 | 2203 | 2187 | 2200 | 2206 | 2199 | 2209 | 2210 | 2193 | 2214 | 2214 | 2211 | 2212 | 2250 | 2296 | 2341 | 2416 | 2463 | 2481 | 2507 | 2550 | 2564 | 2576 | 2579 | 2602 | 2583 | 2599 | 2604 | 2606 | 2629 | 2637 | 2636 | 2641 | 2643 | 2642 | 2654 | 2642 | 2660 | 2665 | 2663 | 2659 | 2667
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 6c42e1f8-1209-ed11-85f4-5cf370a20867 | 1d8c0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.22 | [21 empty] | 010081427802027710118705 2105880 17240328 | 2 | 3 | V790 | 2022-07-21 11:34 | 513 | 269 | 19.7261 | 2017.5217 | 5 | 27 | 2208 | 2199 | 2200 | 2226 | 2210 | 2202 | 2214 | 2218 | 2204 | 2220 | 2210 | 2220 | 2202 | 2209 | 2207 | 2228 | 2259 | 2303 | 2366 | 2424 | 2476 | 2488 | 2517 | 2545 | 2563 | 2589 | 2573 | 2597 | 2587 | 2623 | 2630 | 2637 | 2642 | 2654 | 2648 | 2657 | 2666 | 2654 | 2660 | 2660 | 2642 | 2670 | 2682 | 2671 | 2691
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 7242e1f8-1209-ed11-85f4-5cf370a20867 | 918b0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold), 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed) | 16.37 | 60.68 | 1.34 | 745 | [20 empty] | 010081427802027710118705 2105880 17240328 | 3 | 3 | V790 | 2022-07-21 11:34 | 535 | 271 | 0.5 | 2230.2333 | 2 | 4 | 2237 | 2228 | 2230 | 2225 | 2237 | 2242 | 2237 | 2242 | 2230 | 2258 | 2242 | 2255 | 2251 | 2261 | 2275 | 2289 | 2325 | 2396 | 2489 | 2581 | 2659 | 2728 | 2775 | 2798 | 2826 | 2862 | 2870 | 2890 | 2902 | 2915 | 2933 | 2956 | 2943 | 2956 | 2959 | 2963 | 2960 | 2993 | 2996 | 2992 | 3007 | 3003 | 3000 | 2997 | 2998
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 7842e1f8-1209-ed11-85f4-5cf370a20867 | 328b0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.27 | 56.35 | 1.32 | 874 | [20 empty] | 010081427802027710118705 2105880 17240328 | 4 | 3 | V790 | 2022-07-21 11:34 | 513 | 267 | -4.1 | 2200.7667 | 2 | 4 | 2204 | 2184 | 2178 | 2191 | 2189 | 2178 | 2198 | 2193 | 2185 | 2180 | 2207 | 2198 | 2195 | 2207 | 2217 | 2232 | 2248 | 2324 | 2414 | 2483 | 2571 | 2614 | 2675 | 2692 | 2715 | 2739 | 2751 | 2792 | 2800 | 2806 | 2807 | 2833 | 2849 | 2852 | 2855 | 2859 | 2849 | 2863 | 2875 | 2870 | 2875 | 2897 | 2901 | 2909 | 2881
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 7e42e1f8-1209-ed11-85f4-5cf370a20867 | bdaadb1d-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 2610 (Informational, Delays in sample processing), 1020 (Informational, Peak Not Detected) |  |  | 1.18 | [21 empty] | 010081427802027710118705 2105880 17240328 | 5 | 3 | V790 | 2022-07-21 11:34 | 499 | 266 | 12.6579 | 1361.7431 | 5 | 27 | 1482 | 1472 | 1488 | 1488 | 1492 | 1481 | 1496 | 1494 | 1465 | 1473 | 1484 | 1482 | 1494 | 1487 | 1479 | 1509 | 1514 | 1558 | 1587 | 1621 | 1654 | 1672 | 1707 | 1691 | 1719 | 1693 | 1711 | 1738 | 1716 | 1744 | 1733 | 1732 | 1728 | 1723 | 1742 | 1741 | 1749 | 1748 | 1757 | 1743 | 1705 | 1727 | 1732 | 1735 | 1758
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 8442e1f8-1209-ed11-85f4-5cf370a20867 | 29c954cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.18 | 50.9 | 1.36 | 699 | [20 empty] | 010081427802027710118705 2105880 17240328 | 6 | 3 | V790 | 2022-07-21 11:34 | 498 | 265 | 2 | 2167.2667 | 2 | 4 | 2170 | 2164 | 2172 | 2177 | 2190 | 2170 | 2185 | 2172 | 2194 | 2165 | 2171 | 2189 | 2195 | 2204 | 2200 | 2263 | 2294 | 2358 | 2431 | 2523 | 2586 | 2645 | 2679 | 2710 | 2753 | 2757 | 2760 | 2760 | 2795 | 2815 | 2819 | 2820 | 2848 | 2869 | 2872 | 2907 | 2892 | 2922 | 2939 | 2941 | 2943 | 2954 | 2939 | 2960 | 2961
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 8a42e1f8-1209-ed11-85f4-5cf370a20867 | bfc854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.18 | [21 empty] | 010081427802027710118705 2105880 17240328 | 7 | 3 | V790 | 2022-07-21 11:34 | 494 | 262 | 16.7034 | 1997.2158 | 5 | 27 | 2132 | 2131 | 2131 | 2144 | 2144 | 2168 | 2165 | 2158 | 2155 | 2164 | 2154 | 2158 | 2168 | 2159 | 2177 | 2196 | 2220 | 2235 | 2294 | 2345 | 2388 | 2410 | 2409 | 2429 | 2448 | 2455 | 2478 | 2489 | 2499 | 2512 | 2511 | 2510 | 2518 | 2514 | 2530 | 2533 | 2541 | 2547 | 2550 | 2551 | 2546 | 2545 | 2559 | 2560 | 2541
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 5a7b815a-1309-ed11-85f4-5cf370a20867 | 503c8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.17 | [21 empty] | 010081427802027710118705 2105881 17240328 | 7 | 2 | V834 | 2022-07-21 11:37 | 563 | 270 | 16.2757 | 2127.102 | 5 | 27 | 2288 | 2264 | 2283 | 2274 | 2279 | 2301 | 2273 | 2295 | 2297 | 2281 | 2294 | 2285 | 2284 | 2284 | 2289 | 2304 | 2336 | 2360 | 2410 | 2449 | 2502 | 2519 | 2555 | 2555 | 2596 | 2578 | 2583 | 2595 | 2601 | 2615 | 2619 | 2626 | 2631 | 2624 | 2631 | 2639 | 2633 | 2639 | 2648 | 2642 | 2652 | 2641 | 2657 | 2665 | 2665
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 607b815a-1309-ed11-85f4-5cf370a20867 | 22597999-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.16 | [21 empty] | 010081427802027710118705 2105881 17240328 | 8 | 2 | V834 | 2022-07-21 11:37 | 569 | 268 | 14.8879 | 2018.9233 | 5 | 27 | 2156 | 2167 | 2165 | 2168 | 2159 | 2170 | 2172 | 2152 | 2169 | 2156 | 2161 | 2180 | 2165 | 2180 | 2170 | 2185 | 2177 | 2238 | 2280 | 2317 | 2359 | 2370 | 2404 | 2421 | 2432 | 2434 | 2457 | 2446 | 2473 | 2463 | 2480 | 2488 | 2465 | 2469 | 2486 | 2483 | 2486 | 2476 | 2489 | 2502 | 2514 | 2495 | 2497 | 2499 | 2506
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 667b815a-1309-ed11-85f4-5cf370a20867 | 6e3e8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.15 | [21 empty] | 010081427802027710118705 2105881 17240328 | 9 | 2 | V834 | 2022-07-21 11:37 | 664 | 265 | 14.7126 | 2116.7802 | 5 | 27 | 2252 | 2267 | 2258 | 2262 | 2253 | 2260 | 2266 | 2269 | 2245 | 2262 | 2253 | 2261 | 2258 | 2272 | 2280 | 2275 | 2305 | 2332 | 2367 | 2415 | 2448 | 2476 | 2494 | 2516 | 2530 | 2520 | 2534 | 2532 | 2554 | 2559 | 2569 | 2564 | 2562 | 2578 | 2560 | 2570 | 2569 | 2578 | 2581 | 2581 | 2581 | 2581 | 2604 | 2602 | 2591
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 6c7b815a-1309-ed11-85f4-5cf370a20867 | 103e8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105881 17240328 | 10 | 2 | V834 | 2022-07-21 11:37 | 659 | 280 | 21.4891 | 2038.9826 | 5 | 27 | 2248 | 2265 | 2247 | 2237 | 2258 | 2245 | 2242 | 2258 | 2248 | 2254 | 2257 | 2251 | 2249 | 2251 | 2262 | 2275 | 2303 | 2343 | 2395 | 2464 | 2545 | 2563 | 2586 | 2631 | 2624 | 2642 | 2653 | 2668 | 2680 | 2676 | 2695 | 2688 | 2695 | 2701 | 2708 | 2702 | 2698 | 2706 | 2707 | 2718 | 2703 | 2714 | 2738 | 2717 | 2730
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 727b815a-1309-ed11-85f4-5cf370a20867 | b23d8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.22 | [21 empty] | 010081427802027710118705 2105881 17240328 | 11 | 2 | V834 | 2022-07-21 11:37 | 627 | 270 | 20.9273 | 2048.0593 | 5 | 27 | 2265 | 2253 | 2260 | 2243 | 2252 | 2265 | 2250 | 2246 | 2256 | 2266 | 2255 | 2266 | 2243 | 2270 | 2261 | 2267 | 2287 | 2347 | 2394 | 2470 | 2526 | 2554 | 2588 | 2598 | 2636 | 2650 | 2652 | 2670 | 2675 | 2665 | 2664 | 2697 | 2691 | 2703 | 2717 | 2707 | 2704 | 2694 | 2719 | 2709 | 2715 | 2696 | 2723 | 2714 | 2754
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 787b815a-1309-ed11-85f4-5cf370a20867 | f0fe3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.24 | [21 empty] | 010081427802027710118705 2105881 17240328 | 4 | 2 | V834 | 2022-07-21 11:37 | 633 | 295 | 24.0296 | 2104.43 | 5 | 27 | 2311 | 2329 | 2332 | 2312 | 2325 | 2329 | 2331 | 2342 | 2345 | 2344 | 2338 | 2345 | 2348 | 2360 | 2349 | 2386 | 2408 | 2452 | 2521 | 2582 | 2652 | 2705 | 2725 | 2722 | 2767 | 2784 | 2782 | 2791 | 2803 | 2816 | 2838 | 2844 | 2837 | 2832 | 2847 | 2840 | 2854 | 2835 | 2863 | 2865 | 2860 | 2875 | 2860 | 2878 | 2883
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 7e7b815a-1309-ed11-85f4-5cf370a20867 | 543d8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.27 | [21 empty] | 010081427802027710118705 2105881 17240328 | 5 | 2 | V834 | 2022-07-21 11:37 | 596 | 273 | 27.6573 | 2095.9352 | 5 | 27 | 2334 | 2333 | 2345 | 2340 | 2351 | 2365 | 2368 | 2369 | 2364 | 2374 | 2372 | 2387 | 2360 | 2391 | 2385 | 2398 | 2435 | 2499 | 2566 | 2638 | 2716 | 2770 | 2821 | 2823 | 2865 | 2881 | 2880 | 2902 | 2911 | 2907 | 2896 | 2923 | 2927 | 2948 | 2952 | 2948 | 2941 | 2949 | 2948 | 2956 | 2974 | 2953 | 2971 | 2972 | 3003
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 847b815a-1309-ed11-85f4-5cf370a20867 | 953c8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.23 | 58.23 | 1.25 | 315 | [20 empty] | 010081427802027710118705 2105881 17240328 | 6 | 2 | V834 | 2022-07-21 11:37 | 617 | 276 | 6.4 | 2303.3333 | 2 | 4 | 2296 | 2335 | 2323 | 2333 | 2321 | 2335 | 2329 | 2334 | 2325 | 2344 | 2327 | 2342 | 2354 | 2338 | 2340 | 2371 | 2389 | 2446 | 2541 | 2612 | 2671 | 2715 | 2759 | 2762 | 2811 | 2801 | 2821 | 2826 | 2845 | 2866 | 2855 | 2883 | 2875 | 2882 | 2859 | 2884 | 2889 | 2873 | 2883 | 2888 | 2897 | 2905 | 2912 | 2906 | 2904
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | b3de19db-1309-ed11-85f4-5cf370a20867 | 3a5f6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.74 | 58.9 | 1.27 | 541 | [20 empty] | 010081427802027710118705 2105881 17240328 | 3 | 2 | V834 | 2022-07-21 11:41 | 592 | 269 | 2.5 | 2335.5 | 2 | 4 | 2338 | 2335 | 2342 | 2348 | 2355 | 2346 | 2357 | 2342 | 2358 | 2354 | 2343 | 2350 | 2367 | 2361 | 2356 | 2374 | 2414 | 2468 | 2540 | 2624 | 2709 | 2763 | 2797 | 2834 | 2853 | 2880 | 2893 | 2910 | 2911 | 2929 | 2936 | 2943 | 2955 | 2945 | 2970 | 2965 | 2967 | 2992 | 2973 | 2979 | 3001 | 3000 | 2997 | 2986 | 2987
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | b9de19db-1309-ed11-85f4-5cf370a20867 | a75e6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.19 | [21 empty] | 010081427802027710118705 2105882 17240328 | 10 | 1 | V639 | 2022-07-21 11:41 | 614 | 260 | 19.8022 | 2195.6174 | 5 | 27 | 2358 | 2387 | 2380 | 2376 | 2379 | 2368 | 2391 | 2397 | 2393 | 2400 | 2396 | 2395 | 2387 | 2402 | 2392 | 2420 | 2445 | 2482 | 2539 | 2597 | 2631 | 2679 | 2702 | 2722 | 2738 | 2746 | 2782 | 2770 | 2781 | 2789 | 2805 | 2813 | 2814 | 2808 | 2811 | 2823 | 2840 | 2826 | 2833 | 2847 | 2837 | 2846 | 2829 | 2859 | 2843
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | bfde19db-1309-ed11-85f4-5cf370a20867 | 185e6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.26 | 56.9 | 1.29 | 510 | [20 empty] | 010081427802027710118705 2105882 17240328 | 11 | 1 | V639 | 2022-07-21 11:41 | 588 | 272 | 3.9 | 2333.5 | 2 | 4 | 2326 | 2352 | 2347 | 2356 | 2344 | 2347 | 2348 | 2352 | 2350 | 2363 | 2373 | 2360 | 2372 | 2370 | 2372 | 2396 | 2447 | 2509 | 2586 | 2674 | 2747 | 2800 | 2825 | 2856 | 2890 | 2886 | 2924 | 2914 | 2934 | 2948 | 2960 | 2966 | 2972 | 2989 | 2980 | 3004 | 2994 | 3012 | 3005 | 2996 | 3010 | 3009 | 3009 | 3016 | 3024
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | c5de19db-1309-ed11-85f4-5cf370a20867 | bffe3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105882 17240328 | 12 | 1 | V639 | 2022-07-21 11:41 | 672 | 260 | 21.7893 | 2250.3186 | 5 | 27 | 2445 | 2458 | 2446 | 2455 | 2461 | 2454 | 2458 | 2449 | 2471 | 2450 | 2471 | 2466 | 2487 | 2486 | 2481 | 2496 | 2524 | 2576 | 2635 | 2671 | 2748 | 2767 | 2799 | 2831 | 2847 | 2859 | 2882 | 2887 | 2903 | 2891 | 2922 | 2908 | 2909 | 2935 | 2911 | 2931 | 2935 | 2937 | 2930 | 2957 | 2950 | 2973 | 2952 | 2966 | 2962
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | cbde19db-1309-ed11-85f4-5cf370a20867 | 32fe3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.23 | [21 empty] | 010081427802027710118705 2105881 17240328 | 1 | 2 | V834 | 2022-07-21 11:41 | 679 | 291 | 23.8111 | 2216.2577 | 5 | 27 | 2440 | 2434 | 2446 | 2445 | 2444 | 2451 | 2440 | 2450 | 2462 | 2446 | 2451 | 2454 | 2448 | 2467 | 2472 | 2481 | 2519 | 2551 | 2622 | 2675 | 2734 | 2783 | 2834 | 2856 | 2892 | 2886 | 2916 | 2908 | 2917 | 2932 | 2953 | 2940 | 2963 | 2959 | 2967 | 2967 | 2972 | 2974 | 2972 | 2973 | 2974 | 3005 | 2994 | 3001 | 3003
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | d1de19db-1309-ed11-85f4-5cf370a20867 | c65f6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.22 | [21 empty] | 010081427802027710118705 2105881 17240328 | 2 | 2 | V834 | 2022-07-21 11:41 | 600 | 276 | 24.1291 | 2174.6047 | 5 | 27 | 2370 | 2387 | 2386 | 2384 | 2387 | 2405 | 2398 | 2410 | 2419 | 2410 | 2423 | 2423 | 2407 | 2447 | 2417 | 2453 | 2473 | 2524 | 2600 | 2659 | 2722 | 2773 | 2788 | 2826 | 2830 | 2840 | 2855 | 2869 | 2869 | 2893 | 2883 | 2888 | 2887 | 2906 | 2900 | 2912 | 2911 | 2929 | 2923 | 2921 | 2918 | 2940 | 2928 | 2936 | 2934
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | d7de19db-1309-ed11-85f4-5cf370a20867 | 674775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105882 17240328 | 8 | 1 | V639 | 2022-07-21 11:41 | 620 | 261 | 21.2957 | 2243.913 | 5 | 27 | 2434 | 2439 | 2439 | 2439 | 2441 | 2442 | 2460 | 2449 | 2438 | 2452 | 2461 | 2433 | 2458 | 2467 | 2476 | 2497 | 2522 | 2561 | 2620 | 2690 | 2716 | 2748 | 2786 | 2821 | 2809 | 2854 | 2838 | 2858 | 2860 | 2869 | 2890 | 2873 | 2900 | 2894 | 2893 | 2891 | 2918 | 2918 | 2916 | 2932 | 2911 | 2933 | 2929 | 2933 | 2937
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | ddde19db-1309-ed11-85f4-5cf370a20867 | 095f6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105882 17240328 | 9 | 1 | V639 | 2022-07-21 11:41 | 661 | 258 | 21.6571 | 2230.5644 | 5 | 27 | 2398 | 2431 | 2422 | 2447 | 2436 | 2430 | 2434 | 2429 | 2424 | 2434 | 2442 | 2462 | 2447 | 2457 | 2461 | 2488 | 2513 | 2547 | 2635 | 2688 | 2731 | 2752 | 2788 | 2787 | 2808 | 2829 | 2844 | 2850 | 2837 | 2852 | 2857 | 2868 | 2870 | 2874 | 2858 | 2894 | 2875 | 2883 | 2901 | 2888 | 2899 | 2892 | 2910 | 2917 | 2908
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 65fdae7a-1409-ed11-85f4-5cf370a20867 | 084775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105882 17240328 | 4 | 1 | V639 | 2022-07-21 11:45 | 684 | 263 | 21.7674 | 2166.2783 | 5 | 27 | 2362 | 2366 | 2361 | 2378 | 2381 | 2387 | 2373 | 2369 | 2372 | 2378 | 2380 | 2379 | 2399 | 2386 | 2396 | 2408 | 2434 | 2485 | 2536 | 2624 | 2653 | 2690 | 2727 | 2743 | 2768 | 2771 | 2799 | 2795 | 2795 | 2826 | 2817 | 2820 | 2836 | 2844 | 2832 | 2851 | 2847 | 2862 | 2850 | 2859 | 2854 | 2866 | 2841 | 2861 | 2863
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 6bfdae7a-1409-ed11-85f4-5cf370a20867 | d74675b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105882 17240328 | 5 | 1 | V639 | 2022-07-21 11:45 | 703 | 280 | 23.3328 | 2329.7447 | 5 | 27 | 2542 | 2523 | 2524 | 2540 | 2549 | 2555 | 2553 | 2541 | 2548 | 2557 | 2548 | 2564 | 2564 | 2578 | 2591 | 2599 | 2644 | 2691 | 2746 | 2814 | 2859 | 2876 | 2935 | 2944 | 2962 | 2964 | 2988 | 2996 | 3000 | 3005 | 3025 | 3044 | 3030 | 3056 | 3061 | 3044 | 3058 | 3046 | 3061 | 3059 | 3065 | 3051 | 3076 | 3087 | 3078
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 71fdae7a-1409-ed11-85f4-5cf370a20867 | 494675b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 15.89 | 59.99 | 1.29 | 667 | [20 empty] | 010081427802027710118705 2105882 17240328 | 6 | 1 | V639 | 2022-07-21 11:45 | 749 | 264 | 1.2 | 2464.1333 | 2 | 4 | 2466 | 2457 | 2463 | 2480 | 2474 | 2470 | 2487 | 2473 | 2482 | 2500 | 2471 | 2489 | 2480 | 2481 | 2501 | 2514 | 2575 | 2635 | 2737 | 2812 | 2878 | 2946 | 2994 | 3051 | 3046 | 3059 | 3087 | 3081 | 3111 | 3103 | 3129 | 3118 | 3142 | 3146 | 3151 | 3142 | 3144 | 3160 | 3171 | 3158 | 3179 | 3180 | 3176 | 3196 | 3185
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 77fdae7a-1409-ed11-85f4-5cf370a20867 | f44775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105882 17240328 | 7 | 1 | V639 | 2022-07-21 11:45 | 611 | 265 | 21.8399 | 2179.8047 | 5 | 27 | 2375 | 2373 | 2373 | 2376 | 2389 | 2388 | 2394 | 2379 | 2382 | 2389 | 2406 | 2391 | 2407 | 2402 | 2417 | 2435 | 2471 | 2505 | 2564 | 2617 | 2675 | 2707 | 2731 | 2765 | 2765 | 2790 | 2804 | 2814 | 2809 | 2814 | 2821 | 2831 | 2830 | 2830 | 2839 | 2851 | 2851 | 2874 | 2840 | 2859 | 2869 | 2874 | 2874 | 2860 | 2863
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | f47fc1ae-1609-ed11-85f4-5cf370a20867 | 536b47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.19 | [21 empty] | 010081427802027710118705 2105882 17240328 | 1 | 1 | V639 | 2022-07-21 12:01 | 725 | 269 | 21.8492 | 2389.1866 | 5 | 27 | 2573 | 2580 | 2583 | 2589 | 2580 | 2599 | 2593 | 2591 | 2606 | 2603 | 2605 | 2606 | 2621 | 2624 | 2622 | 2643 | 2657 | 2723 | 2802 | 2835 | 2873 | 2891 | 2946 | 2949 | 2992 | 3019 | 3006 | 3014 | 3030 | 3031 | 3043 | 3044 | 3060 | 3061 | 3054 | 3061 | 3070 | 3072 | 3066 | 3082 | 3115 | 3100 | 3074 | 3082 | 3088
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | fa7fc1ae-1609-ed11-85f4-5cf370a20867 | f46a47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105882 17240328 | 2 | 1 | V639 | 2022-07-21 12:01 | 719 | 260 | 23.9073 | 2499.5613 | 5 | 27 | 2709 | 2709 | 2714 | 2715 | 2729 | 2714 | 2737 | 2717 | 2729 | 2735 | 2728 | 2739 | 2743 | 2746 | 2761 | 2778 | 2815 | 2850 | 2924 | 2996 | 3038 | 3085 | 3113 | 3137 | 3144 | 3145 | 3174 | 3183 | 3209 | 3205 | 3215 | 3228 | 3231 | 3238 | 3243 | 3243 | 3249 | 3258 | 3268 | 3256 | 3260 | 3267 | 3268 | 3271 | 3271
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | 0080c1ae-1609-ed11-85f4-5cf370a20867 | b36b47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16 | 50.4 | 1.21 | 402 | [20 empty] | 010081427802027710118705 2105882 17240328 | 3 | 1 | V639 | 2022-07-21 12:01 | 714 | 262 | 3.7 | 2589.5667 | 2 | 4 | 2588 | 2601 | 2597 | 2613 | 2600 | 2613 | 2624 | 2631 | 2620 | 2618 | 2638 | 2625 | 2631 | 2630 | 2641 | 2660 | 2685 | 2746 | 2826 | 2895 | 2946 | 2980 | 2996 | 3037 | 3046 | 3054 | 3070 | 3072 | 3113 | 3100 | 3116 | 3120 | 3126 | 3116 | 3136 | 3132 | 3120 | 3136 | 3150 | 3141 | 3157 | 3165 | 3160 | 3150 | 3160
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | a79dad68-1f09-ed11-85f4-5cf370a20867 | b16dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | ad9dad68-1f09-ed11-85f4-5cf370a20867 | b56dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | b39dad68-1f09-ed11-85f4-5cf370a20867 | b96dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | b99dad68-1f09-ed11-85f4-5cf370a20867 | bd6dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | bf9dad68-1f09-ed11-85f4-5cf370a20867 | c16dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | c59dad68-1f09-ed11-85f4-5cf370a20867 | c66dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | cb9dad68-1f09-ed11-85f4-5cf370a20867 | 010bca23-2809-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.17 | [21 empty] | 010081427802027710118705 2105883 17240328 | 11 | 4 | V1169 | 2022-07-21 13:03 | 870 | 289 | 20.3816 | 2650.1897 | 5 | 27 | 2808 | 2838 | 2836 | 2830 | 2848 | 2850 | 2853 | 2859 | 2855 | 2855 | 2857 | 2864 | 2847 | 2864 | 2860 | 2870 | 2904 | 2930 | 2978 | 3024 | 3101 | 3139 | 3190 | 3199 | 3210 | 3244 | 3250 | 3273 | 3267 | 3268 | 3288 | 3287 | 3283 | 3295 | 3293 | 3304 | 3313 | 3315 | 3314 | 3322 | 3319 | 3341 | 3335 | 3349 | 3348
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | d29dad68-1f09-ed11-85f4-5cf370a20867 | a30aca23-2809-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.18 | [21 empty] | 010081427802027710118705 2105883 17240328 | 12 | 4 | V1169 | 2022-07-21 13:03 | 839 | 276 | 21.0879 | 2618.4016 | 5 | 27 | 2824 | 2818 | 2808 | 2800 | 2809 | 2821 | 2829 | 2815 | 2831 | 2832 | 2825 | 2832 | 2841 | 2844 | 2845 | 2847 | 2892 | 2926 | 2978 | 3031 | 3077 | 3120 | 3155 | 3162 | 3205 | 3228 | 3233 | 3246 | 3261 | 3248 | 3261 | 3279 | 3261 | 3277 | 3299 | 3293 | 3295 | 3292 | 3294 | 3304 | 3327 | 3333 | 3317 | 3332 | 3335
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 1f42d920-2009-ed11-85f4-5cf370a20867 | 13e5bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.45 | 54.87 | 1.25 | 703 | [20 empty] | 010081427802027710118705 2105883 17240328 | 5 | 4 | V1169 | 2022-07-21 13:09 | 716 | 286 | 0.1 | 2777.1 | 2 | 4 | 2777 | 2775 | 2776 | 2784 | 2770 | 2785 | 2797 | 2806 | 2788 | 2816 | 2796 | 2802 | 2805 | 2806 | 2821 | 2838 | 2859 | 2926 | 3001 | 3093 | 3161 | 3218 | 3261 | 3288 | 3305 | 3334 | 3341 | 3364 | 3377 | 3379 | 3391 | 3396 | 3395 | 3421 | 3421 | 3442 | 3447 | 3462 | 3464 | 3450 | 3466 | 3485 | 3473 | 3479 | 3490
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 2642d920-2009-ed11-85f4-5cf370a20867 | e1e4bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.28 | 52.81 | 1.23 | 438 | [20 empty] | 010081427802027710118705 2105883 17240328 | 6 | 4 | V1169 | 2022-07-21 13:09 | 732 | 276 | 5 | 2706.5333 | 2 | 4 | 2710 | 2715 | 2723 | 2724 | 2737 | 2733 | 2730 | 2740 | 2741 | 2728 | 2746 | 2741 | 2759 | 2764 | 2771 | 2783 | 2826 | 2888 | 2984 | 3043 | 3124 | 3167 | 3181 | 3211 | 3240 | 3254 | 3274 | 3278 | 3292 | 3303 | 3308 | 3323 | 3323 | 3322 | 3334 | 3333 | 3359 | 3339 | 3357 | 3352 | 3353 | 3340 | 3378 | 3385 | 3361
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 2c42d920-2009-ed11-85f4-5cf370a20867 | 8fe6bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.18 | [21 empty] | 010081427802027710118705 2105883 17240328 | 7 | 4 | V1169 | 2022-07-21 13:09 | 837 | 291 | 20.8789 | 2564.3731 | 5 | 27 | 2744 | 2759 | 2745 | 2764 | 2759 | 2770 | 2770 | 2771 | 2754 | 2772 | 2773 | 2780 | 2776 | 2783 | 2785 | 2814 | 2820 | 2853 | 2918 | 2968 | 3037 | 3061 | 3095 | 3131 | 3142 | 3158 | 3159 | 3194 | 3178 | 3182 | 3191 | 3192 | 3232 | 3209 | 3225 | 3224 | 3226 | 3232 | 3248 | 3242 | 3252 | 3242 | 3267 | 3250 | 3254
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 3242d920-2009-ed11-85f4-5cf370a20867 | 30e6bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.19 | [21 empty] | 010081427802027710118705 2105883 17240328 | 8 | 4 | V1169 | 2022-07-21 13:09 | 768 | 291 | 21.1269 | 2541.8308 | 5 | 27 | 2733 | 2728 | 2743 | 2723 | 2732 | 2749 | 2745 | 2746 | 2738 | 2747 | 2741 | 2753 | 2766 | 2769 | 2781 | 2779 | 2814 | 2863 | 2900 | 2960 | 3021 | 3065 | 3079 | 3089 | 3107 | 3137 | 3148 | 3162 | 3172 | 3162 | 3188 | 3195 | 3183 | 3199 | 3212 | 3226 | 3233 | 3216 | 3228 | 3225 | 3227 | 3239 | 3249 | 3256 | 3253
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 3842d920-2009-ed11-85f4-5cf370a20867 | 00e6bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.78 | 57.59 | 1.25 | 502 | [20 empty] | 010081427802027710118705 2105883 17240328 | 9 | 4 | V1169 | 2022-07-21 13:09 | 775 | 295 | 4.4 | 2755.0667 | 2 | 4 | 2763 | 2754 | 2765 | 2775 | 2783 | 2787 | 2784 | 2774 | 2771 | 2767 | 2787 | 2775 | 2796 | 2790 | 2803 | 2806 | 2817 | 2879 | 2956 | 3047 | 3115 | 3170 | 3207 | 3260 | 3271 | 3294 | 3307 | 3334 | 3337 | 3339 | 3365 | 3365 | 3376 | 3386 | 3402 | 3388 | 3398 | 3418 | 3417 | 3443 | 3439 | 3430 | 3430 | 3452 | 3464
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 3e42d920-2009-ed11-85f4-5cf370a20867 | 73e5bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.61 | 57.38 | 1.23 | 706 | [20 empty] | 010081427802027710118705 2105883 17240328 | 10 | 4 | V1169 | 2022-07-21 13:09 | 849 | 286 | -0.7 | 2830.2333 | 2 | 4 | 2843 | 2826 | 2819 | 2815 | 2834 | 2845 | 2848 | 2849 | 2835 | 2843 | 2840 | 2844 | 2848 | 2848 | 2865 | 2870 | 2907 | 2945 | 3039 | 3126 | 3204 | 3242 | 3296 | 3328 | 3325 | 3365 | 3387 | 3386 | 3411 | 3412 | 3423 | 3441 | 3443 | 3449 | 3441 | 3473 | 3468 | 3467 | 3472 | 3488 | 3490 | 3505 | 3495 | 3501 | 3509
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 4442d920-2009-ed11-85f4-5cf370a20867 | 803d7504-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:09 |  |  |  |  | 0
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | b7477ca2-3609-ed11-85f4-5cf370a20867 | f2c2a3da-3609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:50 |  |  |  |  | 0
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | f5c2a3da-3609-ed11-85f4-5cf370a20867 | 8db7eb13-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:51 |  |  |  |  | 0
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 90b7eb13-3709-ed11-85f4-5cf370a20867 | b7bbce4d-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:53 |  |  |  |  | 0
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | f2bbce4d-3709-ed11-85f4-5cf370a20867 | 0b592f87-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:54 |  |  |  |  | 0
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 0e592f87-3709-ed11-85f4-5cf370a20867 | 4e3e36c2-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Aborted | TargetAborted | 7162 (Warning, Quantity not sufficient), 1000 (Error, User Aborted Test) |  |  | 0 | [25 empty] | 2022-07-21 15:56 |  |  |  |  | 0
Normalized
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 0d5c9e96-1209-ed11-85f4-5cf370a20867 | 5fc854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105880 17240328 | 8 | 3 | V790 | 2022-07-21 11:32 | 501 | 259 | 18.9314 | 2168.8451 | 5 | 27 | 159.22 | 137.49 | 117.76 | 98.43 | 79.7 | 63.97 | 47.63 | 30.5 | 11.37 | -3.96 | -19.09 | -33.02 | -49.55 | -62.08 | -70.82 | -75.75 | -69.48 | -55.81 | -30.54 | -7.07 | 19 | 34.86 | 43.13 | 43.8 | 41.47 | 34.14 | 27.61 | 24.48 | 12.94 | 3.81 | -9.32 | -22.65 | -38.98 | -49.51 | -63.84 | -75.18 | -89.91 | -104.04 | -118.37 | -134.9 | -152.43 | -167.76 | -180.9 | -197.63 | -212.36
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 145c9e96-1209-ed11-85f4-5cf370a20867 | 01c854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.27 | [21 empty] | 010081427802027710118705 2105880 17240328 | 9 | 3 | V790 | 2022-07-21 11:32 | 555 | 259 | 27.7921 | 2118.1265 | 5 | 27 | 213.88 | 190.49 | 168.9 | 150.11 | 125.51 | 100.52 | 73.73 | 48.94 | 20.94 | -4.45 | -29.64 | -53.23 | -79.82 | -100.62 | -114.21 | -117.6 | -106.39 | -80.38 | -49.58 | -10.37 | 23.64 | 47.85 | 63.46 | 72.26 | 65.47 | 58.28 | 45.69 | 30.09 | 11.7 | -4.09 | -22.88 | -40.67 | -65.27 | -87.46 | -110.45 | -135.04 | -160.43 | -177.23 | -200.62 | -224.41 | -243.8 | -264.79 | -293.19 | -317.98 | -343.97
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 1b5c9e96-1209-ed11-85f4-5cf370a20867 | 4aca54cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105880 17240328 | 10 | 3 | V790 | 2022-07-21 11:32 | 524 | 267 | 20.2032 | 2087.4016 | 5 | 27 | 159 | 138.39 | 118.99 | 100.59 | 79.38 | 65.58 | 51.58 | 32.97 | 15.57 | -3.23 | -23.44 | -41.24 | -55.04 | -68.45 | -72.65 | -75.85 | -68.86 | -49.26 | -27.86 | -3.66 | 20.93 | 37.13 | 41.73 | 45.72 | 42.52 | 33.32 | 23.11 | 13.71 | 1.91 | -14.7 | -22.1 | -31.5 | -48.91 | -62.51 | -76.11 | -97.12 | -120.52 | -136.52 | -155.12 | -172.93 | -187.13 | -202.73 | -221.74 | -243.34 | -263.74
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 225c9e96-1209-ed11-85f4-5cf370a20867 | e9c954cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105880 17240328 | 11 | 3 | V790 | 2022-07-21 11:32 | 509 | 246 | 19.9302 | 2120.551 | 5 | 27 | 150.32 | 137.39 | 119.26 | 104.13 | 80 | 65.07 | 47.14 | 30.81 | 15.28 | 1.75 | -17.58 | -36.91 | -53.44 | -68.37 | -77.1 | -79.83 | -69.57 | -48.9 | -27.83 | -3.76 | 19.31 | 32.38 | 41.05 | 45.12 | 44.19 | 35.26 | 25.93 | 16 | 3.47 | -12.86 | -24.99 | -37.12 | -55.85 | -70.78 | -84.71 | -95.44 | -113.57 | -126.7 | -144.63 | -162.36 | -180.89 | -197.22 | -213.35 | -227.08 | -241.01
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 295c9e96-1209-ed11-85f4-5cf370a20867 | 87c954cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.52 | 55.01 | 1.28 | 468 | [20 empty] | 010081427802027710118705 2105880 17240328 | 12 | 3 | V790 | 2022-07-21 11:32 | 499 | 265 | 4.9 | 2369.2333 | 2 | 4 | 1.27 | 0.17 | -0.33 | 0.17 | 2.07 | -0.43 | -3.13 | -7.83 | -12.33 | -17.63 | -18.93 | -18.23 | -16.93 | -14.43 | -5.13 | 15.17 | 48.87 | 99.97 | 161.47 | 225.17 | 284.87 | 333.57 | 370.47 | 399.57 | 420.07 | 433.17 | 443.87 | 455.97 | 458.67 | 467.37 | 472.67 | 474.57 | 469.67 | 475.17 | 470.47 | 470.77 | 472.67 | 475.97 | 473.67 | 475.57 | 475.07 | 471.97 | 471.47 | 470.57 | 468.47
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 305c9e96-1209-ed11-85f4-5cf370a20867 | e454c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105879 17240328 | 1 | 4 | V1169 | 2022-07-21 11:32 | 785 | 294 | 23.1577 | 2506.8071 | 5 | 27 | 186.44 | 163.48 | 141.32 | 121.16 | 101.6 | 82.65 | 60.69 | 40.73 | 19.37 | -3.58 | -20.34 | -42.3 | -65.26 | -84.02 | -92.77 | -101.13 | -91.49 | -68.45 | -41 | -10.76 | 20.68 | 43.32 | 52.57 | 56.01 | 55.85 | 48.69 | 38.93 | 27.38 | 14.02 | -0.94 | -14.5 | -26.85 | -42.81 | -59.97 | -80.33 | -99.88 | -122.44 | -140.6 | -158.96 | -175.72 | -199.07 | -215.43 | -234.59 | -253.95 | -275.9
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 375c9e96-1209-ed11-85f4-5cf370a20867 | 8554c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105879 17240328 | 2 | 4 | V1169 | 2022-07-21 11:32 | 729 | 292 | 23.2439 | 2407.3328 | 5 | 27 | 182.62 | 160.58 | 139.74 | 118.09 | 98.65 | 81.8 | 58.36 | 38.72 | 18.87 | -2.77 | -23.22 | -38.66 | -59.5 | -74.95 | -88.59 | -92.03 | -88.48 | -71.32 | -47.97 | -17.21 | 10.75 | 33.9 | 50.66 | 58.61 | 57.37 | 54.33 | 42.68 | 31.84 | 17.79 | 7.55 | -6.09 | -17.94 | -32.18 | -46.02 | -64.87 | -87.91 | -106.16 | -129.4 | -150.04 | -170.29 | -188.13 | -206.58 | -226.82 | -248.66 | -271.91
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 3e5c9e96-1209-ed11-85f4-5cf370a20867 | 4455c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105879 17240328 | 3 | 4 | V1169 | 2022-07-21 11:32 | 739 | 274 | 22.2113 | 2319.6806 | 5 | 27 | 176.51 | 155.5 | 136.09 | 113.07 | 95.26 | 74.05 | 54.04 | 35.43 | 18.02 | 0.01 | -19.4 | -39.02 | -63.83 | -79.24 | -91.45 | -94.26 | -85.47 | -57.28 | -31.29 | -0.91 | 23.88 | 43.07 | 50.66 | 52.45 | 46.04 | 40.63 | 28.62 | 16.2 | 3.99 | -8.62 | -24.63 | -38.44 | -57.05 | -71.46 | -87.47 | -104.89 | -123.1 | -141.91 | -160.12 | -180.33 | -199.74 | -218.75 | -238.57 | -262.38 | -283.79
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 6042e1f8-1209-ed11-85f4-5cf370a20867 | f53c8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.52 | 53.44 | 1.22 | 471 | [20 empty] | 010081427802027710118705 2105881 17240328 | 12 | 2 | V834 | 2022-07-21 11:34 | 614 | 257 | 0.7 | 2263.5 | 2 | 4 | -1.8 | 0.1 | -0.2 | 0.1 | -1.4 | -1.7 | -5.2 | -2.3 | -4.4 | -4.7 | -3 | 1.3 | -1.8 | 4.7 | 9.6 | 21.7 | 50.6 | 97.3 | 148.8 | 205.5 | 260.2 | 302.3 | 339.2 | 362.1 | 382.6 | 397.5 | 411.2 | 419.7 | 429.8 | 437.3 | 444 | 449.5 | 449.8 | 453.3 | 457.6 | 458.1 | 458.2 | 459.9 | 462.4 | 464.9 | 467.4 | 470.9 | 473 | 473.5 | 471.2
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 6642e1f8-1209-ed11-85f4-5cf370a20867 | 508c0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105880 17240328 | 1 | 3 | V790 | 2022-07-21 11:34 | 494 | 260 | 20.1514 | 2002.9866 | 5 | 27 | 160.66 | 141.91 | 125.56 | 106.01 | 87.46 | 73.51 | 54.95 | 36 | 20.45 | -1.1 | -19.65 | -36.8 | -56.55 | -76.31 | -85.06 | -88.81 | -83.56 | -62.71 | -32.66 | -6.61 | 15.43 | 37.08 | 46.53 | 48.98 | 48.43 | 47.28 | 33.73 | 20.57 | 6.02 | -8.73 | -23.48 | -32.83 | -45.58 | -58.33 | -71.08 | -88.64 | -105.39 | -124.34 | -140.69 | -156.44 | -172.39 | -191.54 | -206.7 | -225.45 | -245.2
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 6c42e1f8-1209-ed11-85f4-5cf370a20867 | 1d8c0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.22 | [21 empty] | 010081427802027710118705 2105880 17240328 | 2 | 3 | V790 | 2022-07-21 11:34 | 513 | 269 | 19.7261 | 2017.5217 | 5 | 27 | 167.35 | 151.23 | 131.9 | 110.97 | 94.25 | 78.12 | 54 | 36.27 | 18.14 | -0.38 | -23.31 | -42.03 | -64.36 | -80.49 | -92.41 | -91.94 | -80.27 | -56.59 | -26.72 | -0.64 | 22.43 | 38.5 | 46.58 | 49.45 | 46.73 | 43 | 31.67 | 23.95 | 12.42 | 5.5 | -5.23 | -11.56 | -26.28 | -40.61 | -54.53 | -71.86 | -90.39 | -107.71 | -130.44 | -149.37 | -163.49 | -181.02 | -194.54 | -204.47 | -220
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 7242e1f8-1209-ed11-85f4-5cf370a20867 | 918b0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold), 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed) | 16.37 | 60.68 | 1.34 | 745 | [20 empty] | 010081427802027710118705 2105880 17240328 | 3 | 3 | V790 | 2022-07-21 11:34 | 535 | 271 | 0.5 | 2230.2333 | 2 | 4 | 3.07 | 0.17 | -0.33 | 0.17 | 1.47 | 3.37 | 3.87 | 7.57 | 7.07 | 10.17 | 11.47 | 17.17 | 20.07 | 28.97 | 42.47 | 70.97 | 116.07 | 176.77 | 250.27 | 330.37 | 405.67 | 466.97 | 515.47 | 555.57 | 583.47 | 605.97 | 626.27 | 643.57 | 657.27 | 673.97 | 684.07 | 694.37 | 702.67 | 708.17 | 708.47 | 717.97 | 725.47 | 731.57 | 739.87 | 747.97 | 748.87 | 748.57 | 749.27 | 746.97 | 745.47
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 7842e1f8-1209-ed11-85f4-5cf370a20867 | 328b0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.27 | 56.35 | 1.32 | 874 | [20 empty] | 010081427802027710118705 2105880 17240328 | 4 | 3 | V790 | 2022-07-21 11:34 | 513 | 267 | -4.1 | 2200.7667 | 2 | 4 | -1.87 | -0.37 | 0.73 | -0.37 | 6.53 | 13.63 | 16.53 | 18.83 | 28.73 | 32.83 | 37.33 | 45.83 | 57.33 | 66.43 | 80.53 | 110.43 | 155.93 | 213.23 | 285.13 | 362.43 | 436.73 | 496.43 | 546.93 | 584.63 | 616.13 | 643.63 | 669.33 | 691.63 | 709.33 | 729.83 | 745.33 | 759.83 | 773.73 | 788.23 | 795.53 | 802.43 | 811.13 | 818.23 | 825.53 | 839.23 | 850.93 | 861.83 | 868.13 | 873.43 | 874.33
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 7e42e1f8-1209-ed11-85f4-5cf370a20867 | bdaadb1d-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 2610 (Informational, Delays in sample processing), 1020 (Informational, Peak Not Detected) |  |  | 1.18 | [21 empty] | 010081427802027710118705 2105880 17240328 | 5 | 3 | V790 | 2022-07-21 11:34 | 499 | 266 | 12.6579 | 1361.7431 | 5 | 27 | 106.8 | 95.34 | 84.68 | 71.83 | 63.97 | 52.51 | 35.25 | 18.79 | 6.74 | -8.72 | -21.38 | -29.64 | -41.1 | -48.75 | -55.01 | -54.87 | -47.53 | -31.79 | -15.44 | 3.5 | 20.64 | 28.78 | 35.73 | 30.87 | 26.01 | 19.55 | 11.89 | 4.24 | -0.42 | -8.88 | -23.54 | -34.8 | -47.85 | -58.91 | -68.17 | -76.83 | -82.69 | -95.14 | -115 | -132.06 | -147.92 | -164.98 | -174.63 | -176.69 | -183.15
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 8442e1f8-1209-ed11-85f4-5cf370a20867 | 29c954cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.18 | 50.9 | 1.36 | 699 | [20 empty] | 010081427802027710118705 2105880 17240328 | 6 | 3 | V790 | 2022-07-21 11:34 | 498 | 265 | 2 | 2167.2667 | 2 | 4 | -0.07 | -0.67 | 1.33 | -0.67 | 1.53 | -0.47 | 0.93 | -6.07 | -7.87 | -9.07 | -6.47 | -6.47 | -1.47 | 14.93 | 33.93 | 64.53 | 107.93 | 170.53 | 233.13 | 301.33 | 363.53 | 417.33 | 461.33 | 493.53 | 514.53 | 528.73 | 543.73 | 554.13 | 564.53 | 574.53 | 590.13 | 602.93 | 612.33 | 627.93 | 640.33 | 653.13 | 665.13 | 676.93 | 682.13 | 692.53 | 693.93 | 696.13 | 698.13 | 699.73 | 699.13
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 8a42e1f8-1209-ed11-85f4-5cf370a20867 | bfc854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.18 | [21 empty] | 010081427802027710118705 2105880 17240328 | 7 | 3 | V790 | 2022-07-21 11:34 | 494 | 262 | 16.7034 | 1997.2158 | 5 | 27 | 117.68 | 103.38 | 89.07 | 79.57 | 69.67 | 58.36 | 43.86 | 31.16 | 11.65 | -6.45 | -21.15 | -37.06 | -51.16 | -59.46 | -63.77 | -67.07 | -56.77 | -39.88 | -18.18 | 3.12 | 21.21 | 31.51 | 35.41 | 32.1 | 29 | 28.3 | 25.59 | 21.69 | 16.19 | 5.88 | -5.02 | -18.72 | -31.83 | -44.13 | -54.63 | -65.54 | -75.04 | -87.54 | -101.65 | -117.55 | -131.85 | -146.56 | -165.26 | -182.96 | -200.47
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 5a7b815a-1309-ed11-85f4-5cf370a20867 | 503c8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.17 | [21 empty] | 010081427802027710118705 2105881 17240328 | 7 | 2 | V834 | 2022-07-21 11:37 | 563 | 270 | 16.2757 | 2127.102 | 5 | 27 | 138.82 | 119.75 | 101.67 | 88 | 73.52 | 59.64 | 47.97 | 32.09 | 14.42 | 0.54 | -17.93 | -36.81 | -51.49 | -65.76 | -71.84 | -72.91 | -63.99 | -48.26 | -24.94 | -4.62 | 18.11 | 30.83 | 43.96 | 42.88 | 39.41 | 31.13 | 24.05 | 11.58 | 3.5 | -4.17 | -13.25 | -24.92 | -38 | -50.28 | -65.15 | -79.83 | -91.3 | -105.38 | -119.05 | -133.73 | -146.41 | -159.28 | -170.96 | -184.63 | -196.11
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 607b815a-1309-ed11-85f4-5cf370a20867 | 22597999-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.16 | [21 empty] | 010081427802027710118705 2105881 17240328 | 8 | 2 | V834 | 2022-07-21 11:37 | 569 | 268 | 14.8879 | 2018.9233 | 5 | 27 | 126.19 | 113.7 | 99.41 | 87.32 | 73.44 | 55.95 | 41.26 | 25.77 | 9.09 | -4.2 | -16.49 | -29.18 | -41.27 | -51.35 | -66.84 | -67.13 | -62.02 | -47.51 | -27.59 | -3.88 | 14.43 | 27.74 | 35.85 | 35.97 | 38.48 | 31.99 | 27.5 | 18.81 | 13.13 | 4.44 | -6.65 | -22.34 | -32.63 | -46.91 | -62.2 | -74.89 | -85.78 | -97.47 | -106.15 | -119.24 | -129.93 | -142.82 | -156.9 | -173.39 | -186.08
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 667b815a-1309-ed11-85f4-5cf370a20867 | 6e3e8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.15 | [21 empty] | 010081427802027710118705 2105881 17240328 | 9 | 2 | V834 | 2022-07-21 11:37 | 664 | 265 | 14.7126 | 2116.7802 | 5 | 27 | 124.71 | 111.99 | 97.48 | 84.37 | 69.46 | 56.94 | 38.83 | 25.92 | 9.81 | -5.91 | -22.82 | -32.13 | -43.24 | -53.56 | -59.47 | -59.38 | -55.1 | -42.81 | -22.92 | -3.43 | 14.25 | 29.34 | 37.63 | 37.32 | 34.2 | 27.09 | 19.98 | 11.07 | 6.15 | -2.56 | -11.27 | -21.18 | -35.7 | -50.21 | -63.92 | -75.44 | -89.55 | -100.06 | -112.57 | -124.89 | -134.4 | -144.91 | -157.62 | -170.34 | -183.05
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 6c7b815a-1309-ed11-85f4-5cf370a20867 | 103e8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105881 17240328 | 10 | 2 | V834 | 2022-07-21 11:37 | 659 | 280 | 21.4891 | 2038.9826 | 5 | 27 | 190.73 | 167.04 | 147.55 | 125.46 | 99.37 | 80.08 | 60.79 | 38.5 | 19.42 | -0.27 | -23.56 | -44.45 | -64.34 | -82.23 | -93.32 | -96.01 | -88.7 | -69.79 | -37.28 | -6.77 | 20.35 | 46.06 | 56.57 | 54.48 | 50.99 | 45.9 | 34.21 | 23.12 | 12.23 | -2.26 | -18.35 | -35.63 | -50.72 | -70.81 | -90.3 | -109.59 | -129.88 | -149.37 | -170.66 | -188.95 | -204.04 | -223.53 | -242.62 | -258.7 | -276.99
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 727b815a-1309-ed11-85f4-5cf370a20867 | b23d8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.22 | [21 empty] | 010081427802027710118705 2105881 17240328 | 11 | 2 | V834 | 2022-07-21 11:37 | 627 | 270 | 20.9273 | 2048.0593 | 5 | 27 | 192.61 | 167.29 | 143.76 | 122.83 | 101.3 | 77.58 | 59.25 | 41.12 | 18.2 | 0.47 | -21.06 | -39.19 | -61.11 | -79.64 | -96.37 | -96.5 | -92.62 | -71.75 | -40.88 | -8.4 | 18.87 | 38.74 | 51.01 | 54.89 | 53.56 | 49.03 | 43.5 | 28.38 | 10.25 | -1.68 | -18.4 | -33.73 | -44.26 | -56.59 | -76.11 | -96.44 | -114.17 | -136.7 | -156.02 | -178.55 | -193.68 | -215.6 | -227.53 | -240.66 | -249.99
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 787b815a-1309-ed11-85f4-5cf370a20867 | f0fe3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.24 | [21 empty] | 010081427802027710118705 2105881 17240328 | 4 | 2 | V834 | 2022-07-21 11:37 | 633 | 295 | 24.0296 | 2104.43 | 5 | 27 | 190.34 | 166.51 | 145.28 | 124.85 | 101.22 | 79.19 | 61.76 | 41.53 | 19.3 | -1.93 | -24.76 | -45.79 | -68.82 | -83.25 | -94.67 | -97.9 | -89.73 | -67.16 | -37.99 | -2.62 | 27.95 | 44.12 | 57.09 | 59.46 | 50.83 | 40 | 32.17 | 17.94 | 4.71 | -6.92 | -21.75 | -39.98 | -57.81 | -81.44 | -103.47 | -127.9 | -145.73 | -166.16 | -186.19 | -206.02 | -225.05 | -246.08 | -266.5 | -285.93 | -308.36
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 7e7b815a-1309-ed11-85f4-5cf370a20867 | 543d8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.27 | [21 empty] | 010081427802027710118705 2105881 17240328 | 5 | 2 | V834 | 2022-07-21 11:37 | 596 | 273 | 27.6573 | 2095.9352 | 5 | 27 | 212.41 | 185.95 | 161.69 | 140.24 | 119.58 | 96.72 | 73.86 | 50.81 | 24.55 | 0.69 | -28.77 | -51.02 | -76.48 | -98.94 | -116.99 | -116.85 | -109.51 | -86.57 | -50.62 | -11.28 | 25.46 | 49.2 | 66.95 | 72.29 | 66.63 | 55.17 | 45.12 | 25.86 | 1.2 | -17.85 | -40.51 | -60.77 | -79.43 | -96.68 | -120.74 | -144 | -171.66 | -198.51 | -220.97 | -246.23 | -269.48 | -292.34 | -310.6 | -332.46 | -350.11
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 847b815a-1309-ed11-85f4-5cf370a20867 | 953c8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.23 | 58.23 | 1.25 | 315 | [20 empty] | 010081427802027710118705 2105881 17240328 | 6 | 2 | V834 | 2022-07-21 11:37 | 617 | 276 | 6.4 | 2303.3333 | 2 | 4 | -0.53 | 0.47 | -0.93 | 0.47 | -7.13 | -11.33 | -19.33 | -21.13 | -29.13 | -32.93 | -35.33 | -39.13 | -46.33 | -43.93 | -40.93 | -28.93 | 5.27 | 53.27 | 106.87 | 165.67 | 221.87 | 259.67 | 293.07 | 312.67 | 327.47 | 334.47 | 344.67 | 349.27 | 353.67 | 359.67 | 363.07 | 364.07 | 356.27 | 355.67 | 350.47 | 343.67 | 337.47 | 336.87 | 333.07 | 329.87 | 331.27 | 329.47 | 326.27 | 321.27 | 314.67
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | b3de19db-1309-ed11-85f4-5cf370a20867 | 3a5f6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.74 | 58.9 | 1.27 | 541 | [20 empty] | 010081427802027710118705 2105881 17240328 | 3 | 2 | V834 | 2022-07-21 11:41 | 592 | 269 | 2.5 | 2335.5 | 2 | 4 | 0.2 | -0.3 | 0.6 | -0.3 | 1.6 | -0.9 | -1.4 | -4.1 | -7.2 | -11.1 | -8.6 | -10.5 | -12.6 | -8.9 | 1.4 | 19.1 | 52.4 | 103.5 | 168 | 235.3 | 298.6 | 354.9 | 398.2 | 429.9 | 453.4 | 473.5 | 486.4 | 499.1 | 507.8 | 515.3 | 521.8 | 526.1 | 531.8 | 535.1 | 537.4 | 542.3 | 545.4 | 544.7 | 549.4 | 553.5 | 552 | 552.1 | 551.2 | 545.9 | 540.8
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | b9de19db-1309-ed11-85f4-5cf370a20867 | a75e6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.19 | [21 empty] | 010081427802027710118705 2105882 17240328 | 10 | 1 | V639 | 2022-07-21 11:41 | 614 | 260 | 19.8022 | 2195.6174 | 5 | 27 | 152.78 | 136.58 | 120.98 | 103.17 | 84.17 | 67.77 | 51.37 | 35.77 | 21.56 | 2.56 | -19.24 | -37.24 | -58.65 | -73.65 | -83.45 | -84.25 | -76.65 | -55.46 | -33.06 | -6.06 | 18.14 | 34.93 | 43.33 | 46.53 | 47.33 | 41.13 | 33.12 | 23.52 | 15.52 | 1.92 | -9.08 | -23.49 | -38.89 | -55.09 | -69.49 | -86.9 | -101.7 | -114.3 | -131.3 | -149.9 | -169.11 | -183.71 | -204.31 | -222.91 | -243.32
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | bfde19db-1309-ed11-85f4-5cf370a20867 | 185e6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.26 | 56.9 | 1.29 | 510 | [20 empty] | 010081427802027710118705 2105882 17240328 | 11 | 1 | V639 | 2022-07-21 11:41 | 588 | 272 | 3.9 | 2333.5 | 2 | 4 | -2 | 0.1 | -0.2 | 0.1 | -4.6 | -7.5 | -12.6 | -12.7 | -11.4 | -12.9 | -12.8 | -12.7 | -14.8 | -14.1 | -0.6 | 22.9 | 62.2 | 118.7 | 185 | 251.7 | 311 | 361.1 | 400.4 | 424.3 | 445.2 | 459.1 | 470.8 | 478.5 | 489.4 | 493.9 | 501.6 | 508.7 | 511.2 | 516.1 | 517.8 | 521.9 | 521.2 | 520.5 | 517.8 | 516.9 | 512.4 | 510.7 | 512.4 | 511.3 | 510.4
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | c5de19db-1309-ed11-85f4-5cf370a20867 | bffe3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105882 17240328 | 12 | 1 | V639 | 2022-07-21 11:41 | 672 | 260 | 21.7893 | 2250.3186 | 5 | 27 | 175.69 | 155.9 | 137.31 | 117.32 | 95.53 | 74.35 | 55.76 | 31.77 | 13.38 | -6.81 | -21 | -39.79 | -55.38 | -72.17 | -82.36 | -86.35 | -78.34 | -62.13 | -33.52 | -6.71 | 16.11 | 33.52 | 46.93 | 47.34 | 48.55 | 44.36 | 36.97 | 23.98 | 14.79 | -1.8 | -19.19 | -34.58 | -52.37 | -72.36 | -88.75 | -104.93 | -127.72 | -140.31 | -158.3 | -172.49 | -191.28 | -205.87 | -226.66 | -246.05 | -270.04
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | cbde19db-1309-ed11-85f4-5cf370a20867 | 32fe3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.23 | [21 empty] | 010081427802027710118705 2105881 17240328 | 1 | 2 | V834 | 2022-07-21 11:41 | 679 | 291 | 23.8111 | 2216.2577 | 5 | 27 | 199.93 | 177.12 | 154.11 | 132.5 | 109.89 | 86.88 | 66.46 | 43.05 | 19.24 | -1.77 | -25.98 | -48.79 | -67.4 | -85.21 | -96.02 | -99.23 | -92.05 | -75.26 | -48.47 | -19.48 | 13.31 | 36.3 | 55.89 | 62.48 | 65.27 | 56.25 | 44.64 | 28.83 | 18.42 | -0.59 | -13.4 | -28.81 | -45.62 | -66.63 | -84.05 | -105.66 | -126.87 | -149.48 | -171.89 | -189.1 | -208.91 | -226.92 | -244.73 | -262.74 | -286.96
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | d1de19db-1309-ed11-85f4-5cf370a20867 | c65f6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.22 | [21 empty] | 010081427802027710118705 2105881 17240328 | 2 | 2 | V834 | 2022-07-21 11:41 | 600 | 276 | 24.1291 | 2174.6047 | 5 | 27 | 177.87 | 156.54 | 135.81 | 118.68 | 96.75 | 77.42 | 60.29 | 40.76 | 20.23 | 1.1 | -23.62 | -42.15 | -64.88 | -83.01 | -97.14 | -97.87 | -91.4 | -67.13 | -37.46 | -1.59 | 27.09 | 48.16 | 58.23 | 57.7 | 49.97 | 42.04 | 26.51 | 14.98 | -0.55 | -18.08 | -38.61 | -55.33 | -78.06 | -96.39 | -115.92 | -131.65 | -152.38 | -172.31 | -195.24 | -213.57 | -237.9 | -259.42 | -280.95 | -301.88 | -327.21
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | d7de19db-1309-ed11-85f4-5cf370a20867 | 674775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105882 17240328 | 8 | 1 | V639 | 2022-07-21 11:41 | 620 | 261 | 21.2957 | 2243.913 | 5 | 27 | 170.79 | 150.5 | 130.6 | 110.9 | 93.81 | 74.51 | 53.02 | 33.92 | 16.43 | -10.27 | -29.77 | -45.26 | -61.76 | -75.85 | -79.35 | -80.04 | -70.74 | -49.23 | -26.73 | -2.83 | 20.88 | 39.78 | 42.29 | 48.59 | 45.3 | 38.4 | 24.9 | 15.61 | 1.51 | -12.78 | -25.68 | -40.17 | -56.67 | -77.77 | -90.06 | -107.76 | -124.65 | -138.15 | -155.44 | -173.74 | -192.83 | -210.73 | -231.03 | -247.12 | -267.62
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | ddde19db-1309-ed11-85f4-5cf370a20867 | 095f6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105882 17240328 | 9 | 1 | V639 | 2022-07-21 11:41 | 661 | 258 | 21.6571 | 2230.5644 | 5 | 27 | 157.18 | 145.32 | 131.26 | 116.01 | 94.95 | 74.69 | 48.44 | 26.38 | 7.12 | -8.94 | -26.99 | -42.05 | -58.31 | -70.76 | -82.22 | -83.88 | -69.94 | -46.19 | -19.25 | 6.89 | 33.44 | 42.18 | 44.52 | 42.46 | 39.21 | 29.95 | 18.29 | 5.44 | -10.62 | -27.48 | -45.13 | -59.39 | -79.85 | -94.11 | -114.36 | -133.42 | -149.68 | -165.33 | -185.99 | -204.25 | -220.51 | -238.96 | -256.62 | -276.48 | -294.93
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 65fdae7a-1409-ed11-85f4-5cf370a20867 | 084775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105882 17240328 | 4 | 1 | V639 | 2022-07-21 11:45 | 684 | 263 | 21.7674 | 2166.2783 | 5 | 27 | 174.55 | 155.99 | 138.02 | 121.25 | 100.88 | 80.72 | 57.75 | 35.38 | 12.22 | -8.35 | -24.12 | -43.09 | -61.25 | -77.42 | -88.19 | -92.76 | -84.52 | -60.69 | -33.46 | -4.03 | 22.61 | 42.24 | 49.27 | 51.1 | 51.14 | 42.97 | 31.6 | 21.43 | 8.87 | -8.7 | -22.27 | -34.23 | -54.8 | -69.77 | -86.14 | -102.7 | -123.27 | -139.64 | -160.81 | -178.77 | -204.74 | -224.31 | -245.28 | -265.24 | -287.61
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 6bfdae7a-1409-ed11-85f4-5cf370a20867 | d74675b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105882 17240328 | 5 | 1 | V639 | 2022-07-21 11:45 | 703 | 280 | 23.3328 | 2329.7447 | 5 | 27 | 181.52 | 157.79 | 135.86 | 115.12 | 97.79 | 77.86 | 56.13 | 34.39 | 9.66 | -11.47 | -30.21 | -47.54 | -64.07 | -77.2 | -84.54 | -82.47 | -72.2 | -50.94 | -22.27 | 0.8 | 26.27 | 42.53 | 48.8 | 46.47 | 45.54 | 34.4 | 22.27 | 7.54 | -3.6 | -15.73 | -32.26 | -44.39 | -56.53 | -76.06 | -96.59 | -116.73 | -139.06 | -162.79 | -181.92 | -206.66 | -223.99 | -242.12 | -261.66 | -282.39 | -300.32
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 71fdae7a-1409-ed11-85f4-5cf370a20867 | 494675b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 15.89 | 59.99 | 1.29 | 667 | [20 empty] | 010081427802027710118705 2105882 17240328 | 6 | 1 | V639 | 2022-07-21 11:45 | 749 | 264 | 1.2 | 2464.1333 | 2 | 4 | -1.73 | -0.13 | 0.27 | -0.13 | 4.67 | 5.47 | 4.67 | 8.67 | 7.67 | 6.87 | 7.07 | 5.67 | 4.67 | 12.07 | 28.07 | 57.87 | 107.87 | 168.87 | 240.47 | 313.47 | 384.07 | 445.67 | 491.27 | 526.27 | 553.27 | 569.47 | 580.27 | 590.47 | 603.27 | 608.27 | 619.27 | 625.07 | 633.47 | 634.87 | 638.87 | 641.27 | 645.07 | 645.27 | 651.47 | 657.47 | 659.47 | 663.27 | 667.47 | 667.47 | 667.27
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 77fdae7a-1409-ed11-85f4-5cf370a20867 | f44775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105882 17240328 | 7 | 1 | V639 | 2022-07-21 11:45 | 611 | 265 | 21.8399 | 2179.8047 | 5 | 27 | 172.56 | 150.92 | 131.88 | 112.64 | 95 | 74.36 | 53.72 | 31.88 | 13.64 | -8.8 | -25.04 | -42.88 | -59.12 | -75.16 | -81 | -83.24 | -72.68 | -54.52 | -28.36 | -3 | 20.36 | 38.72 | 46.48 | 47.64 | 45.2 | 39.96 | 26.92 | 14.88 | -0.76 | -17.2 | -35.84 | -53.48 | -70.32 | -86.16 | -104 | -117.04 | -136.88 | -154.72 | -172.96 | -190.2 | -212.04 | -229.88 | -250.92 | -273.96 | -298
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | f47fc1ae-1609-ed11-85f4-5cf370a20867 | 536b47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.19 | [21 empty] | 010081427802027710118705 2105882 17240328 | 1 | 1 | V639 | 2022-07-21 12:01 | 725 | 269 | 21.8492 | 2389.1866 | 5 | 27 | 165.36 | 146.72 | 126.27 | 109.62 | 90.37 | 70.12 | 51.67 | 34.42 | 13.77 | -5.48 | -21.33 | -39.58 | -57.63 | -71.88 | -83.52 | -84.97 | -71.22 | -50.47 | -26.32 | -1.37 | 21.38 | 28.93 | 38.48 | 45.83 | 46.98 | 38.73 | 33.08 | 19.04 | 1.99 | -12.26 | -24.91 | -40.56 | -57.81 | -76.06 | -92.71 | -112.16 | -133.01 | -149.26 | -160.31 | -176.15 | -197.6 | -216.25 | -236.9 | -264.15 | -288.4
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | fa7fc1ae-1609-ed11-85f4-5cf370a20867 | f46a47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105882 17240328 | 2 | 1 | V639 | 2022-07-21 12:01 | 719 | 260 | 23.9073 | 2499.5613 | 5 | 27 | 186.53 | 163.82 | 143.92 | 121.01 | 102.7 | 79.39 | 58.29 | 35.58 | 14.47 | -9.03 | -27.74 | -48.25 | -66.96 | -80.86 | -89.57 | -92.08 | -80.39 | -57.29 | -29.2 | 0.89 | 29.59 | 48.28 | 53.97 | 51.46 | 45.36 | 35.45 | 25.94 | 14.23 | 4.33 | -8.78 | -23.09 | -41.2 | -57.5 | -75.81 | -95.52 | -114.02 | -131.93 | -153.24 | -173.75 | -194.05 | -215.96 | -239.27 | -260.18 | -281.88 | -304.99
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | 0080c1ae-1609-ed11-85f4-5cf370a20867 | b36b47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16 | 50.4 | 1.21 | 402 | [20 empty] | 010081427802027710118705 2105882 17240328 | 3 | 1 | V639 | 2022-07-21 12:01 | 714 | 262 | 3.7 | 2589.5667 | 2 | 4 | -0.87 | 0.43 | -0.87 | 0.43 | 1.33 | 4.43 | 2.13 | 2.03 | 3.33 | -0.17 | -3.87 | -5.57 | -4.67 | -3.97 | 4.33 | 23.63 | 59.13 | 106.23 | 159.73 | 215.03 | 261.33 | 299.83 | 326.33 | 344.23 | 358.53 | 370.03 | 381.53 | 388.63 | 397.33 | 403.63 | 410.73 | 407.63 | 411.13 | 410.63 | 406.93 | 405.23 | 408.33 | 405.63 | 406.93 | 412.23 | 413.33 | 409.63 | 409.73 | 406.63 | 401.93
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | a79dad68-1f09-ed11-85f4-5cf370a20867 | b16dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | ad9dad68-1f09-ed11-85f4-5cf370a20867 | b56dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | b39dad68-1f09-ed11-85f4-5cf370a20867 | b96dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | b99dad68-1f09-ed11-85f4-5cf370a20867 | bd6dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | bf9dad68-1f09-ed11-85f4-5cf370a20867 | c16dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | c59dad68-1f09-ed11-85f4-5cf370a20867 | c66dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | cb9dad68-1f09-ed11-85f4-5cf370a20867 | 010bca23-2809-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.17 | [21 empty] | 010081427802027710118705 2105883 17240328 | 11 | 4 | V1169 | 2022-07-21 13:03 | 870 | 289 | 20.3816 | 2650.1897 | 5 | 27 | 149.03 | 133.05 | 120.67 | 108.68 | 91.3 | 75.52 | 60.14 | 41.16 | 22.18 | 3.99 | -18.79 | -37.37 | -56.75 | -74.53 | -86.91 | -90.7 | -88.28 | -75.86 | -50.04 | -23.42 | 8.2 | 32.01 | 48.83 | 57.05 | 58.87 | 55.09 | 48.31 | 39.52 | 27.94 | 14.96 | -3.42 | -18.2 | -33.58 | -50.76 | -65.95 | -79.93 | -96.51 | -111.09 | -128.47 | -143.25 | -159.64 | -173.02 | -188.2 | -202.78 | -221.76
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | d29dad68-1f09-ed11-85f4-5cf370a20867 | a30aca23-2809-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.18 | [21 empty] | 010081427802027710118705 2105883 17240328 | 12 | 4 | V1169 | 2022-07-21 13:03 | 839 | 276 | 21.0879 | 2618.4016 | 5 | 27 | 180.11 | 154.22 | 130.13 | 108.45 | 89.56 | 69.87 | 54.98 | 38.49 | 18.21 | -2.28 | -18.17 | -36.66 | -55.14 | -71.83 | -80.92 | -85.01 | -79.3 | -63.18 | -38.27 | -13.76 | 10.95 | 26.66 | 40.38 | 49.49 | 51 | 48.11 | 46.82 | 34.34 | 19.85 | 7.96 | -10.13 | -28.02 | -38.9 | -53.59 | -71.48 | -86.37 | -104.06 | -124.14 | -138.43 | -151.92 | -168.01 | -181.5 | -196.38 | -215.87 | -236.56
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 1f42d920-2009-ed11-85f4-5cf370a20867 | 13e5bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.45 | 54.87 | 1.25 | 703 | [20 empty] | 010081427802027710118705 2105883 17240328 | 5 | 4 | V1169 | 2022-07-21 13:09 | 716 | 286 | 0.1 | 2777.1 | 2 | 4 | -0.8 | 0.5 | -1 | 0.5 | 4.8 | 10.7 | 11.4 | 20.5 | 22.6 | 23.5 | 23.2 | 26.7 | 27.6 | 35.9 | 47.2 | 71.3 | 110.2 | 164.5 | 229 | 300.7 | 367.6 | 424.9 | 467.2 | 501.7 | 526.2 | 546.7 | 564.4 | 579.1 | 590.4 | 601.3 | 607.4 | 616.1 | 624.4 | 634.5 | 644.6 | 657.9 | 666.4 | 672.1 | 676.8 | 684.3 | 686.4 | 689.3 | 697.2 | 701.9 | 702.8
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 2642d920-2009-ed11-85f4-5cf370a20867 | e1e4bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.28 | 52.81 | 1.23 | 438 | [20 empty] | 010081427802027710118705 2105883 17240328 | 6 | 4 | V1169 | 2022-07-21 13:09 | 732 | 276 | 5 | 2706.5333 | 2 | 4 | 2.07 | -0.13 | 0.27 | -0.13 | -2.13 | -3.73 | -5.33 | -12.13 | -14.53 | -17.33 | -18.53 | -18.93 | -15.33 | -12.93 | -0.93 | 19.87 | 58.87 | 108.27 | 171.47 | 234.67 | 288.27 | 328.67 | 363.07 | 384.07 | 400.47 | 414.87 | 426.07 | 433.67 | 439.47 | 444.27 | 448.27 | 449.27 | 450.47 | 450.47 | 452.67 | 450.87 | 452.87 | 451.47 | 450.47 | 441.67 | 444.47 | 445.07 | 441.87 | 438.47 | 437.67
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 2c42d920-2009-ed11-85f4-5cf370a20867 | 8fe6bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.18 | [21 empty] | 010081427802027710118705 2105883 17240328 | 7 | 4 | V1169 | 2022-07-21 13:09 | 837 | 291 | 20.8789 | 2564.3731 | 5 | 27 | 161.95 | 145.07 | 127.19 | 111.51 | 92.83 | 77.15 | 54.27 | 36 | 15.72 | -3.16 | -23.04 | -38.12 | -56.4 | -69.08 | -81.96 | -87.43 | -81.31 | -65.59 | -41.87 | -14.55 | 12.97 | 34.69 | 48.61 | 51.93 | 50.66 | 49.58 | 38.1 | 25.22 | 10.94 | -3.34 | -16.62 | -31.3 | -43.58 | -57.85 | -71.93 | -92.81 | -105.89 | -123.37 | -138.65 | -156.33 | -170.21 | -190.68 | -209.16 | -229.64 | -248.12
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 3242d920-2009-ed11-85f4-5cf370a20867 | 30e6bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.19 | [21 empty] | 010081427802027710118705 2105883 17240328 | 8 | 4 | V1169 | 2022-07-21 13:09 | 768 | 291 | 21.1269 | 2541.8308 | 5 | 27 | 171.04 | 147.92 | 126.59 | 108.66 | 90.93 | 70.41 | 52.28 | 34.15 | 11.43 | -8.1 | -25.23 | -40.15 | -54.48 | -68.01 | -76.93 | -78.66 | -73.59 | -58.91 | -31.64 | -2.57 | 19.5 | 36.18 | 44.45 | 46.52 | 42 | 37.47 | 32.94 | 22.82 | 11.89 | 0.16 | -16.76 | -32.49 | -43.62 | -57.14 | -70.67 | -85.2 | -100.53 | -119.05 | -139.98 | -159.91 | -174.43 | -189.96 | -205.49 | -221.41 | -239.74
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 3842d920-2009-ed11-85f4-5cf370a20867 | 00e6bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.78 | 57.59 | 1.25 | 502 | [20 empty] | 010081427802027710118705 2105883 17240328 | 9 | 4 | V1169 | 2022-07-21 13:09 | 775 | 295 | 4.4 | 2755.0667 | 2 | 4 | 2.13 | 0.13 | -0.27 | 0.13 | 1.73 | -0.87 | -6.07 | -13.67 | -18.07 | -24.27 | -24.27 | -24.87 | -22.07 | -22.67 | -18.67 | -6.47 | 22.33 | 66.73 | 124.13 | 190.33 | 251.53 | 307.93 | 348.33 | 379.73 | 402.73 | 423.73 | 434.73 | 443.93 | 453.73 | 460.93 | 464.93 | 470.33 | 478.53 | 478.73 | 480.93 | 484.93 | 486.73 | 490.53 | 496.33 | 498.33 | 496.33 | 498.93 | 498.73 | 499.33 | 501.73
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 3e42d920-2009-ed11-85f4-5cf370a20867 | 73e5bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.61 | 57.38 | 1.23 | 706 | [20 empty] | 010081427802027710118705 2105883 17240328 | 10 | 4 | V1169 | 2022-07-21 13:09 | 849 | 286 | -0.7 | 2830.2333 | 2 | 4 | 5.27 | 0.37 | -0.73 | 0.37 | 5.47 | 12.17 | 16.87 | 19.37 | 19.07 | 18.97 | 19.47 | 22.77 | 27.87 | 34.57 | 47.87 | 67.97 | 106.87 | 159.77 | 227.27 | 294.97 | 365.87 | 424.37 | 464.87 | 497.77 | 527.47 | 546.17 | 563.47 | 581.57 | 593.87 | 605.37 | 617.47 | 625.77 | 632.27 | 642.97 | 649.07 | 654.57 | 659.87 | 669.97 | 674.07 | 682.17 | 688.47 | 694.97 | 699.87 | 704.37 | 705.87
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 4442d920-2009-ed11-85f4-5cf370a20867 | 803d7504-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:09 |  |  |  |  | 0
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | b7477ca2-3609-ed11-85f4-5cf370a20867 | f2c2a3da-3609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:50 |  |  |  |  | 0
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | f5c2a3da-3609-ed11-85f4-5cf370a20867 | 8db7eb13-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:51 |  |  |  |  | 0
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 90b7eb13-3709-ed11-85f4-5cf370a20867 | b7bbce4d-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:53 |  |  |  |  | 0
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | f2bbce4d-3709-ed11-85f4-5cf370a20867 | 0b592f87-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:54 |  |  |  |  | 0
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 0e592f87-3709-ed11-85f4-5cf370a20867 | 4e3e36c2-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Aborted | TargetAborted | 7162 (Warning, Quantity not sufficient), 1000 (Error, User Aborted Test) |  |  | 0 | [25 empty] | 2022-07-21 15:56 |  |  |  |  | 0
SecondDerivative
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 0d5c9e96-1209-ed11-85f4-5cf370a20867 | 5fc854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105880 17240328 | 8 | 3 | V790 | 2022-07-21 11:32 | 501 | 259 | 18.9314 | 2168.8451 | 5 | 27 | 0 | 0.85 | 2.85 | 4.65 | 4.4 | 0.95 | -2.1 | -0.1 | 4 | 4.15 | 1.35 | 2.3 | 8.45 | 15.65 | 23.55 | 31.85 | 34.4 | 26.45 | 10.8 | -7.1 | -22.65 | -29.3 | -25.8 | -18.8 | -10.75 | -3.25 | -3.15 | -7.6 | -8 | -6.65 | -5.25 | -0.85 | 2.6 | 1.5 | -1.1 | -2.5 | -2.5 | -3 | -3.8 | -1 | 3.05 | 2.15 | -1.7 | -2.15 | 0.55
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 145c9e96-1209-ed11-85f4-5cf370a20867 | 01c854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.27 | [21 empty] | 010081427802027710118705 2105880 17240328 | 9 | 3 | V790 | 2022-07-21 11:32 | 555 | 259 | 27.7921 | 2118.1265 | 5 | 27 | 0 | 0.4 | -1.5 | -6.3 | -7 | -3.35 | -1.45 | -0.6 | 1.8 | 2.95 | 1.7 | 4.75 | 15.85 | 29.7 | 42.25 | 49.9 | 46.25 | 32.75 | 13.45 | -10.15 | -28.1 | -34.7 | -36.95 | -34.1 | -24.05 | -16.1 | -12.15 | -6.7 | -2.4 | -3.3 | -7.05 | -7.75 | -4.15 | -2.3 | -1.25 | 3.95 | 5 | -0.8 | -1.05 | 1.1 | -4.6 | -8.3 | -4.05 | -0.95 | -1.05
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 1b5c9e96-1209-ed11-85f4-5cf370a20867 | 4aca54cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105880 17240328 | 10 | 3 | V790 | 2022-07-21 11:32 | 524 | 267 | 20.2032 | 2087.4016 | 5 | 27 | 0 | 0.1 | 0.9 | 4.45 | 7.2 | 2.1 | -4.4 | -4.6 | -2.85 | 0.2 | 5.95 | 10.75 | 14.65 | 18.75 | 24.15 | 31.65 | 31.85 | 20.75 | 7.45 | -7.45 | -23.25 | -28.1 | -23.3 | -20.55 | -17.15 | -9.1 | -4.9 | -5.55 | -0.7 | 4.4 | -2.05 | -7.9 | -4.65 | -6.2 | -10.7 | -4.25 | 4.45 | 4.6 | 3.7 | 3.3 | -2.35 | -7.9 | -6.4 | -1.55 | 0.9
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 225c9e96-1209-ed11-85f4-5cf370a20867 | e9c954cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105880 17240328 | 11 | 3 | V790 | 2022-07-21 11:32 | 509 | 246 | 19.9302 | 2120.551 | 5 | 27 | 0 | -2.05 | -5.55 | -3.35 | 2.95 | 4.25 | 3 | 2.6 | -1.6 | -5.8 | -2.1 | 5.9 | 12.9 | 20.85 | 31.2 | 37.55 | 31.25 | 17 | 4 | -9.5 | -20.8 | -22.7 | -20.8 | -21.3 | -18.7 | -11.1 | -6.85 | -7.35 | -4.25 | 0.2 | -2.4 | -4.8 | 0.9 | 5 | 0.6 | -3.85 | -3.85 | -4.05 | -3.5 | 0.05 | 3.35 | 4.65 | 2.65 | -0.85 | -1.15
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 295c9e96-1209-ed11-85f4-5cf370a20867 | 87c954cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.52 | 55.01 | 1.28 | 468 | [20 empty] | 010081427802027710118705 2105880 17240328 | 12 | 3 | V790 | 2022-07-21 11:32 | 499 | 265 | 4.9 | 2369.2333 | 2 | 4 | 0 | 1 | 1.85 | -1.2 | -5.65 | -6.3 | -4.3 | -1.55 | 3 | 7.4 | 7.7 | 6.8 | 12.45 | 25.9 | 41.35 | 52.8 | 53.2 | 37.55 | 11.3 | -15.15 | -33.7 | -39.65 | -36.7 | -31.65 | -23.7 | -14.1 | -10.05 | -8.15 | -4.3 | -6.55 | -11.2 | -6.6 | -1 | -1.2 | 1.85 | 4.85 | 0.4 | -3 | -2 | -2.75 | -2.75 | 0 | -0.55 | -2.95 | -2.3
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 305c9e96-1209-ed11-85f4-5cf370a20867 | e454c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105879 17240328 | 1 | 4 | V1169 | 2022-07-21 11:32 | 785 | 294 | 23.1577 | 2506.8071 | 5 | 27 | 0 | 1.35 | 3.65 | 2.95 | -0.5 | -2.1 | -1.65 | -0.9 | 1.6 | 1.9 | -1.95 | 1.55 | 14.1 | 23.85 | 33 | 44.4 | 43.3 | 27.6 | 10.95 | -6.45 | -26.15 | -35.3 | -29.65 | -22.2 | -18.6 | -14.05 | -9.25 | -6.4 | -2.95 | 0.35 | -1.1 | -5.85 | -8.1 | -6.85 | -4.2 | -0.15 | 4 | 3.3 | -1.55 | -2.05 | 1.45 | 0.3 | -4.25 | -5.4 | -3.6
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 375c9e96-1209-ed11-85f4-5cf370a20867 | 8554c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105879 17240328 | 2 | 4 | V1169 | 2022-07-21 11:32 | 729 | 292 | 23.2439 | 2407.3328 | 5 | 27 | 0 | 0.45 | 2.45 | 3.75 | 0.25 | -3 | -0.9 | 0.35 | 0.5 | 3.6 | 4.2 | 3.05 | 8.3 | 18.7 | 28.85 | 36.3 | 38 | 31.35 | 16.7 | -1.65 | -16.75 | -26.2 | -30.45 | -28.15 | -22.5 | -17 | -10.1 | -3.2 | -0.25 | -0.9 | -2.05 | -3.5 | -7.4 | -10.7 | -7.65 | -2.6 | -1.05 | 1.1 | 4.2 | 3.6 | -0.6 | -4.65 | -5.75 | -4.15 | -2.15
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 3e5c9e96-1209-ed11-85f4-5cf370a20867 | 4455c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105879 17240328 | 3 | 4 | V1169 | 2022-07-21 11:32 | 739 | 274 | 22.2113 | 2319.6806 | 5 | 27 | 0 | -0.1 | 0.65 | 1.5 | 0.75 | 1.4 | 3.5 | 2.55 | -0.8 | -3.9 | -4.7 | 1.85 | 14.4 | 25.2 | 36.1 | 46.45 | 41.95 | 22 | 2.25 | -13.05 | -25.95 | -32.25 | -29.65 | -21.65 | -14.85 | -11.3 | -6.85 | -3 | -3.35 | -4.45 | -3.95 | -2.05 | 0.1 | -1 | -3.6 | -3.45 | -1.95 | -1.7 | -1.65 | -0.45 | -0.9 | -4 | -4.5 | -1 | 1.8
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 6042e1f8-1209-ed11-85f4-5cf370a20867 | f53c8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.52 | 53.44 | 1.22 | 471 | [20 empty] | 010081427802027710118705 2105881 17240328 | 12 | 2 | V834 | 2022-07-21 11:34 | 614 | 257 | 0.7 | 2263.5 | 2 | 4 | 0 | -0.7 | -1.85 | -2.25 | -1.45 | 1.1 | 2.15 | 0.4 | 1.95 | 4.3 | 1.35 | 1.2 | 7.85 | 16.75 | 32.85 | 51 | 51.4 | 33.9 | 11.9 | -10.5 | -28.3 | -35.5 | -33.15 | -24.8 | -16.8 | -12.8 | -9.45 | -5.9 | -4.7 | -5.9 | -7.65 | -5.8 | -0.85 | -0.8 | -4.1 | -2.4 | 1.85 | 2.7 | 1.45 | 0.85 | -0.3 | -3.4 | -5.95 | -5.35 | -3.5
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 6642e1f8-1209-ed11-85f4-5cf370a20867 | 508c0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105880 17240328 | 1 | 3 | V790 | 2022-07-21 11:34 | 494 | 260 | 20.1514 | 2002.9866 | 5 | 27 | 0 | -0.75 | -0.65 | 2.35 | 2.4 | -1.6 | -2.15 | -1.7 | -2.35 | 0.1 | 1 | 1 | 10 | 23.1 | 31.4 | 39.15 | 41.85 | 26.65 | 3 | -11.15 | -19.55 | -27.45 | -25.95 | -18.25 | -17.95 | -19.9 | -13.4 | -5 | -0.25 | 4 | 4.65 | 0.3 | -3.25 | -5.45 | -6.95 | -5.15 | -0.95 | 2.45 | 1.95 | -1.25 | -1.75 | -1.1 | -3.3 | -4.3 | -2.25
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 6c42e1f8-1209-ed11-85f4-5cf370a20867 | 1d8c0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.22 | [21 empty] | 010081427802027710118705 2105880 17240328 | 2 | 3 | V790 | 2022-07-21 11:34 | 513 | 269 | 19.7261 | 2017.5217 | 5 | 27 | 0 | -0.55 | 0.75 | 2.5 | -1.7 | -4.05 | 1.25 | 2.3 | -2.75 | -3.8 | -0.25 | 4.95 | 14.05 | 26.8 | 38.55 | 43.8 | 37.55 | 19.55 | -1.25 | -15.75 | -23.75 | -26.35 | -23.4 | -18.5 | -15.1 | -11.15 | -5.1 | -0.35 | 1.3 | 0.25 | -4.35 | -8.65 | -7.15 | -4.8 | -5.5 | -5.25 | -4.7 | -2.2 | 4.55 | 7.25 | 5.55 | 6 | 1.9 | -5.9 | -7.15
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 7242e1f8-1209-ed11-85f4-5cf370a20867 | 918b0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold), 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed) | 16.37 | 60.68 | 1.34 | 745 | [20 empty] | 010081427802027710118705 2105880 17240328 | 3 | 3 | V790 | 2022-07-21 11:34 | 535 | 271 | 0.5 | 2230.2333 | 2 | 4 | 0 | 1.3 | 3.4 | 3.05 | 1.35 | 0.85 | 0.25 | -0.3 | 1.3 | 3.55 | 4.4 | 6.9 | 15.65 | 31.35 | 49.1 | 59.85 | 58.2 | 44.35 | 18.3 | -14.6 | -39.05 | -45.85 | -42.45 | -35.85 | -25.35 | -15.65 | -10.9 | -7.6 | -6.4 | -8.1 | -8.25 | -8.55 | -9.05 | -2.4 | 5.55 | 4.05 | 0.35 | -0.6 | -5.95 | -11.4 | -8.8 | -4.3 | -2.75 | -1.05 | 0.5
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 7842e1f8-1209-ed11-85f4-5cf370a20867 | 328b0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.27 | 56.35 | 1.32 | 874 | [20 empty] | 010081427802027710118705 2105880 17240328 | 4 | 3 | V790 | 2022-07-21 11:34 | 513 | 267 | -4.1 | 2200.7667 | 2 | 4 | 0 | 0.8 | 5.1 | 8.85 | 3.4 | -2.8 | 1.1 | 4.05 | 0.15 | 1.45 | 7.35 | 7.45 | 9.35 | 25.55 | 46.65 | 55.9 | 53.2 | 42.25 | 19 | -12.35 | -35.95 | -43.5 | -39.25 | -28.85 | -18.05 | -12.8 | -11.8 | -9.2 | -6.5 | -7 | -6.25 | -4.35 | -7.25 | -10.3 | -6.25 | -1.05 | 1.1 | 5.05 | 7.2 | 1.5 | -6.45 | -10.3 | -9.9 | -7.15 | -4.95
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 7e42e1f8-1209-ed11-85f4-5cf370a20867 | bdaadb1d-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 2610 (Informational, Delays in sample processing), 1020 (Informational, Peak Not Detected) |  |  | 1.18 | [21 empty] | 010081427802027710118705 2105880 17240328 | 5 | 3 | V790 | 2022-07-21 11:34 | 499 | 266 | 12.6579 | 1361.7431 | 5 | 27 | 0 | 0.35 | 1.75 | 0.45 | -6.55 | -9.15 | -1.95 | 3.25 | 3.4 | 5.5 | 6.3 | 4.45 | 6.6 | 13.3 | 21.25 | 26.1 | 22.65 | 13.25 | 2.55 | -9.25 | -18.8 | -23.05 | -21.55 | -14 | -6.55 | -2.65 | 0.45 | -1.15 | -8.05 | -9.4 | -3.35 | 1.6 | 4 | 5.55 | 4.35 | -3.2 | -13.65 | -13.9 | -3.95 | 3.4 | 9.4 | 16.7 | 12.55 | -0.25 | -5.15
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 8442e1f8-1209-ed11-85f4-5cf370a20867 | 29c954cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.18 | 50.9 | 1.36 | 699 | [20 empty] | 010081427802027710118705 2105880 17240328 | 6 | 3 | V790 | 2022-07-21 11:34 | 498 | 265 | 2 | 2167.2667 | 2 | 4 | 0 | -0.3 | -0.55 | -0.4 | -1.8 | -5.15 | -4.9 | 1.8 | 7.15 | 6.25 | 7.9 | 17.9 | 26.95 | 31.35 | 40.45 | 50.65 | 45.9 | 26.5 | 5.1 | -14.25 | -29.95 | -39.2 | -42.5 | -37.65 | -24.7 | -13 | -7.9 | -3.4 | 3.15 | 4.35 | -0.55 | -1.1 | 2.1 | 0.75 | -1.9 | -3.45 | -6.3 | -7.35 | -7.65 | -9.2 | -6.8 | -2.7 | -3.15 | -4.45 | -3.2
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 8a42e1f8-1209-ed11-85f4-5cf370a20867 | bfc854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.18 | [21 empty] | 010081427802027710118705 2105880 17240328 | 7 | 3 | V790 | 2022-07-21 11:34 | 494 | 262 | 16.7034 | 1997.2158 | 5 | 27 | 0 | 2.3 | 5.25 | 2 | -4.05 | -6.2 | -7.3 | -6.95 | -1.15 | 4.05 | 5.2 | 9.15 | 14.45 | 16.65 | 22.2 | 30.2 | 28.45 | 16 | 0.7 | -13.4 | -23.2 | -25.35 | -18.35 | -6.6 | -0.3 | -2.15 | -6 | -9.05 | -10.4 | -8.75 | -5.2 | -0.8 | 2.8 | 3.6 | 2.05 | -1.25 | -5.25 | -6.45 | -3.55 | -1.2 | -3.2 | -4.95 | -2.2 | 0.7 | 0.7
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 5a7b815a-1309-ed11-85f4-5cf370a20867 | 503c8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.17 | [21 empty] | 010081427802027710118705 2105881 17240328 | 7 | 2 | V834 | 2022-07-21 11:37 | 563 | 270 | 16.2757 | 2127.102 | 5 | 27 | 0 | 2.25 | 5.35 | 4.6 | 2.35 | -0.95 | -4.8 | -3.7 | -1.85 | -2.9 | 0.05 | 7.2 | 14.15 | 21.25 | 27.5 | 30.75 | 28.3 | 18.3 | 4.7 | -7.4 | -16.5 | -23.6 | -27 | -22.2 | -13.3 | -7.9 | -3.6 | 1.55 | 1.6 | -3.55 | -6.4 | -4.9 | -3.4 | -2.45 | 0.45 | 1.85 | -0.5 | -1.7 | 0.2 | 2.2 | 2.15 | 0.65 | -0.85 | -0.8 | 1.05
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 607b815a-1309-ed11-85f4-5cf370a20867 | 22597999-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.16 | [21 empty] | 010081427802027710118705 2105881 17240328 | 8 | 2 | V834 | 2022-07-21 11:37 | 569 | 268 | 14.8879 | 2018.9233 | 5 | 27 | 0 | 0.2 | -0.85 | -3.85 | -4.05 | -0.95 | 0.35 | 1.75 | 4.6 | 4.35 | 2.35 | 1.4 | 1.9 | 10.6 | 25.65 | 32.7 | 29.65 | 21.3 | 6.8 | -9.25 | -19.15 | -21.55 | -18.3 | -14.2 | -12.15 | -8.85 | -4.3 | -2.8 | -6.1 | -9.1 | -5.65 | -1.35 | -2.1 | -1.1 | 3.5 | 5 | 3.15 | 0.35 | -1.95 | -2.55 | -3.8 | -4.85 | -2.1 | 1.9 | 3.6
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 667b815a-1309-ed11-85f4-5cf370a20867 | 6e3e8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.15 | [21 empty] | 010081427802027710118705 2105881 17240328 | 9 | 2 | V834 | 2022-07-21 11:37 | 664 | 265 | 14.7126 | 2116.7802 | 5 | 27 | 0 | -0.2 | -0.35 | -0.75 | -2.15 | -2.05 | -0.3 | -0.9 | -0.6 | 4.95 | 8.7 | 6.5 | 7.2 | 14 | 19.8 | 23.3 | 25.2 | 19.6 | 6.55 | -5.5 | -14.75 | -22.55 | -24.15 | -18.5 | -11.4 | -5.5 | -0.65 | 0.4 | -2.45 | -5.15 | -7.35 | -7.9 | -3.55 | 1.6 | 2.4 | 1.6 | 1.4 | 0.85 | 1.75 | 2.35 | -0.85 | -3.6 | -2.95 | -1.55 | -0.5
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 6c7b815a-1309-ed11-85f4-5cf370a20867 | 103e8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105881 17240328 | 10 | 2 | V834 | 2022-07-21 11:37 | 659 | 280 | 21.4891 | 2038.9826 | 5 | 27 | 0 | -1.25 | -3.45 | -0.75 | 4.8 | 3.6 | 0.25 | 0.3 | -1.45 | -2.55 | 1.35 | 6.7 | 13.5 | 23.35 | 32.8 | 40.1 | 42.6 | 31.65 | 9.75 | -8.9 | -24.35 | -38 | -36.25 | -21.75 | -13.4 | -11.2 | -6.8 | -4.75 | -6.95 | -6.6 | -3.35 | -3.15 | -4.95 | -3.6 | -1.15 | -0.8 | -0.8 | 1.65 | 5 | 3.05 | -1.5 | -0.55 | 0.8 | -1.8 | -2.8
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 727b815a-1309-ed11-85f4-5cf370a20867 | b23d8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.22 | [21 empty] | 010081427802027710118705 2105881 17240328 | 11 | 2 | V834 | 2022-07-21 11:37 | 627 | 270 | 20.9273 | 2048.0593 | 5 | 27 | 0 | 1.6 | 3 | 1.3 | 2.2 | 4.75 | 1.65 | -1.4 | 0.1 | 0.75 | -0.35 | 0.6 | 8.1 | 22.75 | 35.8 | 42.55 | 44.05 | 33.3 | 9.6 | -13 | -25.6 | -29.8 | -28.05 | -21.55 | -15.5 | -16.35 | -17.65 | -9.35 | -0.55 | 0.85 | 3.2 | 3.8 | -4.95 | -11.55 | -7.45 | -2.7 | -2.4 | -0.7 | 2.9 | 4.4 | 6.1 | 9.8 | 7.4 | 1.2 | 1.55
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 787b815a-1309-ed11-85f4-5cf370a20867 | f0fe3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.24 | [21 empty] | 010081427802027710118705 2105881 17240328 | 4 | 2 | V834 | 2022-07-21 11:37 | 633 | 295 | 24.0296 | 2104.43 | 5 | 27 | 0 | 0.25 | -0.5 | -0.6 | 3.3 | 4.4 | -0.95 | -4.05 | -2.35 | -0.6 | 1.5 | 7.75 | 16.4 | 23.65 | 32.45 | 42.1 | 43.2 | 32.15 | 11.1 | -14.55 | -30.7 | -33.75 | -34.25 | -29.35 | -15.55 | -6.6 | -5.75 | -3.35 | -2.25 | -6.25 | -8.95 | -9 | -8.15 | -4.05 | 2.5 | 5.4 | 2.55 | 0.05 | 0.35 | -0.35 | -1.3 | -0.65 | -1.8 | -4.15 | -3.9
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 7e7b815a-1309-ed11-85f4-5cf370a20867 | 543d8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.27 | [21 empty] | 010081427802027710118705 2105881 17240328 | 5 | 2 | V834 | 2022-07-21 11:37 | 596 | 273 | 27.6573 | 2095.9352 | 5 | 27 | 0 | 2.15 | 4.85 | 2.35 | -1.85 | -3 | -3.45 | -4 | -3.45 | -0.4 | 3.35 | 5.1 | 12.05 | 28.8 | 43.55 | 48.95 | 49 | 39.65 | 16.15 | -11.75 | -30.35 | -37.8 | -40.3 | -35.85 | -23.7 | -17 | -17.85 | -12.25 | -2.3 | 1.65 | 3.35 | 3.6 | -2.3 | -8.7 | -9.45 | -5.6 | 0.7 | 4 | 2.5 | 2.85 | 5.6 | 5.25 | 1 | -1.15 | 1.85
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 847b815a-1309-ed11-85f4-5cf370a20867 | 953c8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.23 | 58.23 | 1.25 | 315 | [20 empty] | 010081427802027710118705 2105881 17240328 | 6 | 2 | V834 | 2022-07-21 11:37 | 617 | 276 | 6.4 | 2303.3333 | 2 | 4 | 0 | -1.45 | -5.85 | -8.85 | -5.45 | 0.1 | 1.2 | 0.5 | 2.7 | 2.5 | -0.65 | 3.6 | 13.5 | 24.2 | 42.15 | 57.65 | 52.05 | 32.3 | 9.65 | -16.8 | -36.75 | -40.65 | -36.45 | -29.1 | -18.15 | -9.85 | -6.95 | -4.15 | -2.4 | -6.95 | -12.8 | -10.2 | -3.6 | -3.35 | -3.2 | 2.95 | 5.55 | 2.7 | 2.9 | 3.15 | -1.9 | -6.35 | -6.45 | -4.45 | -2.4
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | b3de19db-1309-ed11-85f4-5cf370a20867 | 3a5f6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.74 | 58.9 | 1.27 | 541 | [20 empty] | 010081427802027710118705 2105881 17240328 | 3 | 2 | V834 | 2022-07-21 11:41 | 592 | 269 | 2.5 | 2335.5 | 2 | 4 | 0 | 0.15 | 0.15 | -1.15 | -2.8 | -3 | -3 | -1.5 | 3.15 | 4.25 | 0.85 | 4.35 | 15.85 | 26.95 | 39.2 | 53.6 | 58.25 | 43.6 | 16.3 | -10.1 | -29.7 | -41.15 | -41.2 | -32.35 | -23.45 | -17.45 | -12.65 | -9.45 | -7.4 | -5.55 | -3.8 | -3 | -3.1 | -1.4 | -0.45 | -2.8 | -1.6 | 1.85 | -1.65 | -6.3 | -5.45 | -5.65 | -6.35 | -3.05 | -0.2
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | b9de19db-1309-ed11-85f4-5cf370a20867 | a75e6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.19 | [21 empty] | 010081427802027710118705 2105882 17240328 | 10 | 1 | V639 | 2022-07-21 11:41 | 614 | 260 | 19.8022 | 2195.6174 | 5 | 27 | 0 | -1.25 | -3 | -1.25 | 2.35 | 3.45 | 2.05 | -2.6 | -7.45 | -5.7 | -0.1 | 5.7 | 14.6 | 24.45 | 32.1 | 36.8 | 33.4 | 21.4 | 6.85 | -8.8 | -22.45 | -26.5 | -22.2 | -18.35 | -16.4 | -11.5 | -5.75 | -4.6 | -5.45 | -4.95 | -5.1 | -4.6 | -2 | -0.7 | 0.25 | 2.45 | 0.35 | -5.5 | -5.7 | -0.5 | 0.4 | -3 | -3.5 | -1.9 | -1.6
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | bfde19db-1309-ed11-85f4-5cf370a20867 | 185e6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.26 | 56.9 | 1.29 | 510 | [20 empty] | 010081427802027710118705 2105882 17240328 | 11 | 1 | V639 | 2022-07-21 11:41 | 588 | 272 | 3.9 | 2333.5 | 2 | 4 | 0 | -1.55 | -5 | -6.25 | -3.1 | 2.6 | 6.45 | 4.15 | 0.05 | -0.6 | -0.6 | 3.1 | 17.3 | 35.4 | 48.6 | 56.55 | 54 | 34.4 | 5 | -20.15 | -35.75 | -43.4 | -40.95 | -30.15 | -20.55 | -14.25 | -8.35 | -5.35 | -4.35 | -2.55 | -3.3 | -5.1 | -3.75 | -2.35 | -3.8 | -6.1 | -5.75 | -3.3 | -2.2 | -0.45 | 3.75 | 4.25 | -0.65 | -3.15 | -1.2
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | c5de19db-1309-ed11-85f4-5cf370a20867 | bffe3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105882 17240328 | 12 | 1 | V639 | 2022-07-21 11:41 | 672 | 260 | 21.7893 | 2250.3186 | 5 | 27 | 0 | -0.85 | -2.8 | -2.55 | 0 | 0.05 | -0.2 | 3.35 | 6.4 | 4.8 | 1.55 | 2.15 | 8.4 | 18.7 | 29.65 | 37.15 | 37.8 | 27 | 6.4 | -11.1 | -19.8 | -25.2 | -25.4 | -19 | -15.15 | -14.65 | -11 | -8.3 | -9 | -6.25 | -2.6 | -3.1 | -1.55 | 1.15 | -0.7 | 0.05 | 4.3 | 3.1 | -0.7 | -1.8 | -3.2 | -6 | -7.1 | -5.95 | -4.85
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | cbde19db-1309-ed11-85f4-5cf370a20867 | 32fe3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.23 | [21 empty] | 010081427802027710118705 2105881 17240328 | 1 | 2 | V834 | 2022-07-21 11:41 | 679 | 291 | 23.8111 | 2216.2577 | 5 | 27 | 0 | 0.4 | 0.55 | 0.1 | 0.6 | 0.15 | -1.7 | -0.95 | 0.2 | 0.35 | 4 | 9.45 | 14.65 | 22.55 | 31.4 | 37.05 | 37.25 | 30.35 | 17.05 | -0.25 | -17 | -27.9 | -32.1 | -32 | -28.4 | -19.5 | -8.6 | -3.8 | -3 | -0.9 | -2.3 | -6.45 | -5.8 | -3.25 | -3.7 | -4.05 | -1.25 | 3.55 | 5.5 | 3.2 | 1.55 | -0.3 | -5.75 | -9.25 | -7.7
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | d1de19db-1309-ed11-85f4-5cf370a20867 | c65f6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.22 | [21 empty] | 010081427802027710118705 2105881 17240328 | 2 | 2 | V834 | 2022-07-21 11:41 | 600 | 276 | 24.1291 | 2174.6047 | 5 | 27 | 0 | 0.75 | 0.65 | -0.3 | 1.6 | 2.05 | -1.4 | -3.35 | -3.55 | -2.1 | 1 | 4.1 | 11.6 | 24.75 | 36.9 | 44.35 | 44.2 | 32.1 | 10.05 | -13.6 | -30.7 | -38.3 | -36.05 | -26.25 | -16.75 | -10.4 | -6.15 | -6.65 | -8.05 | -5.2 | -2.7 | -1.85 | 1.3 | 3.65 | 0.8 | -3.95 | -4.7 | -1.85 | -1.2 | -2.35 | -0.5 | 0.8 | -2.5 | -5.4 | -5.05
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | d7de19db-1309-ed11-85f4-5cf370a20867 | 674775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105882 17240328 | 8 | 1 | V639 | 2022-07-21 11:41 | 620 | 261 | 21.2957 | 2243.913 | 5 | 27 | 0 | 0.85 | 2.5 | 1.45 | -2.25 | -2.05 | 0.15 | -3.15 | -3.4 | 5.75 | 10.5 | 9.35 | 14.9 | 23.35 | 28.45 | 32.9 | 30.35 | 17.55 | 4.75 | -7.55 | -22.5 | -28.05 | -22.4 | -19.95 | -19.6 | -12.9 | -6.05 | -4.5 | -3.05 | -1.4 | -4.2 | -6.65 | -1.85 | 2.9 | 1.2 | 0.45 | 0.5 | -3.3 | -4.95 | -2.55 | -0.6 | 0.5 | -0.4 | -3.4 | -4.65
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | ddde19db-1309-ed11-85f4-5cf370a20867 | 095f6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105882 17240328 | 9 | 1 | V639 | 2022-07-21 11:41 | 661 | 258 | 21.6571 | 2230.5644 | 5 | 27 | 0 | -2.6 | -8.2 | -11.15 | -9.85 | -4.75 | 4.1 | 9.6 | 7.4 | 3.6 | 3.05 | 4.75 | 8.7 | 18.7 | 34.7 | 44.05 | 35.75 | 17.8 | 0.4 | -18.8 | -34 | -32 | -20.15 | -14.4 | -14 | -11.9 | -9.1 | -7.6 | -4.65 | -0.95 | -0.55 | -1.4 | -1.75 | -2.45 | 0.3 | 3.25 | -0.4 | -3.3 | 0.2 | 2.15 | -0.45 | -2.35 | -2.45 | -1.5 | 0.4
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 65fdae7a-1409-ed11-85f4-5cf370a20867 | 084775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105882 17240328 | 4 | 1 | V639 | 2022-07-21 11:45 | 684 | 263 | 21.7674 | 2166.2783 | 5 | 27 | 0 | -0.15 | -1.75 | -4.55 | -5.65 | -4.5 | -2 | 2.5 | 7.25 | 6.6 | 1.95 | 2.55 | 9.95 | 19.7 | 31.9 | 43.2 | 41.7 | 25.4 | 6.05 | -11.3 | -26.65 | -32.25 | -26 | -20.05 | -18.3 | -12.85 | -7.1 | -7.2 | -5.2 | -0.15 | -2.05 | -5.05 | -1.25 | 0.15 | -3.25 | -3.55 | -1.75 | -2.8 | -5.35 | -3.95 | 1.25 | 2.7 | -0.55 | -2.65 | -2.6
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 6bfdae7a-1409-ed11-85f4-5cf370a20867 | d74675b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105882 17240328 | 5 | 1 | V639 | 2022-07-21 11:45 | 703 | 280 | 23.3328 | 2329.7447 | 5 | 27 | 0 | 1.9 | 5.15 | 3.7 | -2 | -5.2 | -4.55 | -0.75 | 5.15 | 8.05 | 7.05 | 8.05 | 14.4 | 23.75 | 31.7 | 36 | 33.05 | 19.15 | 1.85 | -11.85 | -24.95 | -31.85 | -26.35 | -19.45 | -17.7 | -13.05 | -4.1 | 0.45 | -1.85 | -2.3 | 0.1 | -4.35 | -10.9 | -9.05 | -4.8 | -3.4 | -1 | 1.2 | 3.05 | 5.5 | 3.1 | -1.55 | -1.95 | -0.05 | 1.85
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 71fdae7a-1409-ed11-85f4-5cf370a20867 | 494675b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 15.89 | 59.99 | 1.29 | 667 | [20 empty] | 010081427802027710118705 2105882 17240328 | 6 | 1 | V639 | 2022-07-21 11:45 | 749 | 264 | 1.2 | 2464.1333 | 2 | 4 | 0 | 0.6 | 2.6 | 2.3 | -1.4 | -1.55 | -0.35 | -2.65 | -2.9 | -1.05 | 1.15 | 9.8 | 24.65 | 40.25 | 54.35 | 59.9 | 51.1 | 32.75 | 10.8 | -12.55 | -34.2 | -46.2 | -44.85 | -38.75 | -32.4 | -20.85 | -8.35 | -4.35 | -4.55 | -2.7 | -2.9 | -6.15 | -7 | -3.7 | -1.05 | -0.95 | 1.55 | 4.55 | 1.25 | -2.8 | -2 | -2.85 | -5.9 | -4.9 | -1.55
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 77fdae7a-1409-ed11-85f4-5cf370a20867 | f44775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105882 17240328 | 7 | 1 | V639 | 2022-07-21 11:45 | 611 | 265 | 21.8399 | 2179.8047 | 5 | 27 | 0 | 0.95 | 1.9 | -0.15 | -3.25 | -2.9 | 0 | 1.55 | 2.8 | 4.8 | 4.4 | 5.1 | 12.6 | 22.7 | 30.35 | 34.95 | 32.9 | 21.5 | 5.45 | -9.45 | -21.95 | -28.9 | -26.05 | -19.4 | -17 | -15.3 | -10.8 | -7.7 | -6.5 | -3.8 | 0.15 | 2.15 | 1.75 | 1.3 | -0.85 | -4 | -2.75 | -0.45 | -2 | -2.8 | -2.05 | -4.2 | -5.65 | -3.2 | -1.2
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | f47fc1ae-1609-ed11-85f4-5cf370a20867 | 536b47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.19 | [21 empty] | 010081427802027710118705 2105882 17240328 | 1 | 1 | V639 | 2022-07-21 12:01 | 725 | 269 | 21.8492 | 2389.1866 | 5 | 27 | 0 | 0.8 | 1 | -1.1 | -1.05 | 1.4 | 0.3 | -1.2 | 1.8 | 3.3 | 1.3 | 3.2 | 10.45 | 21.75 | 35.8 | 41.5 | 31.85 | 16.15 | 0.35 | -16.35 | -23.35 | -16.5 | -13.65 | -19.75 | -20.35 | -16.2 | -14.65 | -9.05 | -0.05 | 1.05 | -4.05 | -5.6 | -2.95 | -2.15 | -3.1 | 1.4 | 8.8 | 5.85 | -5.75 | -9.6 | -6.25 | -7.45 | -7.6 | -1.4 | 2.7
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | fa7fc1ae-1609-ed11-85f4-5cf370a20867 | f46a47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105882 17240328 | 2 | 1 | V639 | 2022-07-21 12:01 | 719 | 260 | 23.9073 | 2499.5613 | 5 | 27 | 0 | 0.35 | 1 | 0.15 | -1.85 | -1.75 | -0.45 | 0.15 | 1.95 | 3.85 | 4.5 | 8.2 | 15.3 | 22.8 | 32.75 | 41.45 | 38.35 | 24.1 | 6.95 | -12.1 | -30.95 | -38.95 | -32.35 | -20.55 | -11.5 | -5.85 | -2.85 | -2.9 | -5.7 | -7.9 | -6.4 | -3.75 | -3.25 | -2.3 | -0.35 | -1.45 | -3.35 | -2.25 | -1.7 | -2.8 | -1.45 | 0.65 | -0.75 | -2.7 | -2.15
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | 0080c1ae-1609-ed11-85f4-5cf370a20867 | b36b47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16 | 50.4 | 1.21 | 402 | [20 empty] | 010081427802027710118705 2105882 17240328 | 3 | 1 | V639 | 2022-07-21 12:01 | 714 | 262 | 3.7 | 2589.5667 | 2 | 4 | 0 | 0.55 | 2.1 | 2.2 | -1.3 | -3.45 | -1.35 | -1.9 | -4.95 | -2.1 | 4.15 | 7.55 | 13.15 | 26.9 | 43.15 | 50.4 | 43.2 | 24.8 | 1.05 | -20.9 | -34.4 | -37.55 | -31.15 | -19.8 | -11.05 | -7.7 | -6.3 | -4.2 | -4.85 | -9.35 | -9.5 | -4.9 | -4.65 | -3.95 | 2.15 | 3.6 | 1.6 | 4.35 | 3.15 | -5.15 | -7.4 | -3.75 | -3.55 | -4.15 | -2.55
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | a79dad68-1f09-ed11-85f4-5cf370a20867 | b16dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | ad9dad68-1f09-ed11-85f4-5cf370a20867 | b56dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | b39dad68-1f09-ed11-85f4-5cf370a20867 | b96dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | b99dad68-1f09-ed11-85f4-5cf370a20867 | bd6dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | bf9dad68-1f09-ed11-85f4-5cf370a20867 | c16dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | c59dad68-1f09-ed11-85f4-5cf370a20867 | c66dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | cb9dad68-1f09-ed11-85f4-5cf370a20867 | 010bca23-2809-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.17 | [21 empty] | 010081427802027710118705 2105883 17240328 | 11 | 4 | V1169 | 2022-07-21 13:03 | 870 | 289 | 20.3816 | 2650.1897 | 5 | 27 | 0 | -0.25 | -2.7 | -5.1 | -3.4 | -2.75 | -4.4 | -3.85 | -3.25 | -2.1 | 1.5 | 4.8 | 10.2 | 19.65 | 27.4 | 32.6 | 36.95 | 33.7 | 20.15 | 2.1 | -15.65 | -27.25 | -29.65 | -26.3 | -20.45 | -14.4 | -10.55 | -9.7 | -9.9 | -6.75 | -1.4 | 0.05 | -0.1 | 1.6 | 1.25 | -0.9 | -1.45 | -0.65 | 0.75 | 2.05 | 1.9 | -0.6 | -4.9 | -7.15 | -5.7
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | d29dad68-1f09-ed11-85f4-5cf370a20867 | a30aca23-2809-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.18 | [21 empty] | 010081427802027710118705 2105883 17240328 | 12 | 4 | V1169 | 2022-07-21 13:03 | 839 | 276 | 21.0879 | 2618.4016 | 5 | 27 | 0 | 2.35 | 6.5 | 7.45 | 6.6 | 4.55 | -1.65 | -5.15 | -0.55 | 3.15 | 1.1 | 2.25 | 10.9 | 20.65 | 27.95 | 34.2 | 35.35 | 25.7 | 8.75 | -7.4 | -16.55 | -18.45 | -19.85 | -20.5 | -16.55 | -15.6 | -17.65 | -12.75 | -7.05 | -5.25 | 0.8 | 4.55 | -1.1 | -4.55 | -3.05 | -2.95 | 0.35 | 5.75 | 4.45 | 0.6 | -1.05 | -5.05 | -8.95 | -6.95 | -2.75
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 1f42d920-2009-ed11-85f4-5cf370a20867 | 13e5bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.45 | 54.87 | 1.25 | 703 | [20 empty] | 010081427802027710118705 2105883 17240328 | 5 | 4 | V1169 | 2022-07-21 13:09 | 716 | 286 | 0.1 | 2777.1 | 2 | 4 | 0 | 1.5 | 5.55 | 6.8 | 2.85 | 1.15 | 0.5 | -4.9 | -6.95 | -1.6 | 3.45 | 7.75 | 15.65 | 27.75 | 42.7 | 53.7 | 53.1 | 40.4 | 17.65 | -10.8 | -34.35 | -43.6 | -40.1 | -31.25 | -21.5 | -14.55 | -11.8 | -10.4 | -8.9 | -5.95 | -0.95 | 2.6 | 3.75 | 3.7 | -0.25 | -7.05 | -8.5 | -4.05 | -2.7 | -3.5 | 0.7 | 2.8 | -3.65 | -8.15 | -5.8
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 2642d920-2009-ed11-85f4-5cf370a20867 | e1e4bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.28 | 52.81 | 1.23 | 438 | [20 empty] | 010081427802027710118705 2105883 17240328 | 6 | 4 | V1169 | 2022-07-21 13:09 | 732 | 276 | 5 | 2706.5333 | 2 | 4 | 0 | -0.15 | -1.2 | -2.15 | -2.5 | -4.1 | -3.4 | 1.4 | 4.3 | 4.9 | 6.4 | 8.4 | 14.2 | 27.55 | 43.3 | 52.35 | 49.8 | 33.25 | 3.5 | -25.65 | -38.75 | -39.15 | -34.5 | -24.6 | -15.05 | -12 | -11.15 | -8.3 | -5.75 | -5.6 | -5.65 | -3.55 | -1.15 | -0.9 | -1.15 | -1.05 | -3.85 | -6.75 | -1.1 | 6.55 | 2.5 | -4.55 | -3.1 | 0.65 | 2.15
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 2c42d920-2009-ed11-85f4-5cf370a20867 | 8fe6bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.18 | [21 empty] | 010081427802027710118705 2105883 17240328 | 7 | 4 | V1169 | 2022-07-21 13:09 | 837 | 291 | 20.8789 | 2564.3731 | 5 | 27 | 0 | 0.1 | 0 | -1.35 | -4 | -4.45 | -1.2 | 0.95 | 1.45 | 3.4 | 4.75 | 5.3 | 8.05 | 14.8 | 26.3 | 36.35 | 36.75 | 28.15 | 14.55 | -1.85 | -18.05 | -29.2 | -29.7 | -21.85 | -17.7 | -18.3 | -13.85 | -5.85 | -1.1 | 0.35 | 0.8 | 0.5 | -2.45 | -5.95 | -3.8 | 1.1 | 1.1 | -0.6 | -0.35 | -2.25 | -5.2 | -4.15 | -1.25 | 0.3 | 1.4
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 3242d920-2009-ed11-85f4-5cf370a20867 | 30e6bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.19 | [21 empty] | 010081427802027710118705 2105883 17240328 | 8 | 4 | V1169 | 2022-07-21 13:09 | 768 | 291 | 21.1269 | 2541.8308 | 5 | 27 | 0 | 2.2 | 4.65 | 1.95 | -0.75 | -0.3 | -2.1 | -2.5 | 3.15 | 8 | 7.3 | 5.65 | 8.75 | 16.75 | 24.8 | 31.3 | 36.05 | 30.25 | 9.35 | -13.05 | -24.6 | -27.6 | -25.65 | -18.2 | -9.55 | -7.45 | -9.4 | -8.9 | -8.3 | -6.45 | 0.4 | 4.4 | 1.05 | -2.45 | -3.7 | -6 | -8 | -4.65 | 3.45 | 7.5 | 4.05 | -0.65 | -3.95 | -5.4 | -3.8
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 3842d920-2009-ed11-85f4-5cf370a20867 | 00e6bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.78 | 57.59 | 1.25 | 502 | [20 empty] | 010081427802027710118705 2105883 17240328 | 9 | 4 | V1169 | 2022-07-21 13:09 | 775 | 295 | 4.4 | 2755.0667 | 2 | 4 | 0 | 1.1 | 1.95 | -1.85 | -8.1 | -9.4 | -4.5 | 1.5 | 5.95 | 8.55 | 7.4 | 3.8 | 4.8 | 16.7 | 36.55 | 53.1 | 57.25 | 46.8 | 23.9 | -4.25 | -28.25 | -41.15 | -39.6 | -30.1 | -24.1 | -20.75 | -13.25 | -6.8 | -6.6 | -5.15 | -0.95 | -2.7 | -6.4 | -3.05 | 1 | 1.5 | 2.3 | -0.35 | -6.05 | -5.4 | -0.65 | 0.65 | -0.35 | -0.6 | 0.75
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 3e42d920-2009-ed11-85f4-5cf370a20867 | 73e5bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.61 | 57.38 | 1.23 | 706 | [20 empty] | 010081427802027710118705 2105883 17240328 | 10 | 4 | V1169 | 2022-07-21 13:09 | 849 | 286 | -0.7 | 2830.2333 | 2 | 4 | 0 | 3.05 | 9.05 | 10.25 | 4.4 | -3.3 | -7.65 | -6.55 | -1.75 | 3.65 | 7.05 | 8.9 | 13.2 | 23.45 | 39.5 | 54.3 | 56.15 | 41.6 | 18.5 | -8.25 | -35.25 | -47 | -38.45 | -28.25 | -22.8 | -14.55 | -8.95 | -8.9 | -7.15 | -5.6 | -6.05 | -3.3 | -1.2 | -3.8 | -3.45 | 1.25 | 1.85 | -0.7 | -0.55 | -0.4 | -2.2 | -3.8 | -5.35 | -5.7 | -4.05
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 4442d920-2009-ed11-85f4-5cf370a20867 | 803d7504-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:09 |  |  |  |  | 0
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | b7477ca2-3609-ed11-85f4-5cf370a20867 | f2c2a3da-3609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:50 |  |  |  |  | 0
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | f5c2a3da-3609-ed11-85f4-5cf370a20867 | 8db7eb13-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:51 |  |  |  |  | 0
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 90b7eb13-3709-ed11-85f4-5cf370a20867 | b7bbce4d-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:53 |  |  |  |  | 0
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | f2bbce4d-3709-ed11-85f4-5cf370a20867 | 0b592f87-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:54 |  |  |  |  | 0
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 0e592f87-3709-ed11-85f4-5cf370a20867 | 4e3e36c2-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Aborted | TargetAborted | 7162 (Warning, Quantity not sufficient), 1000 (Error, User Aborted Test) |  |  | 0 | [25 empty] | 2022-07-21 15:56 |  |  |  |  | 0
Modulated
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 0d5c9e96-1209-ed11-85f4-5cf370a20867 | 5fc854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105880 17240328 | 8 | 3 | V790 | 2022-07-21 11:32 | 501 | 259 | 18.9314 | 2168.8451 | 5 | 27 | 2353 | 2344 | 2332 | 2339 | 2349 | 2351 | 2345 | 2348 | 2352 | 2358 | 2350 | 2363 | 2367 | 2377 | 2370 | 2382 | 2414 | 2437 | 2503 | 2533 | 2603 | 2626 | 2662 | 2677 | 2669 | 2701 | 2709 | 2720 | 2739 | 2748 | 2738 | 2758 | 2749 | 2767 | 2761 | 2780 | 2781 | 2787 | 2788 | 2785 | 2803 | 2793 | 2794 | 2806 | 2814
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 145c9e96-1209-ed11-85f4-5cf370a20867 | 01c854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.27 | [21 empty] | 010081427802027710118705 2105880 17240328 | 9 | 3 | V790 | 2022-07-21 11:32 | 555 | 259 | 27.7921 | 2118.1265 | 5 | 27 | 2354 | 2369 | 2368 | 2376 | 2385 | 2399 | 2385 | 2382 | 2381 | 2400 | 2398 | 2397 | 2395 | 2402 | 2406 | 2433 | 2468 | 2517 | 2597 | 2675 | 2726 | 2803 | 2826 | 2857 | 2892 | 2909 | 2908 | 2929 | 2933 | 2953 | 2956 | 2968 | 2974 | 2983 | 2969 | 2984 | 2992 | 2990 | 2995 | 3024 | 3006 | 3012 | 3032 | 3029 | 3021
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 1b5c9e96-1209-ed11-85f4-5cf370a20867 | 4aca54cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105880 17240328 | 10 | 3 | V790 | 2022-07-21 11:32 | 524 | 267 | 20.2032 | 2087.4016 | 5 | 27 | 2265 | 2276 | 2262 | 2263 | 2269 | 2274 | 2271 | 2294 | 2294 | 2277 | 2288 | 2278 | 2294 | 2306 | 2309 | 2322 | 2358 | 2379 | 2442 | 2508 | 2530 | 2580 | 2603 | 2624 | 2632 | 2651 | 2665 | 2658 | 2674 | 2686 | 2693 | 2683 | 2722 | 2728 | 2700 | 2726 | 2716 | 2718 | 2712 | 2721 | 2734 | 2728 | 2748 | 2735 | 2727
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 225c9e96-1209-ed11-85f4-5cf370a20867 | e9c954cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105880 17240328 | 11 | 3 | V790 | 2022-07-21 11:32 | 509 | 246 | 19.9302 | 2120.551 | 5 | 27 | 2283 | 2311 | 2294 | 2318 | 2292 | 2307 | 2290 | 2319 | 2328 | 2310 | 2329 | 2322 | 2322 | 2331 | 2327 | 2354 | 2378 | 2408 | 2482 | 2530 | 2559 | 2598 | 2623 | 2647 | 2673 | 2679 | 2693 | 2678 | 2700 | 2723 | 2716 | 2711 | 2717 | 2739 | 2729 | 2741 | 2741 | 2763 | 2748 | 2763 | 2751 | 2752 | 2770 | 2766 | 2782
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 295c9e96-1209-ed11-85f4-5cf370a20867 | 87c954cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.52 | 55.01 | 1.28 | 468 | [20 empty] | 010081427802027710118705 2105880 17240328 | 12 | 3 | V790 | 2022-07-21 11:32 | 499 | 265 | 4.9 | 2369.2333 | 2 | 4 | 2373 | 2374 | 2384 | 2392 | 2395 | 2400 | 2408 | 2396 | 2403 | 2396 | 2402 | 2406 | 2414 | 2431 | 2427 | 2439 | 2477 | 2540 | 2624 | 2707 | 2771 | 2820 | 2863 | 2892 | 2916 | 2941 | 2947 | 2953 | 2970 | 3001 | 2979 | 3015 | 3004 | 3004 | 3001 | 3031 | 3016 | 3030 | 3038 | 3042 | 3044 | 3050 | 3052 | 3047 | 3064
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 305c9e96-1209-ed11-85f4-5cf370a20867 | e454c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105879 17240328 | 1 | 4 | V1169 | 2022-07-21 11:32 | 785 | 294 | 23.1577 | 2506.8071 | 5 | 27 | 2720 | 2708 | 2714 | 2721 | 2725 | 2735 | 2726 | 2735 | 2727 | 2741 | 2744 | 2727 | 2767 | 2733 | 2742 | 2766 | 2799 | 2841 | 2897 | 2973 | 3019 | 3066 | 3114 | 3126 | 3135 | 3152 | 3181 | 3194 | 3193 | 3193 | 3201 | 3222 | 3242 | 3247 | 3229 | 3231 | 3236 | 3260 | 3250 | 3254 | 3255 | 3268 | 3259 | 3284 | 3274
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 375c9e96-1209-ed11-85f4-5cf370a20867 | 8554c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105879 17240328 | 2 | 4 | V1169 | 2022-07-21 11:32 | 729 | 292 | 23.2439 | 2407.3328 | 5 | 27 | 2614 | 2623 | 2601 | 2620 | 2626 | 2622 | 2642 | 2633 | 2619 | 2644 | 2639 | 2650 | 2647 | 2658 | 2656 | 2678 | 2698 | 2746 | 2792 | 2858 | 2911 | 2968 | 3002 | 3024 | 3058 | 3067 | 3078 | 3103 | 3082 | 3120 | 3113 | 3143 | 3151 | 3139 | 3165 | 3160 | 3165 | 3152 | 3164 | 3165 | 3173 | 3180 | 3179 | 3188 | 3180
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 3e5c9e96-1209-ed11-85f4-5cf370a20867 | 4455c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105879 17240328 | 3 | 4 | V1169 | 2022-07-21 11:32 | 739 | 274 | 22.2113 | 2319.6806 | 5 | 27 | 2519 | 2507 | 2528 | 2525 | 2533 | 2515 | 2529 | 2533 | 2536 | 2551 | 2539 | 2550 | 2547 | 2549 | 2538 | 2573 | 2600 | 2644 | 2704 | 2790 | 2814 | 2863 | 2879 | 2911 | 2939 | 2934 | 2942 | 2963 | 2962 | 2988 | 2984 | 2990 | 2994 | 3004 | 3006 | 3023 | 3021 | 3018 | 3024 | 3023 | 3043 | 3031 | 3032 | 3040 | 3035
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 6042e1f8-1209-ed11-85f4-5cf370a20867 | f53c8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.52 | 53.44 | 1.22 | 471 | [20 empty] | 010081427802027710118705 2105881 17240328 | 12 | 2 | V834 | 2022-07-21 11:34 | 614 | 257 | 0.7 | 2263.5 | 2 | 4 | 2256 | 2273 | 2271 | 2269 | 2258 | 2261 | 2269 | 2273 | 2255 | 2276 | 2254 | 2271 | 2285 | 2280 | 2264 | 2290 | 2299 | 2349 | 2428 | 2501 | 2551 | 2586 | 2626 | 2642 | 2689 | 2669 | 2692 | 2704 | 2714 | 2735 | 2723 | 2733 | 2741 | 2745 | 2740 | 2744 | 2758 | 2747 | 2749 | 2752 | 2760 | 2774 | 2763 | 2770 | 2766
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 6642e1f8-1209-ed11-85f4-5cf370a20867 | 508c0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105880 17240328 | 1 | 3 | V790 | 2022-07-21 11:34 | 494 | 260 | 20.1514 | 2002.9866 | 5 | 27 | 2184 | 2192 | 2175 | 2191 | 2203 | 2187 | 2200 | 2206 | 2199 | 2209 | 2210 | 2193 | 2214 | 2214 | 2211 | 2212 | 2250 | 2296 | 2341 | 2416 | 2463 | 2481 | 2507 | 2550 | 2564 | 2576 | 2579 | 2602 | 2583 | 2599 | 2604 | 2606 | 2629 | 2637 | 2636 | 2641 | 2643 | 2642 | 2654 | 2642 | 2660 | 2665 | 2663 | 2659 | 2667
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 6c42e1f8-1209-ed11-85f4-5cf370a20867 | 1d8c0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.22 | [21 empty] | 010081427802027710118705 2105880 17240328 | 2 | 3 | V790 | 2022-07-21 11:34 | 513 | 269 | 19.7261 | 2017.5217 | 5 | 27 | 2208 | 2199 | 2200 | 2226 | 2210 | 2202 | 2214 | 2218 | 2204 | 2220 | 2210 | 2220 | 2202 | 2209 | 2207 | 2228 | 2259 | 2303 | 2366 | 2424 | 2476 | 2488 | 2517 | 2545 | 2563 | 2589 | 2573 | 2597 | 2587 | 2623 | 2630 | 2637 | 2642 | 2654 | 2648 | 2657 | 2666 | 2654 | 2660 | 2660 | 2642 | 2670 | 2682 | 2671 | 2691
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 7242e1f8-1209-ed11-85f4-5cf370a20867 | 918b0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold), 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed) | 16.37 | 60.68 | 1.34 | 745 | [20 empty] | 010081427802027710118705 2105880 17240328 | 3 | 3 | V790 | 2022-07-21 11:34 | 535 | 271 | 0.5 | 2230.2333 | 2 | 4 | 2237 | 2228 | 2230 | 2225 | 2237 | 2242 | 2237 | 2242 | 2230 | 2258 | 2242 | 2255 | 2251 | 2261 | 2275 | 2289 | 2325 | 2396 | 2489 | 2581 | 2659 | 2728 | 2775 | 2798 | 2826 | 2862 | 2870 | 2890 | 2902 | 2915 | 2933 | 2956 | 2943 | 2956 | 2959 | 2963 | 2960 | 2993 | 2996 | 2992 | 3007 | 3003 | 3000 | 2997 | 2998
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 7842e1f8-1209-ed11-85f4-5cf370a20867 | 328b0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.27 | 56.35 | 1.32 | 874 | [20 empty] | 010081427802027710118705 2105880 17240328 | 4 | 3 | V790 | 2022-07-21 11:34 | 513 | 267 | -4.1 | 2200.7667 | 2 | 4 | 2204 | 2184 | 2178 | 2191 | 2189 | 2178 | 2198 | 2193 | 2185 | 2180 | 2207 | 2198 | 2195 | 2207 | 2217 | 2232 | 2248 | 2324 | 2414 | 2483 | 2571 | 2614 | 2675 | 2692 | 2715 | 2739 | 2751 | 2792 | 2800 | 2806 | 2807 | 2833 | 2849 | 2852 | 2855 | 2859 | 2849 | 2863 | 2875 | 2870 | 2875 | 2897 | 2901 | 2909 | 2881
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 7e42e1f8-1209-ed11-85f4-5cf370a20867 | bdaadb1d-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 2610 (Informational, Delays in sample processing), 1020 (Informational, Peak Not Detected) |  |  | 1.18 | [21 empty] | 010081427802027710118705 2105880 17240328 | 5 | 3 | V790 | 2022-07-21 11:34 | 499 | 266 | 12.6579 | 1361.7431 | 5 | 27 | 1482 | 1472 | 1488 | 1488 | 1492 | 1481 | 1496 | 1494 | 1465 | 1473 | 1484 | 1482 | 1494 | 1487 | 1479 | 1509 | 1514 | 1558 | 1587 | 1621 | 1654 | 1672 | 1707 | 1691 | 1719 | 1693 | 1711 | 1738 | 1716 | 1744 | 1733 | 1732 | 1728 | 1723 | 1742 | 1741 | 1749 | 1748 | 1757 | 1743 | 1705 | 1727 | 1732 | 1735 | 1758
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 8442e1f8-1209-ed11-85f4-5cf370a20867 | 29c954cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.18 | 50.9 | 1.36 | 699 | [20 empty] | 010081427802027710118705 2105880 17240328 | 6 | 3 | V790 | 2022-07-21 11:34 | 498 | 265 | 2 | 2167.2667 | 2 | 4 | 2170 | 2164 | 2172 | 2177 | 2190 | 2170 | 2185 | 2172 | 2194 | 2165 | 2171 | 2189 | 2195 | 2204 | 2200 | 2263 | 2294 | 2358 | 2431 | 2523 | 2586 | 2645 | 2679 | 2710 | 2753 | 2757 | 2760 | 2760 | 2795 | 2815 | 2819 | 2820 | 2848 | 2869 | 2872 | 2907 | 2892 | 2922 | 2939 | 2941 | 2943 | 2954 | 2939 | 2960 | 2961
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 8a42e1f8-1209-ed11-85f4-5cf370a20867 | bfc854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.18 | [21 empty] | 010081427802027710118705 2105880 17240328 | 7 | 3 | V790 | 2022-07-21 11:34 | 494 | 262 | 16.7034 | 1997.2158 | 5 | 27 | 2132 | 2131 | 2131 | 2144 | 2144 | 2168 | 2165 | 2158 | 2155 | 2164 | 2154 | 2158 | 2168 | 2159 | 2177 | 2196 | 2220 | 2235 | 2294 | 2345 | 2388 | 2410 | 2409 | 2429 | 2448 | 2455 | 2478 | 2489 | 2499 | 2512 | 2511 | 2510 | 2518 | 2514 | 2530 | 2533 | 2541 | 2547 | 2550 | 2551 | 2546 | 2545 | 2559 | 2560 | 2541
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 5a7b815a-1309-ed11-85f4-5cf370a20867 | 503c8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.17 | [21 empty] | 010081427802027710118705 2105881 17240328 | 7 | 2 | V834 | 2022-07-21 11:37 | 563 | 270 | 16.2757 | 2127.102 | 5 | 27 | 2288 | 2264 | 2283 | 2274 | 2279 | 2301 | 2273 | 2295 | 2297 | 2281 | 2294 | 2285 | 2284 | 2284 | 2289 | 2304 | 2336 | 2360 | 2410 | 2449 | 2502 | 2519 | 2555 | 2555 | 2596 | 2578 | 2583 | 2595 | 2601 | 2615 | 2619 | 2626 | 2631 | 2624 | 2631 | 2639 | 2633 | 2639 | 2648 | 2642 | 2652 | 2641 | 2657 | 2665 | 2665
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 607b815a-1309-ed11-85f4-5cf370a20867 | 22597999-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.16 | [21 empty] | 010081427802027710118705 2105881 17240328 | 8 | 2 | V834 | 2022-07-21 11:37 | 569 | 268 | 14.8879 | 2018.9233 | 5 | 27 | 2156 | 2167 | 2165 | 2168 | 2159 | 2170 | 2172 | 2152 | 2169 | 2156 | 2161 | 2180 | 2165 | 2180 | 2170 | 2185 | 2177 | 2238 | 2280 | 2317 | 2359 | 2370 | 2404 | 2421 | 2432 | 2434 | 2457 | 2446 | 2473 | 2463 | 2480 | 2488 | 2465 | 2469 | 2486 | 2483 | 2486 | 2476 | 2489 | 2502 | 2514 | 2495 | 2497 | 2499 | 2506
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 667b815a-1309-ed11-85f4-5cf370a20867 | 6e3e8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.15 | [21 empty] | 010081427802027710118705 2105881 17240328 | 9 | 2 | V834 | 2022-07-21 11:37 | 664 | 265 | 14.7126 | 2116.7802 | 5 | 27 | 2252 | 2267 | 2258 | 2262 | 2253 | 2260 | 2266 | 2269 | 2245 | 2262 | 2253 | 2261 | 2258 | 2272 | 2280 | 2275 | 2305 | 2332 | 2367 | 2415 | 2448 | 2476 | 2494 | 2516 | 2530 | 2520 | 2534 | 2532 | 2554 | 2559 | 2569 | 2564 | 2562 | 2578 | 2560 | 2570 | 2569 | 2578 | 2581 | 2581 | 2581 | 2581 | 2604 | 2602 | 2591
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 6c7b815a-1309-ed11-85f4-5cf370a20867 | 103e8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105881 17240328 | 10 | 2 | V834 | 2022-07-21 11:37 | 659 | 280 | 21.4891 | 2038.9826 | 5 | 27 | 2248 | 2265 | 2247 | 2237 | 2258 | 2245 | 2242 | 2258 | 2248 | 2254 | 2257 | 2251 | 2249 | 2251 | 2262 | 2275 | 2303 | 2343 | 2395 | 2464 | 2545 | 2563 | 2586 | 2631 | 2624 | 2642 | 2653 | 2668 | 2680 | 2676 | 2695 | 2688 | 2695 | 2701 | 2708 | 2702 | 2698 | 2706 | 2707 | 2718 | 2703 | 2714 | 2738 | 2717 | 2730
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 727b815a-1309-ed11-85f4-5cf370a20867 | b23d8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.22 | [21 empty] | 010081427802027710118705 2105881 17240328 | 11 | 2 | V834 | 2022-07-21 11:37 | 627 | 270 | 20.9273 | 2048.0593 | 5 | 27 | 2265 | 2253 | 2260 | 2243 | 2252 | 2265 | 2250 | 2246 | 2256 | 2266 | 2255 | 2266 | 2243 | 2270 | 2261 | 2267 | 2287 | 2347 | 2394 | 2470 | 2526 | 2554 | 2588 | 2598 | 2636 | 2650 | 2652 | 2670 | 2675 | 2665 | 2664 | 2697 | 2691 | 2703 | 2717 | 2707 | 2704 | 2694 | 2719 | 2709 | 2715 | 2696 | 2723 | 2714 | 2754
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 787b815a-1309-ed11-85f4-5cf370a20867 | f0fe3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.24 | [21 empty] | 010081427802027710118705 2105881 17240328 | 4 | 2 | V834 | 2022-07-21 11:37 | 633 | 295 | 24.0296 | 2104.43 | 5 | 27 | 2311 | 2329 | 2332 | 2312 | 2325 | 2329 | 2331 | 2342 | 2345 | 2344 | 2338 | 2345 | 2348 | 2360 | 2349 | 2386 | 2408 | 2452 | 2521 | 2582 | 2652 | 2705 | 2725 | 2722 | 2767 | 2784 | 2782 | 2791 | 2803 | 2816 | 2838 | 2844 | 2837 | 2832 | 2847 | 2840 | 2854 | 2835 | 2863 | 2865 | 2860 | 2875 | 2860 | 2878 | 2883
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 7e7b815a-1309-ed11-85f4-5cf370a20867 | 543d8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.27 | [21 empty] | 010081427802027710118705 2105881 17240328 | 5 | 2 | V834 | 2022-07-21 11:37 | 596 | 273 | 27.6573 | 2095.9352 | 5 | 27 | 2334 | 2333 | 2345 | 2340 | 2351 | 2365 | 2368 | 2369 | 2364 | 2374 | 2372 | 2387 | 2360 | 2391 | 2385 | 2398 | 2435 | 2499 | 2566 | 2638 | 2716 | 2770 | 2821 | 2823 | 2865 | 2881 | 2880 | 2902 | 2911 | 2907 | 2896 | 2923 | 2927 | 2948 | 2952 | 2948 | 2941 | 2949 | 2948 | 2956 | 2974 | 2953 | 2971 | 2972 | 3003
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 847b815a-1309-ed11-85f4-5cf370a20867 | 953c8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.23 | 58.23 | 1.25 | 315 | [20 empty] | 010081427802027710118705 2105881 17240328 | 6 | 2 | V834 | 2022-07-21 11:37 | 617 | 276 | 6.4 | 2303.3333 | 2 | 4 | 2296 | 2335 | 2323 | 2333 | 2321 | 2335 | 2329 | 2334 | 2325 | 2344 | 2327 | 2342 | 2354 | 2338 | 2340 | 2371 | 2389 | 2446 | 2541 | 2612 | 2671 | 2715 | 2759 | 2762 | 2811 | 2801 | 2821 | 2826 | 2845 | 2866 | 2855 | 2883 | 2875 | 2882 | 2859 | 2884 | 2889 | 2873 | 2883 | 2888 | 2897 | 2905 | 2912 | 2906 | 2904
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | b3de19db-1309-ed11-85f4-5cf370a20867 | 3a5f6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.74 | 58.9 | 1.27 | 541 | [20 empty] | 010081427802027710118705 2105881 17240328 | 3 | 2 | V834 | 2022-07-21 11:41 | 592 | 269 | 2.5 | 2335.5 | 2 | 4 | 2338 | 2335 | 2342 | 2348 | 2355 | 2346 | 2357 | 2342 | 2358 | 2354 | 2343 | 2350 | 2367 | 2361 | 2356 | 2374 | 2414 | 2468 | 2540 | 2624 | 2709 | 2763 | 2797 | 2834 | 2853 | 2880 | 2893 | 2910 | 2911 | 2929 | 2936 | 2943 | 2955 | 2945 | 2970 | 2965 | 2967 | 2992 | 2973 | 2979 | 3001 | 3000 | 2997 | 2986 | 2987
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | b9de19db-1309-ed11-85f4-5cf370a20867 | a75e6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.19 | [21 empty] | 010081427802027710118705 2105882 17240328 | 10 | 1 | V639 | 2022-07-21 11:41 | 614 | 260 | 19.8022 | 2195.6174 | 5 | 27 | 2358 | 2387 | 2380 | 2376 | 2379 | 2368 | 2391 | 2397 | 2393 | 2400 | 2396 | 2395 | 2387 | 2402 | 2392 | 2420 | 2445 | 2482 | 2539 | 2597 | 2631 | 2679 | 2702 | 2722 | 2738 | 2746 | 2782 | 2770 | 2781 | 2789 | 2805 | 2813 | 2814 | 2808 | 2811 | 2823 | 2840 | 2826 | 2833 | 2847 | 2837 | 2846 | 2829 | 2859 | 2843
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | bfde19db-1309-ed11-85f4-5cf370a20867 | 185e6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.26 | 56.9 | 1.29 | 510 | [20 empty] | 010081427802027710118705 2105882 17240328 | 11 | 1 | V639 | 2022-07-21 11:41 | 588 | 272 | 3.9 | 2333.5 | 2 | 4 | 2326 | 2352 | 2347 | 2356 | 2344 | 2347 | 2348 | 2352 | 2350 | 2363 | 2373 | 2360 | 2372 | 2370 | 2372 | 2396 | 2447 | 2509 | 2586 | 2674 | 2747 | 2800 | 2825 | 2856 | 2890 | 2886 | 2924 | 2914 | 2934 | 2948 | 2960 | 2966 | 2972 | 2989 | 2980 | 3004 | 2994 | 3012 | 3005 | 2996 | 3010 | 3009 | 3009 | 3016 | 3024
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | c5de19db-1309-ed11-85f4-5cf370a20867 | bffe3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105882 17240328 | 12 | 1 | V639 | 2022-07-21 11:41 | 672 | 260 | 21.7893 | 2250.3186 | 5 | 27 | 2445 | 2458 | 2446 | 2455 | 2461 | 2454 | 2458 | 2449 | 2471 | 2450 | 2471 | 2466 | 2487 | 2486 | 2481 | 2496 | 2524 | 2576 | 2635 | 2671 | 2748 | 2767 | 2799 | 2831 | 2847 | 2859 | 2882 | 2887 | 2903 | 2891 | 2922 | 2908 | 2909 | 2935 | 2911 | 2931 | 2935 | 2937 | 2930 | 2957 | 2950 | 2973 | 2952 | 2966 | 2962
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | cbde19db-1309-ed11-85f4-5cf370a20867 | 32fe3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.23 | [21 empty] | 010081427802027710118705 2105881 17240328 | 1 | 2 | V834 | 2022-07-21 11:41 | 679 | 291 | 23.8111 | 2216.2577 | 5 | 27 | 2440 | 2434 | 2446 | 2445 | 2444 | 2451 | 2440 | 2450 | 2462 | 2446 | 2451 | 2454 | 2448 | 2467 | 2472 | 2481 | 2519 | 2551 | 2622 | 2675 | 2734 | 2783 | 2834 | 2856 | 2892 | 2886 | 2916 | 2908 | 2917 | 2932 | 2953 | 2940 | 2963 | 2959 | 2967 | 2967 | 2972 | 2974 | 2972 | 2973 | 2974 | 3005 | 2994 | 3001 | 3003
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | d1de19db-1309-ed11-85f4-5cf370a20867 | c65f6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.22 | [21 empty] | 010081427802027710118705 2105881 17240328 | 2 | 2 | V834 | 2022-07-21 11:41 | 600 | 276 | 24.1291 | 2174.6047 | 5 | 27 | 2370 | 2387 | 2386 | 2384 | 2387 | 2405 | 2398 | 2410 | 2419 | 2410 | 2423 | 2423 | 2407 | 2447 | 2417 | 2453 | 2473 | 2524 | 2600 | 2659 | 2722 | 2773 | 2788 | 2826 | 2830 | 2840 | 2855 | 2869 | 2869 | 2893 | 2883 | 2888 | 2887 | 2906 | 2900 | 2912 | 2911 | 2929 | 2923 | 2921 | 2918 | 2940 | 2928 | 2936 | 2934
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | d7de19db-1309-ed11-85f4-5cf370a20867 | 674775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105882 17240328 | 8 | 1 | V639 | 2022-07-21 11:41 | 620 | 261 | 21.2957 | 2243.913 | 5 | 27 | 2434 | 2439 | 2439 | 2439 | 2441 | 2442 | 2460 | 2449 | 2438 | 2452 | 2461 | 2433 | 2458 | 2467 | 2476 | 2497 | 2522 | 2561 | 2620 | 2690 | 2716 | 2748 | 2786 | 2821 | 2809 | 2854 | 2838 | 2858 | 2860 | 2869 | 2890 | 2873 | 2900 | 2894 | 2893 | 2891 | 2918 | 2918 | 2916 | 2932 | 2911 | 2933 | 2929 | 2933 | 2937
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | ddde19db-1309-ed11-85f4-5cf370a20867 | 095f6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105882 17240328 | 9 | 1 | V639 | 2022-07-21 11:41 | 661 | 258 | 21.6571 | 2230.5644 | 5 | 27 | 2398 | 2431 | 2422 | 2447 | 2436 | 2430 | 2434 | 2429 | 2424 | 2434 | 2442 | 2462 | 2447 | 2457 | 2461 | 2488 | 2513 | 2547 | 2635 | 2688 | 2731 | 2752 | 2788 | 2787 | 2808 | 2829 | 2844 | 2850 | 2837 | 2852 | 2857 | 2868 | 2870 | 2874 | 2858 | 2894 | 2875 | 2883 | 2901 | 2888 | 2899 | 2892 | 2910 | 2917 | 2908
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 65fdae7a-1409-ed11-85f4-5cf370a20867 | 084775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105882 17240328 | 4 | 1 | V639 | 2022-07-21 11:45 | 684 | 263 | 21.7674 | 2166.2783 | 5 | 27 | 2362 | 2366 | 2361 | 2378 | 2381 | 2387 | 2373 | 2369 | 2372 | 2378 | 2380 | 2379 | 2399 | 2386 | 2396 | 2408 | 2434 | 2485 | 2536 | 2624 | 2653 | 2690 | 2727 | 2743 | 2768 | 2771 | 2799 | 2795 | 2795 | 2826 | 2817 | 2820 | 2836 | 2844 | 2832 | 2851 | 2847 | 2862 | 2850 | 2859 | 2854 | 2866 | 2841 | 2861 | 2863
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 6bfdae7a-1409-ed11-85f4-5cf370a20867 | d74675b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.21 | [21 empty] | 010081427802027710118705 2105882 17240328 | 5 | 1 | V639 | 2022-07-21 11:45 | 703 | 280 | 23.3328 | 2329.7447 | 5 | 27 | 2542 | 2523 | 2524 | 2540 | 2549 | 2555 | 2553 | 2541 | 2548 | 2557 | 2548 | 2564 | 2564 | 2578 | 2591 | 2599 | 2644 | 2691 | 2746 | 2814 | 2859 | 2876 | 2935 | 2944 | 2962 | 2964 | 2988 | 2996 | 3000 | 3005 | 3025 | 3044 | 3030 | 3056 | 3061 | 3044 | 3058 | 3046 | 3061 | 3059 | 3065 | 3051 | 3076 | 3087 | 3078
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 71fdae7a-1409-ed11-85f4-5cf370a20867 | 494675b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 15.89 | 59.99 | 1.29 | 667 | [20 empty] | 010081427802027710118705 2105882 17240328 | 6 | 1 | V639 | 2022-07-21 11:45 | 749 | 264 | 1.2 | 2464.1333 | 2 | 4 | 2466 | 2457 | 2463 | 2480 | 2474 | 2470 | 2487 | 2473 | 2482 | 2500 | 2471 | 2489 | 2480 | 2481 | 2501 | 2514 | 2575 | 2635 | 2737 | 2812 | 2878 | 2946 | 2994 | 3051 | 3046 | 3059 | 3087 | 3081 | 3111 | 3103 | 3129 | 3118 | 3142 | 3146 | 3151 | 3142 | 3144 | 3160 | 3171 | 3158 | 3179 | 3180 | 3176 | 3196 | 3185
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 77fdae7a-1409-ed11-85f4-5cf370a20867 | f44775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105882 17240328 | 7 | 1 | V639 | 2022-07-21 11:45 | 611 | 265 | 21.8399 | 2179.8047 | 5 | 27 | 2375 | 2373 | 2373 | 2376 | 2389 | 2388 | 2394 | 2379 | 2382 | 2389 | 2406 | 2391 | 2407 | 2402 | 2417 | 2435 | 2471 | 2505 | 2564 | 2617 | 2675 | 2707 | 2731 | 2765 | 2765 | 2790 | 2804 | 2814 | 2809 | 2814 | 2821 | 2831 | 2830 | 2830 | 2839 | 2851 | 2851 | 2874 | 2840 | 2859 | 2869 | 2874 | 2874 | 2860 | 2863
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | f47fc1ae-1609-ed11-85f4-5cf370a20867 | 536b47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.19 | [21 empty] | 010081427802027710118705 2105882 17240328 | 1 | 1 | V639 | 2022-07-21 12:01 | 725 | 269 | 21.8492 | 2389.1866 | 5 | 27 | 2573 | 2580 | 2583 | 2589 | 2580 | 2599 | 2593 | 2591 | 2606 | 2603 | 2605 | 2606 | 2621 | 2624 | 2622 | 2643 | 2657 | 2723 | 2802 | 2835 | 2873 | 2891 | 2946 | 2949 | 2992 | 3019 | 3006 | 3014 | 3030 | 3031 | 3043 | 3044 | 3060 | 3061 | 3054 | 3061 | 3070 | 3072 | 3066 | 3082 | 3115 | 3100 | 3074 | 3082 | 3088
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | fa7fc1ae-1609-ed11-85f4-5cf370a20867 | f46a47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.2 | [21 empty] | 010081427802027710118705 2105882 17240328 | 2 | 1 | V639 | 2022-07-21 12:01 | 719 | 260 | 23.9073 | 2499.5613 | 5 | 27 | 2709 | 2709 | 2714 | 2715 | 2729 | 2714 | 2737 | 2717 | 2729 | 2735 | 2728 | 2739 | 2743 | 2746 | 2761 | 2778 | 2815 | 2850 | 2924 | 2996 | 3038 | 3085 | 3113 | 3137 | 3144 | 3145 | 3174 | 3183 | 3209 | 3205 | 3215 | 3228 | 3231 | 3238 | 3243 | 3243 | 3249 | 3258 | 3268 | 3256 | 3260 | 3267 | 3268 | 3271 | 3271
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | 0080c1ae-1609-ed11-85f4-5cf370a20867 | b36b47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Indeterminate | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16 | 50.4 | 1.21 | 402 | [20 empty] | 010081427802027710118705 2105882 17240328 | 3 | 1 | V639 | 2022-07-21 12:01 | 714 | 262 | 3.7 | 2589.5667 | 2 | 4 | 2588 | 2601 | 2597 | 2613 | 2600 | 2613 | 2624 | 2631 | 2620 | 2618 | 2638 | 2625 | 2631 | 2630 | 2641 | 2660 | 2685 | 2746 | 2826 | 2895 | 2946 | 2980 | 2996 | 3037 | 3046 | 3054 | 3070 | 3072 | 3113 | 3100 | 3116 | 3120 | 3126 | 3116 | 3136 | 3132 | 3120 | 3136 | 3150 | 3141 | 3157 | 3165 | 3160 | 3150 | 3160
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | a79dad68-1f09-ed11-85f4-5cf370a20867 | b16dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | ad9dad68-1f09-ed11-85f4-5cf370a20867 | b56dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | b39dad68-1f09-ed11-85f4-5cf370a20867 | b96dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | b99dad68-1f09-ed11-85f4-5cf370a20867 | bd6dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | bf9dad68-1f09-ed11-85f4-5cf370a20867 | c16dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | c59dad68-1f09-ed11-85f4-5cf370a20867 | c66dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | cb9dad68-1f09-ed11-85f4-5cf370a20867 | 010bca23-2809-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.17 | [21 empty] | 010081427802027710118705 2105883 17240328 | 11 | 4 | V1169 | 2022-07-21 13:03 | 870 | 289 | 20.3816 | 2650.1897 | 5 | 27 | 2808 | 2838 | 2836 | 2830 | 2848 | 2850 | 2853 | 2859 | 2855 | 2855 | 2857 | 2864 | 2847 | 2864 | 2860 | 2870 | 2904 | 2930 | 2978 | 3024 | 3101 | 3139 | 3190 | 3199 | 3210 | 3244 | 3250 | 3273 | 3267 | 3268 | 3288 | 3287 | 3283 | 3295 | 3293 | 3304 | 3313 | 3315 | 3314 | 3322 | 3319 | 3341 | 3335 | 3349 | 3348
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | d29dad68-1f09-ed11-85f4-5cf370a20867 | a30aca23-2809-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.18 | [21 empty] | 010081427802027710118705 2105883 17240328 | 12 | 4 | V1169 | 2022-07-21 13:03 | 839 | 276 | 21.0879 | 2618.4016 | 5 | 27 | 2824 | 2818 | 2808 | 2800 | 2809 | 2821 | 2829 | 2815 | 2831 | 2832 | 2825 | 2832 | 2841 | 2844 | 2845 | 2847 | 2892 | 2926 | 2978 | 3031 | 3077 | 3120 | 3155 | 3162 | 3205 | 3228 | 3233 | 3246 | 3261 | 3248 | 3261 | 3279 | 3261 | 3277 | 3299 | 3293 | 3295 | 3292 | 3294 | 3304 | 3327 | 3333 | 3317 | 3332 | 3335
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 1f42d920-2009-ed11-85f4-5cf370a20867 | 13e5bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.45 | 54.87 | 1.25 | 703 | [20 empty] | 010081427802027710118705 2105883 17240328 | 5 | 4 | V1169 | 2022-07-21 13:09 | 716 | 286 | 0.1 | 2777.1 | 2 | 4 | 2777 | 2775 | 2776 | 2784 | 2770 | 2785 | 2797 | 2806 | 2788 | 2816 | 2796 | 2802 | 2805 | 2806 | 2821 | 2838 | 2859 | 2926 | 3001 | 3093 | 3161 | 3218 | 3261 | 3288 | 3305 | 3334 | 3341 | 3364 | 3377 | 3379 | 3391 | 3396 | 3395 | 3421 | 3421 | 3442 | 3447 | 3462 | 3464 | 3450 | 3466 | 3485 | 3473 | 3479 | 3490
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 2642d920-2009-ed11-85f4-5cf370a20867 | e1e4bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.28 | 52.81 | 1.23 | 438 | [20 empty] | 010081427802027710118705 2105883 17240328 | 6 | 4 | V1169 | 2022-07-21 13:09 | 732 | 276 | 5 | 2706.5333 | 2 | 4 | 2710 | 2715 | 2723 | 2724 | 2737 | 2733 | 2730 | 2740 | 2741 | 2728 | 2746 | 2741 | 2759 | 2764 | 2771 | 2783 | 2826 | 2888 | 2984 | 3043 | 3124 | 3167 | 3181 | 3211 | 3240 | 3254 | 3274 | 3278 | 3292 | 3303 | 3308 | 3323 | 3323 | 3322 | 3334 | 3333 | 3359 | 3339 | 3357 | 3352 | 3353 | 3340 | 3378 | 3385 | 3361
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 2c42d920-2009-ed11-85f4-5cf370a20867 | 8fe6bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.18 | [21 empty] | 010081427802027710118705 2105883 17240328 | 7 | 4 | V1169 | 2022-07-21 13:09 | 837 | 291 | 20.8789 | 2564.3731 | 5 | 27 | 2744 | 2759 | 2745 | 2764 | 2759 | 2770 | 2770 | 2771 | 2754 | 2772 | 2773 | 2780 | 2776 | 2783 | 2785 | 2814 | 2820 | 2853 | 2918 | 2968 | 3037 | 3061 | 3095 | 3131 | 3142 | 3158 | 3159 | 3194 | 3178 | 3182 | 3191 | 3192 | 3232 | 3209 | 3225 | 3224 | 3226 | 3232 | 3248 | 3242 | 3252 | 3242 | 3267 | 3250 | 3254
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 3242d920-2009-ed11-85f4-5cf370a20867 | 30e6bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1020 (Informational, Peak Not Detected) |  |  | 1.19 | [21 empty] | 010081427802027710118705 2105883 17240328 | 8 | 4 | V1169 | 2022-07-21 13:09 | 768 | 291 | 21.1269 | 2541.8308 | 5 | 27 | 2733 | 2728 | 2743 | 2723 | 2732 | 2749 | 2745 | 2746 | 2738 | 2747 | 2741 | 2753 | 2766 | 2769 | 2781 | 2779 | 2814 | 2863 | 2900 | 2960 | 3021 | 3065 | 3079 | 3089 | 3107 | 3137 | 3148 | 3162 | 3172 | 3162 | 3188 | 3195 | 3183 | 3199 | 3212 | 3226 | 3233 | 3216 | 3228 | 3225 | 3227 | 3239 | 3249 | 3256 | 3253
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 3842d920-2009-ed11-85f4-5cf370a20867 | 00e6bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.78 | 57.59 | 1.25 | 502 | [20 empty] | 010081427802027710118705 2105883 17240328 | 9 | 4 | V1169 | 2022-07-21 13:09 | 775 | 295 | 4.4 | 2755.0667 | 2 | 4 | 2763 | 2754 | 2765 | 2775 | 2783 | 2787 | 2784 | 2774 | 2771 | 2767 | 2787 | 2775 | 2796 | 2790 | 2803 | 2806 | 2817 | 2879 | 2956 | 3047 | 3115 | 3170 | 3207 | 3260 | 3271 | 3294 | 3307 | 3334 | 3337 | 3339 | 3365 | 3365 | 3376 | 3386 | 3402 | 3388 | 3398 | 3418 | 3417 | 3443 | 3439 | 3430 | 3430 | 3452 | 3464
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 3e42d920-2009-ed11-85f4-5cf370a20867 | 73e5bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Positive | TargetNotAmplified | 1004 (Informational, Fixed Baseline Used), 1006 (Informational, Peak Location Failed), 1005 (Informational, End Point Fluorescence Met Failed), 1025 (Informational, Below EPR Threshold) | 16.61 | 57.38 | 1.23 | 706 | [20 empty] | 010081427802027710118705 2105883 17240328 | 10 | 4 | V1169 | 2022-07-21 13:09 | 849 | 286 | -0.7 | 2830.2333 | 2 | 4 | 2843 | 2826 | 2819 | 2815 | 2834 | 2845 | 2848 | 2849 | 2835 | 2843 | 2840 | 2844 | 2848 | 2848 | 2865 | 2870 | 2907 | 2945 | 3039 | 3126 | 3204 | 3242 | 3296 | 3328 | 3325 | 3365 | 3387 | 3386 | 3411 | 3412 | 3423 | 3441 | 3443 | 3449 | 3441 | 3473 | 3468 | 3467 | 3472 | 3488 | 3490 | 3505 | 3495 | 3501 | 3509
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 4442d920-2009-ed11-85f4-5cf370a20867 | 803d7504-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:09 |  |  |  |  | 0
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | b7477ca2-3609-ed11-85f4-5cf370a20867 | f2c2a3da-3609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:50 |  |  |  |  | 0
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | f5c2a3da-3609-ed11-85f4-5cf370a20867 | 8db7eb13-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:51 |  |  |  |  | 0
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 90b7eb13-3709-ed11-85f4-5cf370a20867 | b7bbce4d-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:53 |  |  |  |  | 0
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | f2bbce4d-3709-ed11-85f4-5cf370a20867 | 0b592f87-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:54 |  |  |  |  | 0
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 0e592f87-3709-ed11-85f4-5cf370a20867 | 4e3e36c2-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | SPC1 | Urine | Aborted | TargetAborted | 7162 (Warning, Quantity not sufficient), 1000 (Error, User Aborted Test) |  |  | 0 | [25 empty] | 2022-07-21 15:56 |  |  |  |  | 0

## Red_625_660
Sample ID | Patient ID | Assay Name | Result Code | Assay Version | Test Guid | Target Result Guid | Test Comment | Replicate Number | Extraction Number | Elution Number | Primer/Probe | Calibration Coefficient | Target Name | Target Setting Specimen Type | Replicate Result | Localized Result | Flags | Ct | Max Peak Height | EPR | End Point Fluorescence | Log Conc. | Log Conc. with Unit Conv. | Conc. | Conc. with Unit Conv. | Log Conc. with Dilution Factor | Log Conc. with Dilution Factor Unit Conv. | Log Conc. Precision | Log Conc. Conv. Precision | Conc. with Dilution Factor | Conc. with Dilution Factor Unit Conv. | Conc. Precision | Conc. Conv. Precision | Log Conc. with Quant Adj. | Log Conc. with Quant Adj. Unit Conv. | Conc. with Quant Adj. | Conc. with Quant Adj. Unit Conv. | Log Intermed. Conc. | Log Intermed. Conc. with Unit Conv. | Intermed. Conc. | Intermed. Conc. with Unit Conv. | Cartridge Barcode | Cartridge Lane | XPCR Module Index | XPCR Module Serial | Start Time | Blank Reading | Dark Reading | Baseline Slope | Baseline Y Intercept | Baseline First Cycle | Baseline Last Cycle | Readings 1 | Readings 2 | Readings 3 | Readings 4 | Readings 5 | Readings 6 | Readings 7 | Readings 8 | Readings 9 | Readings 10 | Readings 11 | Readings 12 | Readings 13 | Readings 14 | Readings 15 | Readings 16 | Readings 17 | Readings 18 | Readings 19 | Readings 20 | Readings 21 | Readings 22 | Readings 23 | Readings 24 | Readings 25 | Readings 26 | Readings 27 | Readings 28 | Readings 29 | Readings 30 | Readings 31 | Readings 32 | Readings 33 | Readings 34 | Readings 35 | Readings 36 | Readings 37 | Readings 38 | Readings 39 | Readings 40 | Readings 41 | Readings 42 | Readings 43 | Readings 44 | Readings 45 | Readings 46 | Readings 47 | Readings 48 | Readings 49 | Readings 50 | Readings 51 | Readings 52 | Readings 53 | Readings 54 | Readings 55 | Readings 56 | Readings 57 | Readings 58 | Readings 59 | Readings 60 | Readings 61 | Readings 62 | Readings 63 | Readings 64 | Readings 65 | Readings 66 | Readings 67 | Readings 68 | Readings 69 | Readings 70 | Readings 71 | Readings 72 | Readings 73 | Readings 74 | Readings 75 | Readings 76 | Readings 77 | Readings 78 | Readings 79 | Readings 80 | Readings 81 | Readings 82 | Readings 83 | Readings 84 | Readings 85 | Readings 86 | Readings 87 | Readings 88 | Readings 89 | Readings 90 | Readings 91 | Readings 92 | Readings 93 | Readings 94 | Readings 95 | Readings 96 | Readings 97 | Readings 98 | Readings 99 | Readings 100 | Readings 101 | Readings 102 | Readings 103 | Readings 104 | Readings 105 | Readings 106 | Readings 107 | Readings 108 | Readings 109 | Readings 110 | Readings 111 | Readings 112 | Readings 113 | Readings 114 | Readings 115 | Readings 116 | Readings 117 | Readings 118 | Readings 119 | Readings 120 | Readings 121 | Readings 122 | Readings 123 | Readings 124 | Readings 125 | Readings 126 | Readings 127 | Readings 128 | Readings 129 | Readings 130 | Readings 131 | Readings 132 | Readings 133 | Readings 134 | Readings 135 | Readings 136 | Readings 137 | Readings 138 | Readings 139 | Readings 140 | Readings 141 | Readings 142 | Readings 143 | Readings 144 | Readings 145 | Readings 146 | Readings 147 | Readings 148 | Readings 149 | Readings 150 | Readings 151 | Readings 152 | Readings 153 | Readings 154 | Readings 155 | Readings 156 | Readings 157 | Readings 158 | Readings 159 | Readings 160 | Readings 161 | Readings 162 | Readings 163 | Readings 164 | Readings 165 | Readings 166 | Readings 167 | Readings 168 | Readings 169 | Readings 170 | Readings 171 | Readings 172 | Readings 173 | Readings 174 | Readings 175 | Readings 176 | Readings 177 | Readings 178 | Readings 179 | Readings 180 | Readings 181 | Readings 182 | Readings 183 | Readings 184 | Readings 185 | Readings 186 | Readings 187 | Readings 188 | Readings 189 | Readings 190 | Readings 191 | Readings 192 | Readings 193 | Readings 194 | Readings 195 | Readings 196 | Readings 197 | Readings 198 | Readings 199 | Readings 200 | Readings 201 | Readings 202 | Readings 203 | Readings 204 | Readings 205 | Readings 206 | Readings 207 | Readings 208 | Readings 209 | Readings 210 | Readings 211 | Readings 212 | Readings 213 | Readings 214 | Readings 215 | Readings 216 | Readings 217 | Readings 218 | Readings 219 | Readings 220 | Readings 221 | Readings 222 | Readings 223 | Readings 224 | Readings 225 | Readings 226 | Readings 227 | Readings 228 | Readings 229 | Readings 230 | Readings 231 | Readings 232 | Readings 233 | Readings 234 | Readings 235 | Readings 236 | Readings 237 | Readings 238 | Readings 239 | Readings 240 | Readings 241 | Readings 242 | Readings 243 | Readings 244 | Readings 245 | Readings 246 | Readings 247 | Readings 248 | Readings 249 | Readings 250 | Readings 251 | Readings 252 | Readings 253 | Readings 254 | Readings 255

## Green_470_510
Sample ID | Patient ID | Assay Name | Result Code | Assay Version | Test Guid | Target Result Guid | Test Comment | Replicate Number | Extraction Number | Elution Number | Primer/Probe | Calibration Coefficient | Target Name | Target Setting Specimen Type | Replicate Result | Localized Result | Flags | Ct | Max Peak Height | EPR | End Point Fluorescence | Log Conc. | Log Conc. with Unit Conv. | Conc. | Conc. with Unit Conv. | Log Conc. with Dilution Factor | Log Conc. with Dilution Factor Unit Conv. | Log Conc. Precision | Log Conc. Conv. Precision | Conc. with Dilution Factor | Conc. with Dilution Factor Unit Conv. | Conc. Precision | Conc. Conv. Precision | Log Conc. with Quant Adj. | Log Conc. with Quant Adj. Unit Conv. | Conc. with Quant Adj. | Conc. with Quant Adj. Unit Conv. | Log Intermed. Conc. | Log Intermed. Conc. with Unit Conv. | Intermed. Conc. | Intermed. Conc. with Unit Conv. | Cartridge Barcode | Cartridge Lane | XPCR Module Index | XPCR Module Serial | Start Time | Blank Reading | Dark Reading | Baseline Slope | Baseline Y Intercept | Baseline First Cycle | Baseline Last Cycle | Readings 1 | Readings 2 | Readings 3 | Readings 4 | Readings 5 | Readings 6 | Readings 7 | Readings 8 | Readings 9 | Readings 10 | Readings 11 | Readings 12 | Readings 13 | Readings 14 | Readings 15 | Readings 16 | Readings 17 | Readings 18 | Readings 19 | Readings 20 | Readings 21 | Readings 22 | Readings 23 | Readings 24 | Readings 25 | Readings 26 | Readings 27 | Readings 28 | Readings 29 | Readings 30 | Readings 31 | Readings 32 | Readings 33 | Readings 34 | Readings 35 | Readings 36 | Readings 37 | Readings 38 | Readings 39 | Readings 40 | Readings 41 | Readings 42 | Readings 43 | Readings 44 | Readings 45 | Readings 46 | Readings 47 | Readings 48 | Readings 49 | Readings 50 | Readings 51 | Readings 52 | Readings 53 | Readings 54 | Readings 55 | Readings 56 | Readings 57 | Readings 58 | Readings 59 | Readings 60 | Readings 61 | Readings 62 | Readings 63 | Readings 64 | Readings 65 | Readings 66 | Readings 67 | Readings 68 | Readings 69 | Readings 70 | Readings 71 | Readings 72 | Readings 73 | Readings 74 | Readings 75 | Readings 76 | Readings 77 | Readings 78 | Readings 79 | Readings 80 | Readings 81 | Readings 82 | Readings 83 | Readings 84 | Readings 85 | Readings 86 | Readings 87 | Readings 88 | Readings 89 | Readings 90 | Readings 91 | Readings 92 | Readings 93 | Readings 94 | Readings 95 | Readings 96 | Readings 97 | Readings 98 | Readings 99 | Readings 100 | Readings 101 | Readings 102 | Readings 103 | Readings 104 | Readings 105 | Readings 106 | Readings 107 | Readings 108 | Readings 109 | Readings 110 | Readings 111 | Readings 112 | Readings 113 | Readings 114 | Readings 115 | Readings 116 | Readings 117 | Readings 118 | Readings 119 | Readings 120 | Readings 121 | Readings 122 | Readings 123 | Readings 124 | Readings 125 | Readings 126 | Readings 127 | Readings 128 | Readings 129 | Readings 130 | Readings 131 | Readings 132 | Readings 133 | Readings 134 | Readings 135 | Readings 136 | Readings 137 | Readings 138 | Readings 139 | Readings 140 | Readings 141 | Readings 142 | Readings 143 | Readings 144 | Readings 145 | Readings 146 | Readings 147 | Readings 148 | Readings 149 | Readings 150 | Readings 151 | Readings 152 | Readings 153 | Readings 154 | Readings 155 | Readings 156 | Readings 157 | Readings 158 | Readings 159 | Readings 160 | Readings 161 | Readings 162 | Readings 163 | Readings 164 | Readings 165 | Readings 166 | Readings 167 | Readings 168 | Readings 169 | Readings 170 | Readings 171 | Readings 172 | Readings 173 | Readings 174 | Readings 175 | Readings 176 | Readings 177 | Readings 178 | Readings 179 | Readings 180 | Readings 181 | Readings 182 | Readings 183 | Readings 184 | Readings 185 | Readings 186 | Readings 187 | Readings 188 | Readings 189 | Readings 190 | Readings 191 | Readings 192 | Readings 193 | Readings 194 | Readings 195 | Readings 196 | Readings 197 | Readings 198 | Readings 199 | Readings 200 | Readings 201 | Readings 202 | Readings 203 | Readings 204 | Readings 205 | Readings 206 | Readings 207 | Readings 208 | Readings 209 | Readings 210 | Readings 211 | Readings 212 | Readings 213 | Readings 214 | Readings 215 | Readings 216 | Readings 217 | Readings 218 | Readings 219 | Readings 220 | Readings 221 | Readings 222 | Readings 223 | Readings 224 | Readings 225 | Readings 226 | Readings 227 | Readings 228 | Readings 229 | Readings 230 | Readings 231 | Readings 232 | Readings 233 | Readings 234 | Readings 235 | Readings 236 | Readings 237 | Readings 238 | Readings 239 | Readings 240 | Readings 241 | Readings 242 | Readings 243 | Readings 244 | Readings 245 | Readings 246 | Readings 247 | Readings 248 | Readings 249 | Readings 250 | Readings 251 | Readings 252 | Readings 253 | Readings 254 | Readings 255
Raw
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 0d5c9e96-1209-ed11-85f4-5cf370a20867 | 8dc854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.3 | 2650.56 | 8.31 | 29213 | [20 empty] | 010081427802027710118705 2105880 17240328 | 8 | 3 | V790 | 2022-07-21 11:32 | 1436 | 266 | 7.9055 | 3994.9491 | 3 | 12 | 4043 | 4024 | 4015 | 4022 | 4148 | 3986 | 4014 | 4026 | 4025 | 4029 | 4052 | 4104 | 4209 | 4410 | 4756 | 5543 | 7164 | 10158 | 14158 | 17999 | 21370 | 23720 | 25748 | 27126 | 28235 | 29116 | 29754 | 30412 | 30998 | 31409 | 31609 | 31979 | 32215 | 32431 | 32640 | 32900 | 33024 | 33159 | 33105 | 33309 | 33392 | 33360 | 33538 | 33464 | 33605
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 145c9e96-1209-ed11-85f4-5cf370a20867 | 2fc854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.24 | 2705.51 | 8.3 | 29527 | [20 empty] | 010081427802027710118705 2105880 17240328 | 9 | 3 | V790 | 2022-07-21 11:32 | 1469 | 268 | 12.2582 | 4050.6436 | 3 | 12 | 4109 | 4124 | 4116 | 4115 | 4112 | 4114 | 4123 | 4116 | 4116 | 4129 | 4168 | 4207 | 4307 | 4507 | 4925 | 5724 | 7445 | 10564 | 14652 | 18604 | 21933 | 24363 | 26242 | 27543 | 28722 | 29599 | 30262 | 30895 | 31477 | 31771 | 32232 | 32522 | 32667 | 33019 | 33286 | 33411 | 33544 | 33682 | 33832 | 33883 | 34051 | 34091 | 34143 | 34066 | 34145
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 1b5c9e96-1209-ed11-85f4-5cf370a20867 | 78ca54cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.2 | 2626.87 | 8.51 | 28852 | [20 empty] | 010081427802027710118705 2105880 17240328 | 10 | 3 | V790 | 2022-07-21 11:32 | 1288 | 265 | 13.2618 | 3841.6764 | 3 | 12 | 3928 | 3907 | 3908 | 3907 | 3920 | 3901 | 3919 | 3908 | 3938 | 3919 | 3950 | 4021 | 4125 | 4325 | 4718 | 5541 | 7259 | 10329 | 14304 | 18151 | 21282 | 23663 | 25398 | 26848 | 27936 | 28812 | 29554 | 30022 | 30571 | 30986 | 31481 | 31521 | 32008 | 32258 | 32452 | 32569 | 32615 | 32875 | 33065 | 33082 | 33144 | 33309 | 33188 | 33351 | 33305
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 225c9e96-1209-ed11-85f4-5cf370a20867 | 17ca54cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.22 | 2575.04 | 8.24 | 28421 | [20 empty] | 010081427802027710118705 2105880 17240328 | 11 | 3 | V790 | 2022-07-21 11:32 | 1419 | 259 | 12.4945 | 3931.0509 | 3 | 12 | 4014 | 4007 | 3996 | 3990 | 3998 | 3986 | 3996 | 4002 | 4013 | 4029 | 4038 | 4072 | 4189 | 4396 | 4782 | 5584 | 7258 | 10236 | 14166 | 17917 | 21050 | 23420 | 25158 | 26637 | 27635 | 28541 | 29200 | 29848 | 30272 | 30687 | 31025 | 31336 | 31651 | 31943 | 32028 | 32211 | 32400 | 32419 | 32603 | 32736 | 32748 | 32749 | 32869 | 32928 | 32925
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 295c9e96-1209-ed11-85f4-5cf370a20867 | b5c954cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.25 | 2811.07 | 8.27 | 29297 | [20 empty] | 010081427802027710118705 2105880 17240328 | 12 | 3 | V790 | 2022-07-21 11:32 | 1453 | 244 | 11.0388 | 4033.9891 | 3 | 12 | 4100 | 4092 | 4088 | 4064 | 4078 | 4194 | 4055 | 4061 | 4090 | 4123 | 4119 | 4172 | 4271 | 4496 | 4873 | 5736 | 7382 | 10533 | 14811 | 19038 | 22182 | 24725 | 26490 | 27958 | 29033 | 29848 | 30468 | 31082 | 31537 | 31910 | 32357 | 32493 | 32742 | 32959 | 33262 | 33182 | 33362 | 33545 | 33574 | 33654 | 33764 | 33751 | 33742 | 33767 | 33876
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 305c9e96-1209-ed11-85f4-5cf370a20867 | 1255c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.65 | 2107.83 | 6.31 | 23255 | 6.73 | 6.55 | 5370317.96 | 3542261.73 | 6.73 | 6.55 |  |  | 5370317.96 | 3542261.73 |  |  | 6.73 | 6.55 | 5370317.96 | 3542261.73 | 6.73 | 6.55 | 5370317.96 | 3542261.73 | 010081427802027710118705 2105879 17240328 | 1 | 4 | V1169 | 2022-07-21 11:32 | 2573 | 296 | 10.5939 | 4394.4255 | 3 | 12 | 4467 | 4426 | 4441 | 4444 | 4447 | 4462 | 4459 | 4445 | 4469 | 4481 | 4506 | 4515 | 4598 | 4696 | 4943 | 5445 | 6473 | 8471 | 11432 | 14636 | 17446 | 19684 | 21295 | 22648 | 23638 | 24381 | 25039 | 25545 | 25915 | 26353 | 26551 | 26794 | 27045 | 27160 | 27311 | 27597 | 27615 | 27790 | 27799 | 27870 | 27938 | 28089 | 28131 | 28193 | 28102
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 375c9e96-1209-ed11-85f4-5cf370a20867 | b354c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.42 | 1950.48 | 6.9 | 23336 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 |  |  | 6309573.44 | 4161794.64 |  |  | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 010081427802027710118705 2105879 17240328 | 2 | 4 | V1169 | 2022-07-21 11:32 | 2058 | 284 | 10.0558 | 3971.7818 | 3 | 12 | 4028 | 4027 | 4009 | 4017 | 4030 | 4028 | 4046 | 4015 | 4022 | 4043 | 4070 | 4117 | 4164 | 4304 | 4558 | 5124 | 6278 | 8394 | 11237 | 14177 | 16687 | 18747 | 20313 | 21638 | 22566 | 23391 | 24115 | 24592 | 25125 | 25534 | 25772 | 26075 | 26301 | 26515 | 26809 | 26994 | 27044 | 27181 | 27312 | 27514 | 27593 | 27574 | 27671 | 27750 | 27793
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 3e5c9e96-1209-ed11-85f4-5cf370a20867 | 7255c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.3 | 2023.88 | 6.93 | 22748 | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 6.84 | 6.66 |  |  | 6918309.71 | 4563317.08 |  |  | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 010081427802027710118705 2105879 17240328 | 3 | 4 | V1169 | 2022-07-21 11:32 | 1979 | 281 | 10.2691 | 3858.8618 | 3 | 12 | 3934 | 3912 | 3912 | 3899 | 3907 | 3911 | 3921 | 3912 | 3944 | 3934 | 3952 | 3987 | 4056 | 4178 | 4465 | 5049 | 6260 | 8487 | 11528 | 14537 | 16963 | 18870 | 20440 | 21511 | 22439 | 23239 | 23777 | 24244 | 24644 | 25038 | 25286 | 25579 | 25920 | 26007 | 26240 | 26321 | 26465 | 26598 | 26677 | 26772 | 26959 | 26915 | 27015 | 27016 | 27105
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 6042e1f8-1209-ed11-85f4-5cf370a20867 | 233d8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.34 | 2623.16 | 6.33 | 27960 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 |  |  | 6760829.75 | 4459443.3 |  |  | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 010081427802027710118705 2105881 17240328 | 12 | 2 | V834 | 2022-07-21 11:34 | 2839 | 265 | 11.2448 | 5279.6436 | 3 | 12 | 5375 | 5371 | 5342 | 5323 | 5342 | 5325 | 5337 | 5330 | 5344 | 5366 | 5384 | 5426 | 5530 | 5670 | 6041 | 6763 | 8262 | 11089 | 15033 | 18933 | 22118 | 24533 | 26450 | 27814 | 28853 | 29716 | 30369 | 30983 | 31350 | 31803 | 32170 | 32399 | 32595 | 32876 | 33003 | 33112 | 33344 | 33349 | 33462 | 33522 | 33601 | 33751 | 33667 | 33758 | 33767
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 6642e1f8-1209-ed11-85f4-5cf370a20867 | 7e8c0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.35 | 2535.4 | 8.04 | 28768 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 |  |  | 6606934.48 | 4357933.98 |  |  | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 010081427802027710118705 2105880 17240328 | 1 | 3 | V790 | 2022-07-21 11:34 | 1595 | 256 | 12.2133 | 4091.92 | 3 | 12 | 4172 | 4146 | 4149 | 4156 | 4155 | 4158 | 4157 | 4168 | 4157 | 4192 | 4193 | 4252 | 4329 | 4520 | 4896 | 5628 | 7128 | 9905 | 13707 | 17505 | 20670 | 23110 | 24952 | 26372 | 27426 | 28335 | 29108 | 29752 | 30204 | 30703 | 31073 | 31422 | 31752 | 31995 | 32264 | 32452 | 32577 | 32791 | 32864 | 33001 | 33160 | 33219 | 33323 | 33346 | 33459
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 6c42e1f8-1209-ed11-85f4-5cf370a20867 | ef8b0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.32 | 2588.44 | 7.64 | 29845 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 |  |  | 6760829.75 | 4459443.3 |  |  | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 010081427802027710118705 2105880 17240328 | 2 | 3 | V790 | 2022-07-21 11:34 | 1882 | 264 | 9.5576 | 4509.8982 | 3 | 12 | 4594 | 4587 | 4547 | 4563 | 4569 | 4545 | 4555 | 4560 | 4562 | 4553 | 4604 | 4639 | 4729 | 4907 | 5284 | 6030 | 7598 | 10494 | 14358 | 18201 | 21406 | 23897 | 25878 | 27428 | 28560 | 29568 | 30308 | 30996 | 31546 | 31971 | 32447 | 32818 | 33206 | 33395 | 33637 | 33787 | 34012 | 34217 | 34247 | 34360 | 34528 | 34604 | 34732 | 34792 | 34801
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 7242e1f8-1209-ed11-85f4-5cf370a20867 | bf8b0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.37 | 2773.69 | 8.52 | 30635 | [20 empty] | 010081427802027710118705 2105880 17240328 | 3 | 3 | V790 | 2022-07-21 11:34 | 1410 | 242 | 14.6582 | 4079.0836 | 3 | 12 | 4144 | 4152 | 4139 | 4159 | 4160 | 4155 | 4164 | 4152 | 4178 | 4195 | 4223 | 4266 | 4351 | 4538 | 4927 | 5683 | 7263 | 10235 | 14392 | 18557 | 21943 | 24553 | 26497 | 28019 | 29105 | 30070 | 30791 | 31420 | 31948 | 32447 | 32905 | 33276 | 33564 | 33869 | 34122 | 34268 | 34510 | 34707 | 34835 | 34960 | 35101 | 35115 | 35276 | 35346 | 35415
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 7842e1f8-1209-ed11-85f4-5cf370a20867 | 608b0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.46 | 2532.36 | 8.41 | 30118 | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 6.79 | 6.61 |  |  | 6165950.02 | 4067060.63 |  |  | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 010081427802027710118705 2105880 17240328 | 4 | 3 | V790 | 2022-07-21 11:34 | 1354 | 259 | 12.6812 | 4069.7309 | 3 | 12 | 4133 | 4139 | 4135 | 4131 | 4128 | 4133 | 4143 | 4146 | 4157 | 4167 | 4197 | 4218 | 4302 | 4486 | 4849 | 5530 | 6984 | 9667 | 13396 | 17217 | 20527 | 23057 | 25178 | 26825 | 28109 | 29049 | 29895 | 30520 | 31157 | 31655 | 32133 | 32419 | 32817 | 33103 | 33356 | 33636 | 33858 | 34050 | 34193 | 34356 | 34406 | 34486 | 34657 | 34751 | 34794
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 7e42e1f8-1209-ed11-85f4-5cf370a20867 | ebaadb1d-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified | 2610 (Informational, Delays in sample processing) | 16.12 | 1603.89 | 6.52 | 16922 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 |  |  | 7762471.17 | 5120125.98 |  |  | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 010081427802027710118705 2105880 17240328 | 5 | 3 | V790 | 2022-07-21 11:34 | 1372 | 268 | 2.2497 | 3104.4473 | 3 | 12 | 3105 | 3103 | 3124 | 3126 | 3132 | 3128 | 3123 | 3094 | 3091 | 3093 | 3084 | 3168 | 3199 | 3336 | 3582 | 4130 | 5141 | 6950 | 9485 | 11869 | 13452 | 14885 | 16099 | 16778 | 17310 | 17890 | 18230 | 18622 | 19012 | 19183 | 19360 | 19480 | 19449 | 19720 | 19563 | 19849 | 19986 | 19853 | 19880 | 19984 | 19914 | 19530 | 19591 | 20062 | 20328
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 8442e1f8-1209-ed11-85f4-5cf370a20867 | 57c954cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.29 | 2625.53 | 8.37 | 30075 | [20 empty] | 010081427802027710118705 2105880 17240328 | 6 | 3 | V790 | 2022-07-21 11:34 | 1388 | 245 | 11.3673 | 4079.7455 | 3 | 12 | 4128 | 4127 | 4124 | 4130 | 4124 | 4235 | 4116 | 4113 | 4130 | 4171 | 4172 | 4229 | 4320 | 4516 | 4914 | 5718 | 7289 | 10243 | 14193 | 18133 | 21287 | 23841 | 25807 | 27324 | 28492 | 29410 | 30185 | 30819 | 31510 | 31879 | 32351 | 32682 | 32920 | 33181 | 33578 | 33667 | 33878 | 34050 | 34125 | 34253 | 34375 | 34480 | 34552 | 34592 | 34729
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 8a42e1f8-1209-ed11-85f4-5cf370a20867 | edc854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.13 | 2178.33 | 8.18 | 26676 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 |  |  | 7762471.17 | 5120125.98 |  |  | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 010081427802027710118705 2105880 17240328 | 7 | 3 | V790 | 2022-07-21 11:34 | 1317 | 270 | 11.3709 | 3719.5782 | 3 | 12 | 3779 | 3769 | 3785 | 3770 | 3790 | 3782 | 3787 | 3772 | 3783 | 3792 | 3822 | 3863 | 3968 | 4148 | 4514 | 5312 | 6887 | 9537 | 12837 | 15993 | 18684 | 20840 | 22557 | 23866 | 25049 | 25937 | 26662 | 27199 | 27796 | 28164 | 28666 | 28911 | 29225 | 29488 | 29729 | 29886 | 30109 | 30238 | 30369 | 30451 | 30673 | 30796 | 30791 | 30888 | 30952
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 5a7b815a-1309-ed11-85f4-5cf370a20867 | 223c8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.28 | 2631.98 | 6.57 | 27149 | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 6.85 | 6.67 |  |  | 7079457.84 | 4669610.39 |  |  | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 010081427802027710118705 2105881 17240328 | 7 | 2 | V834 | 2022-07-21 11:37 | 2538 | 268 | 11.423 | 4907.5673 | 3 | 12 | 4997 | 4971 | 4993 | 4972 | 4954 | 4956 | 4956 | 4973 | 4970 | 4995 | 5012 | 5047 | 5152 | 5343 | 5721 | 6458 | 8061 | 11004 | 14983 | 18892 | 22022 | 24368 | 26070 | 27323 | 28292 | 28978 | 29654 | 30188 | 30575 | 30803 | 31095 | 31354 | 31569 | 31775 | 31876 | 32095 | 32030 | 32266 | 32329 | 32379 | 32446 | 32361 | 32533 | 32564 | 32585
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 607b815a-1309-ed11-85f4-5cf370a20867 | 50597999-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.37 | 2314.8 | 6.2 | 25184 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 |  |  | 6606934.48 | 4357933.98 |  |  | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 010081427802027710118705 2105881 17240328 | 8 | 2 | V834 | 2022-07-21 11:37 | 2749 | 292 | 8.6594 | 4878.6545 | 3 | 12 | 4950 | 4950 | 4924 | 4929 | 4918 | 4913 | 4925 | 4917 | 4925 | 4932 | 4956 | 5001 | 5073 | 5219 | 5533 | 6158 | 7462 | 9936 | 13419 | 16856 | 19686 | 21900 | 23546 | 24796 | 25744 | 26623 | 27401 | 27795 | 28296 | 28641 | 28928 | 29167 | 29473 | 29645 | 29729 | 29934 | 30041 | 30026 | 30150 | 30296 | 30231 | 30299 | 30369 | 30417 | 30492
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 667b815a-1309-ed11-85f4-5cf370a20867 | 9c3e8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.45 | 2392.68 | 5.45 | 25142 | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 6.79 | 6.61 |  |  | 6165950.02 | 4067060.63 |  |  | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 010081427802027710118705 2105881 17240328 | 9 | 2 | V834 | 2022-07-21 11:37 | 3532 | 272 | 7.1382 | 5690.0436 | 3 | 12 | 5760 | 5743 | 5739 | 5723 | 5719 | 5728 | 5736 | 5717 | 5715 | 5723 | 5748 | 5810 | 5846 | 6020 | 6305 | 6910 | 8215 | 10685 | 14208 | 17831 | 20807 | 23050 | 24798 | 25995 | 26995 | 27665 | 28301 | 28715 | 29200 | 29505 | 29674 | 30014 | 30234 | 30394 | 30471 | 30658 | 30650 | 30867 | 30830 | 30892 | 30982 | 30984 | 31019 | 31156 | 31197
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 6c7b815a-1309-ed11-85f4-5cf370a20867 | 3e3e8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.38 | 2436.69 | 5.53 | 25415 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 |  |  | 6606934.48 | 4357933.98 |  |  | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 010081427802027710118705 2105881 17240328 | 10 | 2 | V834 | 2022-07-21 11:37 | 3437 | 275 | 7.8133 | 5650.56 | 3 | 12 | 5704 | 5729 | 5696 | 5710 | 5676 | 5687 | 5681 | 5688 | 5677 | 5702 | 5704 | 5764 | 5859 | 5987 | 6289 | 6951 | 8316 | 10893 | 14529 | 18165 | 21138 | 23344 | 25052 | 26230 | 27204 | 27859 | 28515 | 28937 | 29396 | 29775 | 29980 | 30341 | 30416 | 30567 | 30724 | 30921 | 31003 | 31052 | 31138 | 31203 | 31252 | 31291 | 31413 | 31471 | 31400
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 727b815a-1309-ed11-85f4-5cf370a20867 | e03d8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.44 | 2298.39 | 5.71 | 25080 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 |  |  | 6309573.44 | 4161794.64 |  |  | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 010081427802027710118705 2105881 17240328 | 11 | 2 | V834 | 2022-07-21 11:37 | 3289 | 280 | 6.577 | 5366.3527 | 3 | 12 | 5452 | 5420 | 5404 | 5421 | 5399 | 5390 | 5380 | 5389 | 5397 | 5398 | 5407 | 5468 | 5555 | 5689 | 5971 | 6578 | 7831 | 10260 | 13604 | 17108 | 19982 | 22069 | 23814 | 25092 | 26144 | 26982 | 27578 | 28120 | 28568 | 28888 | 29244 | 29593 | 29650 | 29922 | 30171 | 30149 | 30216 | 30354 | 30546 | 30526 | 30609 | 30680 | 30689 | 30719 | 30768
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 787b815a-1309-ed11-85f4-5cf370a20867 | 1eff3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.43 | 2367.47 | 6.98 | 25496 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 |  |  | 6309573.44 | 4161794.64 |  |  | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 010081427802027710118705 2105881 17240328 | 4 | 2 | V834 | 2022-07-21 11:37 | 2183 | 247 | 10.5467 | 4277.56 | 3 | 12 | 4343 | 4344 | 4343 | 4328 | 4321 | 4334 | 4330 | 4338 | 4338 | 4363 | 4369 | 4409 | 4485 | 4668 | 4953 | 5566 | 6880 | 9355 | 12859 | 16469 | 19359 | 21627 | 23331 | 24541 | 25557 | 26488 | 27073 | 27581 | 28008 | 28455 | 28779 | 29003 | 29217 | 29310 | 29506 | 29601 | 29834 | 29837 | 29991 | 30012 | 30127 | 30197 | 30285 | 30238 | 30239
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 7e7b815a-1309-ed11-85f4-5cf370a20867 | 823d8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.4 | 2435.14 | 6.91 | 25271 | 6.81 | 6.63 | 6456542.29 | 4258735.29 | 6.81 | 6.63 |  |  | 6456542.29 | 4258735.29 |  |  | 6.81 | 6.63 | 6456542.29 | 4258735.29 | 6.81 | 6.63 | 6456542.29 | 4258735.29 | 010081427802027710118705 2105881 17240328 | 5 | 2 | V834 | 2022-07-21 11:37 | 2203 | 258 | 11.3697 | 4297.6873 | 3 | 12 | 4364 | 4351 | 4353 | 4360 | 4355 | 4351 | 4365 | 4371 | 4358 | 4382 | 4398 | 4454 | 4521 | 4677 | 4977 | 5626 | 6978 | 9534 | 13177 | 16805 | 19806 | 22067 | 23725 | 24953 | 25953 | 26658 | 27302 | 27783 | 28060 | 28492 | 28703 | 28974 | 29165 | 29362 | 29467 | 29638 | 29692 | 29749 | 29808 | 29946 | 30005 | 30050 | 29976 | 30069 | 30119
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 847b815a-1309-ed11-85f4-5cf370a20867 | c43c8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.18 | 2359.93 | 6.51 | 24908 | 6.88 | 6.7 | 7585775.75 | 5003577.68 | 6.88 | 6.7 |  |  | 7585775.75 | 5003577.68 |  |  | 6.88 | 6.7 | 7585775.75 | 5003577.68 | 6.88 | 6.7 | 7585775.75 | 5003577.68 | 010081427802027710118705 2105881 17240328 | 6 | 2 | V834 | 2022-07-21 11:37 | 2433 | 253 | 11.0485 | 4546.5564 | 3 | 12 | 4607 | 4604 | 4609 | 4603 | 4604 | 4599 | 4612 | 4608 | 4614 | 4614 | 4636 | 4694 | 4807 | 4922 | 5307 | 6020 | 7534 | 10295 | 13879 | 17293 | 20039 | 22148 | 23739 | 24825 | 25813 | 26490 | 27118 | 27592 | 27930 | 28363 | 28514 | 28864 | 28937 | 29132 | 29245 | 29441 | 29479 | 29544 | 29628 | 29748 | 29717 | 29793 | 29842 | 29992 | 29975
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | b3de19db-1309-ed11-85f4-5cf370a20867 | 685f6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.51 | 2123.17 | 6.78 | 22703 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 |  |  | 6025595.86 | 3974483.03 |  |  | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 010081427802027710118705 2105881 17240328 | 3 | 2 | V834 | 2022-07-21 11:41 | 2093 | 263 | 7.3515 | 3941.8836 | 3 | 12 | 3995 | 3979 | 3977 | 3974 | 3994 | 3980 | 3973 | 3976 | 3992 | 3984 | 4017 | 4025 | 4099 | 4235 | 4490 | 5044 | 6137 | 8272 | 11346 | 14611 | 17260 | 19341 | 21000 | 22109 | 23093 | 23806 | 24271 | 24775 | 25110 | 25411 | 25746 | 25859 | 26056 | 26284 | 26329 | 26492 | 26667 | 26648 | 26765 | 26834 | 26845 | 26898 | 26928 | 27008 | 26980
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | b9de19db-1309-ed11-85f4-5cf370a20867 | 785e6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.61 | 2768.06 | 7.99 | 29669 | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 6.75 | 6.57 |  |  | 5623413.25 | 3709203.38 |  |  | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 010081427802027710118705 2105882 17240328 | 10 | 1 | V639 | 2022-07-21 11:41 | 1720 | 268 | 12.9721 | 4239.3891 | 3 | 12 | 4298 | 4287 | 4291 | 4304 | 4313 | 4314 | 4307 | 4337 | 4319 | 4333 | 4354 | 4407 | 4487 | 4647 | 4968 | 5609 | 6974 | 9603 | 13580 | 17755 | 21402 | 24105 | 26168 | 27739 | 28912 | 29985 | 30687 | 31264 | 31824 | 32180 | 32673 | 32842 | 33144 | 33380 | 33594 | 33791 | 33977 | 34063 | 34104 | 34192 | 34286 | 34432 | 34477 | 34505 | 34493
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | bfde19db-1309-ed11-85f4-5cf370a20867 | 485e6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.29 | 2777.15 | 8.24 | 29533 | [20 empty] | 010081427802027710118705 2105882 17240328 | 11 | 1 | V639 | 2022-07-21 11:41 | 1560 | 266 | 14.7564 | 4080.1273 | 3 | 12 | 4140 | 4148 | 4158 | 4161 | 4145 | 4174 | 4157 | 4163 | 4169 | 4194 | 4225 | 4255 | 4375 | 4561 | 4935 | 5747 | 7426 | 10506 | 14711 | 18795 | 22172 | 24613 | 26533 | 27883 | 28994 | 30747 | 30672 | 31283 | 31754 | 32059 | 32501 | 32708 | 32976 | 33194 | 33362 | 33628 | 34648 | 33908 | 33900 | 34036 | 34061 | 34220 | 34203 | 34321 | 34287
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | c5de19db-1309-ed11-85f4-5cf370a20867 | 90fe3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.47 | 2893.71 | 8.04 | 30993 | [20 empty] | 010081427802027710118705 2105882 17240328 | 12 | 1 | V639 | 2022-07-21 11:41 | 1817 | 263 | 16.143 | 4407.3273 | 3 | 12 | 4471 | 4497 | 4458 | 4491 | 4495 | 4501 | 4497 | 4499 | 4518 | 4541 | 4560 | 4606 | 4703 | 4898 | 5246 | 5996 | 7566 | 10553 | 14817 | 19191 | 22840 | 25630 | 27792 | 29373 | 30643 | 31501 | 32345 | 32976 | 33519 | 33827 | 34261 | 34540 | 34808 | 35039 | 35185 | 35330 | 35592 | 35687 | 35688 | 35801 | 35881 | 35925 | 36009 | 36141 | 36161
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | cbde19db-1309-ed11-85f4-5cf370a20867 | 60fe3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.49 | 2361.31 | 5.95 | 25803 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 |  |  | 6025595.86 | 3974483.03 |  |  | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 010081427802027710118705 2105881 17240328 | 1 | 2 | V834 | 2022-07-21 11:41 | 3061 | 258 | 8.7067 | 5253.04 | 3 | 12 | 5299 | 5300 | 5308 | 5299 | 5310 | 5304 | 5292 | 5269 | 5296 | 5317 | 5346 | 5367 | 5445 | 5583 | 5853 | 6495 | 7770 | 10187 | 13612 | 17209 | 20223 | 22519 | 24374 | 25644 | 26728 | 27565 | 28246 | 28823 | 29320 | 29618 | 29903 | 30149 | 30383 | 30551 | 30743 | 30865 | 30965 | 31079 | 31237 | 31238 | 31338 | 31420 | 31366 | 31391 | 31494
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | d1de19db-1309-ed11-85f4-5cf370a20867 | 985f6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.29 | 2373.91 | 6.32 | 24689 | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 6.84 | 6.66 |  |  | 6918309.71 | 4563317.08 |  |  | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 010081427802027710118705 2105881 17240328 | 2 | 2 | V834 | 2022-07-21 11:41 | 2647 | 271 | 10.5273 | 4659.0655 | 3 | 12 | 4716 | 4707 | 4717 | 4724 | 4708 | 4707 | 4719 | 4726 | 4720 | 4728 | 4748 | 4798 | 4874 | 5050 | 5394 | 6049 | 7462 | 10084 | 13668 | 17232 | 20012 | 22148 | 23743 | 24855 | 25847 | 26501 | 27124 | 27482 | 27943 | 28156 | 28538 | 28620 | 28865 | 29123 | 29129 | 29265 | 29453 | 29464 | 29550 | 29709 | 29638 | 29682 | 29842 | 29872 | 29799
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | d7de19db-1309-ed11-85f4-5cf370a20867 | 954775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.27 | 2874.99 | 8.13 | 29909 | [20 empty] | 010081427802027710118705 2105882 17240328 | 8 | 1 | V639 | 2022-07-21 11:41 | 1689 | 271 | 15.3721 | 4204.9491 | 3 | 12 | 4278 | 4278 | 4271 | 4291 | 4289 | 4294 | 4287 | 4282 | 4303 | 4312 | 4351 | 4402 | 4524 | 4708 | 5117 | 5958 | 7701 | 10907 | 15256 | 19537 | 22891 | 25443 | 27267 | 28657 | 29774 | 30634 | 31338 | 31862 | 32345 | 32686 | 32988 | 33354 | 33536 | 33819 | 34025 | 34164 | 34300 | 34493 | 34588 | 34660 | 34699 | 34819 | 34724 | 34773 | 34844
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | ddde19db-1309-ed11-85f4-5cf370a20867 | db5e6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.06 | 2821.6 | 7.95 | 29515 | [20 empty] | 010081427802027710118705 2105882 17240328 | 9 | 1 | V639 | 2022-07-21 11:41 | 1825 | 274 | 18.3176 | 4254.1782 | 3 | 12 | 4338 | 4363 | 4341 | 4345 | 4345 | 4346 | 4354 | 4369 | 4368 | 4394 | 4424 | 4504 | 4597 | 4829 | 5299 | 6264 | 8246 | 11717 | 16027 | 20075 | 23251 | 25654 | 27434 | 28717 | 29790 | 30498 | 31221 | 31730 | 32196 | 32510 | 32907 | 33201 | 33380 | 33658 | 33713 | 33882 | 34032 | 34084 | 34201 | 34288 | 34486 | 34385 | 34491 | 34646 | 34610
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 65fdae7a-1409-ed11-85f4-5cf370a20867 | 364775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.27 | 2347.62 | 7.47 | 26286 | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 6.85 | 6.67 |  |  | 7079457.84 | 4669610.39 |  |  | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 010081427802027710118705 2105882 17240328 | 4 | 1 | V639 | 2022-07-21 11:45 | 1808 | 260 | 11.5855 | 4075.4291 | 3 | 12 | 4128 | 4123 | 4136 | 4142 | 4139 | 4139 | 4135 | 4136 | 4148 | 4158 | 4174 | 4231 | 4309 | 4484 | 4848 | 5532 | 7002 | 9619 | 13209 | 16663 | 19489 | 21750 | 23518 | 25165 | 25831 | 26683 | 27349 | 27935 | 28353 | 28754 | 29137 | 29316 | 29638 | 29835 | 30070 | 30213 | 30268 | 30437 | 30535 | 30624 | 30685 | 30777 | 30796 | 30858 | 30920
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 6bfdae7a-1409-ed11-85f4-5cf370a20867 | a94675b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.02 | 2809.5 | 7.97 | 30622 | 6.92 | 6.74 | 8317637.71 | 5486313.83 | 6.92 | 6.74 |  |  | 8317637.71 | 5486313.83 |  |  | 6.92 | 6.74 | 8317637.71 | 5486313.83 | 6.92 | 6.74 | 8317637.71 | 5486313.83 | 010081427802027710118705 2105882 17240328 | 5 | 1 | V639 | 2022-07-21 11:45 | 1826 | 257 | 19.0012 | 4401.4909 | 3 | 12 | 4487 | 4499 | 4493 | 4483 | 4506 | 4495 | 4523 | 4511 | 4530 | 4544 | 4573 | 4640 | 4771 | 5005 | 5499 | 6502 | 8573 | 12081 | 16413 | 20388 | 23554 | 26004 | 27790 | 29241 | 30295 | 31199 | 31934 | 32640 | 33130 | 33527 | 34040 | 34383 | 34571 | 34985 | 35007 | 35149 | 35421 | 35406 | 35526 | 35682 | 35747 | 35781 | 35747 | 35871 | 35925
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 71fdae7a-1409-ed11-85f4-5cf370a20867 | 774675b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.14 | 2540.68 | 7.78 | 28700 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 |  |  | 7762471.17 | 5120125.98 |  |  | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 010081427802027710118705 2105882 17240328 | 6 | 1 | V639 | 2022-07-21 11:45 | 1746 | 262 | 16.9309 | 4242.9382 | 3 | 12 | 4331 | 4325 | 4335 | 4329 | 4319 | 4327 | 4343 | 4342 | 4355 | 4370 | 4420 | 4446 | 4561 | 4774 | 5185 | 6013 | 7794 | 10833 | 14718 | 18385 | 21411 | 23720 | 25554 | 26997 | 28003 | 28961 | 29688 | 30374 | 30793 | 31340 | 31611 | 32011 | 32203 | 32494 | 32636 | 32899 | 33055 | 33122 | 33375 | 33447 | 33485 | 33510 | 33618 | 33751 | 33719
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 77fdae7a-1409-ed11-85f4-5cf370a20867 | c64775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.03 | 2805.63 | 8.16 | 29973 | [20 empty] | 010081427802027710118705 2105882 17240328 | 7 | 1 | V639 | 2022-07-21 11:45 | 1646 | 261 | 19.9042 | 4183.5382 | 3 | 12 | 4272 | 4292 | 4277 | 4281 | 4278 | 4275 | 4317 | 4302 | 4323 | 4330 | 4374 | 4430 | 4560 | 4799 | 5290 | 6285 | 8302 | 11805 | 16113 | 20133 | 23250 | 25705 | 27407 | 28751 | 29843 | 30663 | 31354 | 31961 | 32534 | 32870 | 33290 | 33574 | 33773 | 33930 | 34129 | 34300 | 34438 | 34600 | 34665 | 34820 | 34809 | 34900 | 35041 | 35079 | 35047
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | f47fc1ae-1609-ed11-85f4-5cf370a20867 | 816b47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.14 | 2841.9 | 7.99 | 29636 | [20 empty] | 010081427802027710118705 2105882 17240328 | 1 | 1 | V639 | 2022-07-21 12:01 | 1679 | 254 | 18.6121 | 4231.3491 | 3 | 12 | 4309 | 4326 | 4309 | 4317 | 4330 | 4327 | 4346 | 4354 | 4347 | 4382 | 4407 | 4467 | 4577 | 4778 | 5234 | 6139 | 8007 | 11349 | 15710 | 19836 | 23064 | 25485 | 27300 | 28626 | 29731 | 30524 | 31237 | 31737 | 32287 | 32608 | 32912 | 33302 | 33530 | 33734 | 33823 | 33999 | 34178 | 34254 | 34439 | 34421 | 34530 | 34616 | 34726 | 34779 | 34673
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | fa7fc1ae-1609-ed11-85f4-5cf370a20867 | 226b47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.12 | 3139.19 | 8.24 | 32203 | [20 empty] | 010081427802027710118705 2105882 17240328 | 2 | 1 | V639 | 2022-07-21 12:01 | 1768 | 263 | 19.5394 | 4452.2145 | 3 | 12 | 4536 | 4548 | 4549 | 4554 | 4559 | 4554 | 4538 | 4583 | 4578 | 4607 | 4625 | 4705 | 4825 | 5083 | 5575 | 6623 | 8678 | 12387 | 17252 | 21808 | 25295 | 27953 | 29848 | 31222 | 32268 | 33152 | 33899 | 34540 | 34984 | 35357 | 35691 | 35913 | 36161 | 36433 | 36580 | 36792 | 36954 | 37094 | 37197 | 37278 | 37402 | 37404 | 37424 | 37586 | 37554
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | 0080c1ae-1609-ed11-85f4-5cf370a20867 | e16b47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.13 | 3191.32 | 8.32 | 32904 | [20 empty] | 010081427802027710118705 2105882 17240328 | 3 | 1 | V639 | 2022-07-21 12:01 | 1736 | 265 | 18.4218 | 4505.0364 | 3 | 12 | 4614 | 4585 | 4592 | 4608 | 4594 | 4611 | 4610 | 4596 | 4619 | 4647 | 4679 | 4732 | 4887 | 5097 | 5608 | 6602 | 8698 | 12513 | 17410 | 21983 | 25625 | 28212 | 30189 | 31644 | 32789 | 33663 | 34409 | 35009 | 35557 | 35988 | 36420 | 36703 | 36952 | 37210 | 37263 | 37431 | 37634 | 37717 | 37781 | 37977 | 38028 | 38065 | 38138 | 38201 | 38283
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | a79dad68-1f09-ed11-85f4-5cf370a20867 | b26dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | ad9dad68-1f09-ed11-85f4-5cf370a20867 | b66dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | b39dad68-1f09-ed11-85f4-5cf370a20867 | ba6dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | b99dad68-1f09-ed11-85f4-5cf370a20867 | be6dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | bf9dad68-1f09-ed11-85f4-5cf370a20867 | c26dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | c59dad68-1f09-ed11-85f4-5cf370a20867 | c76dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | cb9dad68-1f09-ed11-85f4-5cf370a20867 | 2f0bca23-2809-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 17.03 | 2397.7 | 7.16 | 25393 | 6.62 | 6.44 | 4168693.83 | 2749670.45 | 6.62 | 6.44 |  |  | 4168693.83 | 2749670.45 |  |  | 6.62 | 6.44 | 4168693.83 | 2749670.45 | 6.62 | 6.44 | 4168693.83 | 2749670.45 | 010081427802027710118705 2105883 17240328 | 11 | 4 | V1169 | 2022-07-21 13:03 | 2048 | 285 | 11.6691 | 4140.7564 | 3 | 13 | 4175 | 4235 | 4206 | 4192 | 4192 | 4211 | 4225 | 4200 | 4220 | 4222 | 4219 | 4263 | 4311 | 4411 | 4593 | 5007 | 5872 | 7540 | 10498 | 14239 | 17585 | 20266 | 22219 | 23704 | 24817 | 25780 | 26400 | 27002 | 27557 | 27834 | 28156 | 28451 | 28793 | 28964 | 29076 | 29221 | 29386 | 29543 | 29706 | 29690 | 29775 | 29955 | 29907 | 30047 | 30113
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | d29dad68-1f09-ed11-85f4-5cf370a20867 | d10aca23-2809-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.59 | 2333.73 | 6.98 | 25183 | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 6.75 | 6.57 |  |  | 5623413.25 | 3709203.38 |  |  | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 010081427802027710118705 2105883 17240328 | 12 | 4 | V1169 | 2022-07-21 13:03 | 2088 | 278 | 10.3709 | 4218.2182 | 3 | 12 | 4261 | 4262 | 4267 | 4264 | 4257 | 4292 | 4285 | 4297 | 4265 | 4308 | 4314 | 4341 | 4409 | 4541 | 4797 | 5353 | 6469 | 8682 | 12004 | 15619 | 18541 | 20830 | 22541 | 23840 | 24818 | 25661 | 26257 | 26805 | 27330 | 27610 | 27981 | 28258 | 28590 | 28751 | 28859 | 29152 | 29236 | 29434 | 29584 | 29563 | 29716 | 29813 | 29820 | 29873 | 29882
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 1f42d920-2009-ed11-85f4-5cf370a20867 | 41e5bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.64 | 2405.59 | 7.37 | 25353 | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 6.74 | 6.56 |  |  | 5495408.74 | 3624771.6 |  |  | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 010081427802027710118705 2105883 17240328 | 5 | 4 | V1169 | 2022-07-21 13:09 | 1954 | 288 | 13.0194 | 3986.2545 | 3 | 12 | 4044 | 4045 | 4051 | 4051 | 4043 | 4048 | 4071 | 4074 | 4083 | 4096 | 4089 | 4140 | 4233 | 4387 | 4652 | 5225 | 6425 | 8694 | 12076 | 15787 | 18942 | 21379 | 23028 | 24372 | 25356 | 26111 | 26746 | 27378 | 27630 | 28125 | 28324 | 28555 | 28841 | 29006 | 29179 | 29363 | 29481 | 29442 | 29658 | 29608 | 29756 | 29715 | 29877 | 29949 | 29934
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 2642d920-2009-ed11-85f4-5cf370a20867 | b3e4bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.08 | 2303.83 | 7.28 | 24419 | 6.91 | 6.73 | 8128305.16 | 5361430.08 | 6.91 | 6.73 |  |  | 8128305.16 | 5361430.08 |  |  | 6.91 | 6.73 | 8128305.16 | 5361430.08 | 6.91 | 6.73 | 8128305.16 | 5361430.08 | 010081427802027710118705 2105883 17240328 | 6 | 4 | V1169 | 2022-07-21 13:09 | 1944 | 278 | 15.3091 | 3909.1018 | 3 | 12 | 3993 | 3988 | 3994 | 3979 | 3985 | 3986 | 4003 | 3987 | 4006 | 4040 | 4032 | 4116 | 4208 | 4402 | 4800 | 5600 | 7214 | 10051 | 13586 | 16880 | 19580 | 21558 | 22978 | 24046 | 24891 | 25642 | 26119 | 26670 | 26942 | 27275 | 27551 | 27810 | 28067 | 28226 | 28325 | 28482 | 28575 | 28655 | 28819 | 28809 | 28891 | 29014 | 28966 | 28976 | 29048
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 2c42d920-2009-ed11-85f4-5cf370a20867 | bde6bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.63 | 2270.38 | 7.03 | 24694 | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 6.74 | 6.56 |  |  | 5495408.74 | 3624771.6 |  |  | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 010081427802027710118705 2105883 17240328 | 7 | 4 | V1169 | 2022-07-21 13:09 | 2123 | 298 | 10.1176 | 4103.1782 | 3 | 12 | 4167 | 4143 | 4154 | 4153 | 4156 | 4160 | 4164 | 4166 | 4156 | 4164 | 4196 | 4251 | 4306 | 4447 | 4703 | 5254 | 6362 | 8533 | 11747 | 15197 | 18220 | 20493 | 22141 | 23451 | 24420 | 25224 | 25944 | 26503 | 26852 | 27216 | 27541 | 27910 | 28021 | 28222 | 28467 | 28673 | 28682 | 28790 | 28927 | 28997 | 29033 | 29198 | 29233 | 29225 | 29268
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 3242d920-2009-ed11-85f4-5cf370a20867 | 5ee6bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.32 | 2201.98 | 6.92 | 24142 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 |  |  | 6760829.75 | 4459443.3 |  |  | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 010081427802027710118705 2105883 17240328 | 8 | 4 | V1169 | 2022-07-21 13:09 | 2146 | 296 | 13.8606 | 4100.0255 | 3 | 12 | 4173 | 4177 | 4162 | 4162 | 4162 | 4181 | 4175 | 4201 | 4175 | 4226 | 4216 | 4274 | 4366 | 4520 | 4855 | 5510 | 6804 | 9244 | 12573 | 15859 | 18563 | 20631 | 22139 | 23306 | 24295 | 24964 | 25544 | 26144 | 26470 | 26834 | 27134 | 27422 | 27679 | 27800 | 28055 | 28176 | 28392 | 28489 | 28564 | 28548 | 28642 | 28829 | 28807 | 28811 | 28903
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 3842d920-2009-ed11-85f4-5cf370a20867 | d2e5bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.73 | 2195.79 | 7.21 | 24578 | 6.71 | 6.53 | 5128613.84 | 3382833.69 | 6.71 | 6.53 |  |  | 5128613.84 | 3382833.69 |  |  | 6.71 | 6.53 | 5128613.84 | 3382833.69 | 6.71 | 6.53 | 5128613.84 | 3382833.69 | 010081427802027710118705 2105883 17240328 | 9 | 4 | V1169 | 2022-07-21 13:09 | 1973 | 289 | 7.6679 | 3971.2109 | 3 | 12 | 4028 | 4017 | 4011 | 4017 | 4001 | 4006 | 4013 | 4019 | 4008 | 4037 | 4042 | 4068 | 4118 | 4245 | 4479 | 4957 | 5933 | 7896 | 10952 | 14312 | 17232 | 19572 | 21298 | 22684 | 23787 | 24666 | 25257 | 25875 | 26386 | 26668 | 27007 | 27326 | 27513 | 27731 | 27900 | 28079 | 28288 | 28378 | 28439 | 28525 | 28646 | 28688 | 28816 | 28899 | 28918
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 3e42d920-2009-ed11-85f4-5cf370a20867 | a1e5bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.5 | 2247.9 | 6.88 | 24232 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 |  |  | 6025595.86 | 3974483.03 |  |  | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 010081427802027710118705 2105883 17240328 | 10 | 4 | V1169 | 2022-07-21 13:09 | 2142 | 280 | 10.1648 | 4127.8836 | 3 | 12 | 4174 | 4170 | 4175 | 4180 | 4196 | 4177 | 4176 | 4195 | 4193 | 4187 | 4223 | 4252 | 4349 | 4481 | 4766 | 5329 | 6514 | 8788 | 12102 | 15457 | 18347 | 20496 | 22098 | 23295 | 24218 | 24956 | 25670 | 26048 | 26485 | 26830 | 27150 | 27408 | 27628 | 27810 | 27997 | 28033 | 28300 | 28380 | 28476 | 28640 | 28606 | 28662 | 28797 | 28857 | 28810
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 4442d920-2009-ed11-85f4-5cf370a20867 | 813d7504-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:09 |  |  |  |  | 0
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | b7477ca2-3609-ed11-85f4-5cf370a20867 | f3c2a3da-3609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:50 |  |  |  |  | 0
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | f5c2a3da-3609-ed11-85f4-5cf370a20867 | 8eb7eb13-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:51 |  |  |  |  | 0
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 90b7eb13-3709-ed11-85f4-5cf370a20867 | b8bbce4d-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:53 |  |  |  |  | 0
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | f2bbce4d-3709-ed11-85f4-5cf370a20867 | 0c592f87-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:54 |  |  |  |  | 0
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 0e592f87-3709-ed11-85f4-5cf370a20867 | 4f3e36c2-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | Aborted | TargetAborted | 7162 (Warning, Quantity not sufficient), 1000 (Error, User Aborted Test) |  |  | 0 | [25 empty] | 2022-07-21 15:56 |  |  |  |  | 0
Normalized
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 0d5c9e96-1209-ed11-85f4-5cf370a20867 | 8dc854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.3 | 2650.56 | 8.31 | 29213 | [20 empty] | 010081427802027710118705 2105880 17240328 | 8 | 3 | V790 | 2022-07-21 11:32 | 1436 | 266 | 7.9055 | 3994.9491 | 3 | 12 | 30.75 | 18.64 | 31.73 | 12.43 | 2.52 | -3.18 | -10.49 | -42.19 | -36.9 | -26.8 | 1.89 | 70.99 | 208.48 | 498.77 | 1102.87 | 2284.76 | 4226.46 | 6867.15 | 10024.65 | 13327.94 | 16438.04 | 19023.73 | 21063.03 | 22604.32 | 23803.21 | 24728.11 | 25494.6 | 26121.5 | 26612.19 | 27049.29 | 27401.98 | 27680.68 | 27918.97 | 28169.27 | 28370.36 | 28551.25 | 28678.15 | 28804.04 | 28894.54 | 28953.83 | 29021.73 | 29085.62 | 29136.92 | 29171.61 | 29212.71
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 145c9e96-1209-ed11-85f4-5cf370a20867 | 2fc854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.24 | 2705.51 | 8.3 | 29527 | [20 empty] | 010081427802027710118705 2105880 17240328 | 9 | 3 | V790 | 2022-07-21 11:32 | 1469 | 268 | 12.2582 | 4050.6436 | 3 | 12 | 50.5 | 39.44 | 27.78 | 16.52 | 4.07 | -8.19 | -20.25 | -29.11 | -30.57 | -26.03 | -0.08 | 65.86 | 212.8 | 511.74 | 1147.08 | 2386.23 | 4402.97 | 7126.51 | 10356.05 | 13727.39 | 16850.73 | 19416.68 | 21428.02 | 22948.96 | 24116.5 | 25034.84 | 25809.39 | 26406.93 | 26921.27 | 27361.01 | 27703.15 | 27999.29 | 28290.04 | 28513.58 | 28705.72 | 28896.46 | 29046.8 | 29153.95 | 29269.69 | 29366.83 | 29446.77 | 29481.31 | 29521.45 | 29528 | 29526.54
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 1b5c9e96-1209-ed11-85f4-5cf370a20867 | 78ca54cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.2 | 2626.87 | 8.51 | 28852 | [20 empty] | 010081427802027710118705 2105880 17240328 | 10 | 3 | V790 | 2022-07-21 11:32 | 1288 | 265 | 13.2618 | 3841.6764 | 3 | 12 | 64.86 | 47.4 | 32.54 | 13.88 | 3.01 | -10.25 | -17.31 | -30.77 | -34.23 | -27.09 | 3.04 | 67.18 | 213.72 | 518.66 | 1153 | 2380.53 | 4363.07 | 7036.41 | 10171.35 | 13438.89 | 16439.43 | 18934.96 | 20878.7 | 22371.44 | 23536.38 | 24447.92 | 25179.25 | 25775.99 | 26296.53 | 26676.67 | 27060.61 | 27384.75 | 27664.68 | 27869.02 | 28074.56 | 28234.7 | 28382.84 | 28495.57 | 28597.31 | 28722.85 | 28772.19 | 28816.13 | 28847.47 | 28866.4 | 28852.34
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 225c9e96-1209-ed11-85f4-5cf370a20867 | 17ca54cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.22 | 2575.04 | 8.24 | 28421 | [20 empty] | 010081427802027710118705 2105880 17240328 | 11 | 3 | V790 | 2022-07-21 11:32 | 1419 | 259 | 12.4945 | 3931.0509 | 3 | 12 | 65.45 | 48.16 | 32.47 | 14.37 | -0.32 | -11.62 | -19.51 | -25.81 | -27.9 | -25.2 | -0.29 | 63.81 | 201.92 | 498.63 | 1123.33 | 2320.24 | 4261.74 | 6876.25 | 9956.95 | 13176.86 | 16148.76 | 18630.47 | 20561.57 | 22047.28 | 23190.79 | 24116.29 | 24830.8 | 25428.7 | 25913.01 | 26327.71 | 26675.82 | 26997.52 | 27253.23 | 27477.93 | 27678.24 | 27819.35 | 27938.85 | 28067.96 | 28162.86 | 28220.17 | 28297.67 | 28350.18 | 28375.48 | 28398.39 | 28421.09
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 295c9e96-1209-ed11-85f4-5cf370a20867 | b5c954cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.25 | 2811.07 | 8.27 | 29297 | [20 empty] | 010081427802027710118705 2105880 17240328 | 12 | 3 | V790 | 2022-07-21 11:32 | 1453 | 244 | 11.0388 | 4033.9891 | 3 | 12 | 50.97 | 32.73 | 17.29 | 25.06 | 6.62 | -9.82 | -15.66 | -17.7 | -43.74 | -31.38 | -0.42 | 69.75 | 208.71 | 521.07 | 1152.03 | 2393.39 | 4445.35 | 7267.31 | 10545.47 | 14003.04 | 17183.4 | 19801.76 | 21789.72 | 23311.88 | 24449.44 | 25356.8 | 26061.56 | 26625.92 | 27116.69 | 27510.65 | 27831.61 | 28104.97 | 28364.33 | 28518.29 | 28681.05 | 28830.61 | 28942.58 | 29009.94 | 29115.3 | 29182.06 | 29210.42 | 29237.98 | 29271.34 | 29282.7 | 29296.67
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 305c9e96-1209-ed11-85f4-5cf370a20867 | 1255c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.65 | 2107.83 | 6.31 | 23255 | 6.73 | 6.55 | 5370317.96 | 3542261.73 | 6.73 | 6.55 |  |  | 5370317.96 | 3542261.73 |  |  | 6.73 | 6.55 | 5370317.96 | 3542261.73 | 6.73 | 6.55 | 5370317.96 | 3542261.73 | 010081427802027710118705 2105879 17240328 | 1 | 4 | V1169 | 2022-07-21 11:32 | 2573 | 296 | 10.5939 | 4394.4255 | 3 | 12 | 48.58 | 33.39 | 18.79 | 7.2 | 3.2 | -6.59 | -12.18 | -15.98 | -17.77 | -17.16 | 2.84 | 37.65 | 119.45 | 296.66 | 677.67 | 1441.67 | 2778.28 | 4706.28 | 7095.89 | 9727.5 | 12281.7 | 14514.31 | 16304.11 | 17680.52 | 18740.93 | 19580.33 | 20223.14 | 20755.54 | 21178.95 | 21519.36 | 21808.76 | 22047.17 | 22228.17 | 22426.78 | 22580.39 | 22718.79 | 22836 | 22937.2 | 22994.81 | 23079.02 | 23136.62 | 23204.83 | 23240.64 | 23262.84 | 23254.85
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 375c9e96-1209-ed11-85f4-5cf370a20867 | b354c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.42 | 1950.48 | 6.9 | 23336 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 |  |  | 6309573.44 | 4161794.64 |  |  | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 010081427802027710118705 2105879 17240328 | 2 | 4 | V1169 | 2022-07-21 11:32 | 2058 | 284 | 10.0558 | 3971.7818 | 3 | 12 | 42.16 | 29.91 | 20.25 | 10.2 | 3.94 | -4.92 | -13.97 | -21.43 | -23.08 | -18.94 | 0.8 | 47.15 | 140.09 | 340.84 | 762.98 | 1598.93 | 2975.47 | 4889.21 | 7191.76 | 9675.5 | 12049.25 | 14119.39 | 15787.14 | 17117.88 | 18181.42 | 19027.17 | 19714.51 | 20298.06 | 20764.2 | 21146.15 | 21477.89 | 21745.83 | 21990.78 | 22225.12 | 22408.87 | 22574.81 | 22724.16 | 22855.1 | 22964.84 | 23060.79 | 23148.73 | 23226.28 | 23272.02 | 23301.96 | 23335.71
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 3e5c9e96-1209-ed11-85f4-5cf370a20867 | 7255c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.3 | 2023.88 | 6.93 | 22748 | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 6.84 | 6.66 |  |  | 6918309.71 | 4563317.08 |  |  | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 010081427802027710118705 2105879 17240328 | 3 | 4 | V1169 | 2022-07-21 11:32 | 1979 | 281 | 10.2691 | 3858.8618 | 3 | 12 | 56.07 | 38.8 | 23.13 | 8.26 | -0.21 | -10.48 | -11.75 | -16.61 | -18.68 | -15.75 | 2.78 | 39.31 | 135.24 | 344.37 | 788.7 | 1664.63 | 3124.36 | 5128.49 | 7501.03 | 10012.76 | 12393.09 | 14379.42 | 15949.55 | 17194.48 | 18165.61 | 18916.14 | 19532.47 | 20042 | 20441.13 | 20791.27 | 21116.2 | 21378.53 | 21608.66 | 21805.39 | 21972.32 | 22097.65 | 22221.38 | 22317.51 | 22434.84 | 22514.57 | 22587.71 | 22645.24 | 22701.57 | 22720.5 | 22748.23
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 6042e1f8-1209-ed11-85f4-5cf370a20867 | 233d8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.34 | 2623.16 | 6.33 | 27960 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 |  |  | 6760829.75 | 4459443.3 |  |  | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 010081427802027710118705 2105881 17240328 | 12 | 2 | V834 | 2022-07-21 11:34 | 2839 | 265 | 11.2448 | 5279.6436 | 3 | 12 | 76.71 | 55.07 | 37.22 | 15.98 | -2.07 | -15.71 | -22.76 | -29.2 | -28.65 | -22.09 | 6.66 | 60.62 | 184.37 | 448.93 | 1004.88 | 2105.44 | 3966.79 | 6533.95 | 9593.7 | 12836.66 | 15897.61 | 18442.57 | 20415.32 | 21923.68 | 23079.64 | 23974.99 | 24670.95 | 25249.7 | 25729.26 | 26124.01 | 26435.17 | 26729.12 | 26957.88 | 27135.03 | 27312.79 | 27452.34 | 27558.3 | 27650.85 | 27737.41 | 27807.56 | 27859.92 | 27907.87 | 27945.63 | 27967.58 | 27959.54
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 6642e1f8-1209-ed11-85f4-5cf370a20867 | 7e8c0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.35 | 2535.4 | 8.04 | 28768 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 |  |  | 6606934.48 | 4357933.98 |  |  | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 010081427802027710118705 2105880 17240328 | 1 | 3 | V790 | 2022-07-21 11:34 | 1595 | 256 | 12.2133 | 4091.92 | 3 | 12 | 58.07 | 42.65 | 27.04 | 12.03 | 2.01 | -6.4 | -18.41 | -23.23 | -28.44 | -21.65 | -1.67 | 58.72 | 187.31 | 462.09 | 1025.08 | 2128.07 | 3953.25 | 6462.84 | 9459.03 | 12643.21 | 15640.4 | 18161.19 | 20133.17 | 21653.96 | 22841.35 | 23789.13 | 24543.32 | 25186.51 | 25721.89 | 26172.48 | 26560.27 | 26906.25 | 27206.24 | 27469.83 | 27688.61 | 27884.2 | 28045.79 | 28180.97 | 28310.36 | 28426.55 | 28520.73 | 28604.92 | 28684.31 | 28731.89 | 28767.68
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 6c42e1f8-1209-ed11-85f4-5cf370a20867 | ef8b0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.32 | 2588.44 | 7.64 | 29845 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 |  |  | 6760829.75 | 4459443.3 |  |  | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 010081427802027710118705 2105880 17240328 | 2 | 3 | V790 | 2022-07-21 11:34 | 1882 | 264 | 9.5576 | 4509.8982 | 3 | 12 | 63.74 | 47.99 | 33.43 | 14.07 | -1.89 | -8.84 | -18.6 | -31.36 | -29.12 | -21.87 | 2.37 | 61.81 | 198.45 | 474.1 | 1056.34 | 2199.78 | 4080.42 | 6654.27 | 9719.91 | 12970.15 | 16037.39 | 18641.84 | 20704.08 | 22326.92 | 23599.56 | 24613.6 | 25427.65 | 26100.29 | 26666.53 | 27158.97 | 27591.42 | 27951.66 | 28275.3 | 28533.74 | 28762.99 | 28955.63 | 29116.47 | 29251.51 | 29390.16 | 29499 | 29592.44 | 29691.88 | 29770.53 | 29815.57 | 29845.41
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 7242e1f8-1209-ed11-85f4-5cf370a20867 | bf8b0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.37 | 2773.69 | 8.52 | 30635 | [20 empty] | 010081427802027710118705 2105880 17240328 | 3 | 3 | V790 | 2022-07-21 11:34 | 1410 | 242 | 14.6582 | 4079.0836 | 3 | 12 | 50.86 | 39.2 | 27.74 | 15.28 | 3.03 | -9.03 | -19.89 | -27.55 | -28.61 | -22.87 | 2.28 | 59.62 | 191.36 | 468.7 | 1053.44 | 2215.59 | 4171.73 | 6883.07 | 10120.41 | 13563.75 | 16801.49 | 19512.24 | 21607.18 | 23217.92 | 24450.86 | 25420.8 | 26191.95 | 26845.69 | 27398.03 | 27880.37 | 28294.51 | 28664.05 | 28984.4 | 29242.34 | 29474.48 | 29688.42 | 29866.96 | 30019.91 | 30171.85 | 30278.19 | 30377.33 | 30464.87 | 30541.21 | 30589.36 | 30634.7
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 7842e1f8-1209-ed11-85f4-5cf370a20867 | 608b0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.46 | 2532.36 | 8.41 | 30118 | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 6.79 | 6.61 |  |  | 6165950.02 | 4067060.63 |  |  | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 010081427802027710118705 2105880 17240328 | 4 | 3 | V790 | 2022-07-21 11:34 | 1354 | 259 | 12.6812 | 4069.7309 | 3 | 12 | 52.19 | 39.11 | 25.43 | 12.74 | 0.86 | -9.62 | -17.1 | -21.98 | -21.86 | -19.54 | -1.02 | 52.09 | 175.81 | 429.73 | 970.25 | 2030.57 | 3799.89 | 6260.81 | 9247.53 | 12449.44 | 15538.96 | 18212.08 | 20377.8 | 22069.52 | 23424.44 | 24480.16 | 25333.88 | 26030.4 | 26634.51 | 27126.63 | 27573.35 | 27949.87 | 28277.39 | 28565.31 | 28840.43 | 29074.35 | 29279.66 | 29466.98 | 29608.3 | 29721.22 | 29829.94 | 29928.86 | 30003.78 | 30068.7 | 30117.61
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 7e42e1f8-1209-ed11-85f4-5cf370a20867 | ebaadb1d-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified | 2610 (Informational, Delays in sample processing) | 16.12 | 1603.89 | 6.52 | 16922 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 |  |  | 7762471.17 | 5120125.98 |  |  | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 010081427802027710118705 2105880 17240328 | 5 | 3 | V790 | 2022-07-21 11:34 | 1372 | 268 | 2.2497 | 3104.4473 | 3 | 12 | 1.7 | 3.65 | 6.8 | 9.15 | 10.9 | 2.65 | -6.6 | -16.64 | -27.69 | -20.94 | -2.19 | 44.56 | 140.11 | 347.06 | 739.41 | 1487.36 | 2714.91 | 4370.06 | 6232.21 | 8178.76 | 10006.31 | 11462.66 | 12548.61 | 13433.96 | 14100.71 | 14603.06 | 15047.61 | 15419.96 | 15711.71 | 15959.46 | 16122.61 | 16261.96 | 16335.71 | 16431.26 | 16530.21 | 16608.76 | 16638.51 | 16720.46 | 16731.21 | 16637.76 | 16583.12 | 16617.27 | 16683.82 | 16764.37 | 16921.72
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 8442e1f8-1209-ed11-85f4-5cf370a20867 | 57c954cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.29 | 2625.53 | 8.37 | 30075 | [20 empty] | 010081427802027710118705 2105880 17240328 | 6 | 3 | V790 | 2022-07-21 11:34 | 1388 | 245 | 11.3673 | 4079.7455 | 3 | 12 | 35.89 | 24.92 | 12.75 | 22.79 | 9.22 | -4.35 | -15.72 | -17.68 | -41.65 | -30.42 | -0.39 | 65.45 | 202.68 | 500.51 | 1101.15 | 2274.38 | 4198.41 | 6830.84 | 9933.28 | 13232.31 | 16333.74 | 18948.57 | 21009.01 | 22622.24 | 23879.67 | 24870.71 | 25696.54 | 26362.57 | 26939.4 | 27427.44 | 27836.27 | 28159.1 | 28487.53 | 28739.37 | 28967.2 | 29181.83 | 29359.27 | 29482.9 | 29613.13 | 29722.16 | 29811.2 | 29893.23 | 29977.06 | 30036.49 | 30074.93
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 8a42e1f8-1209-ed11-85f4-5cf370a20867 | edc854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.13 | 2178.33 | 8.18 | 26676 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 |  |  | 7762471.17 | 5120125.98 |  |  | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 010081427802027710118705 2105880 17240328 | 7 | 3 | V790 | 2022-07-21 11:34 | 1317 | 270 | 11.3709 | 3719.5782 | 3 | 12 | 47.25 | 34.08 | 24.91 | 14.14 | 6.37 | -7.6 | -16.37 | -27.35 | -30.72 | -26.89 | 0.94 | 62.57 | 195.6 | 482.23 | 1075.66 | 2178.09 | 3904.52 | 6188.95 | 8851.97 | 11631.2 | 14223.83 | 16418.26 | 18218.09 | 19657.32 | 20810.35 | 21727.38 | 22502.01 | 23113.64 | 23648.07 | 24086.49 | 24480.32 | 24807.35 | 25108.98 | 25341.61 | 25569.84 | 25761.07 | 25925.9 | 26058.93 | 26204.96 | 26330.99 | 26430.21 | 26522.64 | 26611.47 | 26655.9 | 26675.73
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 5a7b815a-1309-ed11-85f4-5cf370a20867 | 223c8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.28 | 2631.98 | 6.57 | 27149 | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 6.85 | 6.67 |  |  | 7079457.84 | 4669610.39 |  |  | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 010081427802027710118705 2105881 17240328 | 7 | 2 | V834 | 2022-07-21 11:37 | 2538 | 268 | 11.423 | 4907.5673 | 3 | 12 | 72.01 | 55.59 | 35.56 | 15.94 | 1.52 | -13.91 | -25.73 | -28.95 | -29.17 | -22.4 | 1.98 | 65.16 | 198.93 | 476.71 | 1068.09 | 2227.06 | 4143.64 | 6766.42 | 9867.8 | 13117.77 | 16119.55 | 18576.13 | 20444.7 | 21824.48 | 22870.26 | 23682.43 | 24321.41 | 24812.19 | 25224.16 | 25552.74 | 25817.52 | 26046.1 | 26249.27 | 26437.85 | 26561.63 | 26689.6 | 26788.98 | 26878.16 | 26936.93 | 26991.71 | 27033.69 | 27069.27 | 27099.04 | 27115.42 | 27148.8
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 607b815a-1309-ed11-85f4-5cf370a20867 | 50597999-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.37 | 2314.8 | 6.2 | 25184 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 |  |  | 6606934.48 | 4357933.98 |  |  | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 010081427802027710118705 2105881 17240328 | 8 | 2 | V834 | 2022-07-21 11:37 | 2749 | 292 | 8.6594 | 4878.6545 | 3 | 12 | 57.49 | 44.63 | 29.57 | 13.51 | -0.15 | -10.21 | -19.67 | -25.53 | -25.59 | -19.05 | 3.49 | 53.63 | 165.17 | 396.91 | 880.45 | 1844.4 | 3475.74 | 5731.68 | 8428.62 | 11307.56 | 14020.9 | 16287.64 | 18056.58 | 19435.32 | 20526.86 | 21368 | 22059.34 | 22630.08 | 23082.42 | 23426.96 | 23753.9 | 24015.04 | 24223.99 | 24416.53 | 24582.67 | 24684.61 | 24776.95 | 24881.69 | 24932.43 | 24975.37 | 25035.31 | 25080.05 | 25110.59 | 25154.13 | 25184.07
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 667b815a-1309-ed11-85f4-5cf370a20867 | 9c3e8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.45 | 2392.68 | 5.45 | 25142 | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 6.79 | 6.61 |  |  | 6165950.02 | 4067060.63 |  |  | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 010081427802027710118705 2105881 17240328 | 9 | 2 | V834 | 2022-07-21 11:37 | 3532 | 272 | 7.1382 | 5690.0436 | 3 | 12 | 55.22 | 40.68 | 25.34 | 11.8 | 3.27 | -8.27 | -17.01 | -23.35 | -26.49 | -18.83 | -0.16 | 53.7 | 162.96 | 388.22 | 862.08 | 1822.75 | 3453.21 | 5751.27 | 8523.53 | 11483.39 | 14298.85 | 16649.12 | 18474.78 | 19839.24 | 20882.3 | 21658.56 | 22292.43 | 22787.29 | 23181.95 | 23517.41 | 23814.07 | 24045.73 | 24231.8 | 24421.46 | 24541.52 | 24660.98 | 24741.04 | 24818.11 | 24875.77 | 24935.43 | 24958.69 | 25016.75 | 25070.61 | 25106.48 | 25141.94
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 6c7b815a-1309-ed11-85f4-5cf370a20867 | 3e3e8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.38 | 2436.69 | 5.53 | 25415 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 |  |  | 6606934.48 | 4357933.98 |  |  | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 010081427802027710118705 2105881 17240328 | 10 | 2 | V834 | 2022-07-21 11:37 | 3437 | 275 | 7.8133 | 5650.56 | 3 | 12 | 49.03 | 42.41 | 29 | 17.79 | 0.37 | -9.04 | -23.45 | -26.07 | -30.48 | -21.69 | 4.69 | 58.88 | 168.47 | 410.05 | 912.64 | 1911.63 | 3612.21 | 5979.6 | 8809.19 | 11806.97 | 14630.96 | 16963.35 | 18763.33 | 20099.72 | 21126.11 | 21895.29 | 22520.68 | 23027.07 | 23443.45 | 23800.84 | 24088.83 | 24315.21 | 24497.2 | 24677.59 | 24802.17 | 24921.56 | 25027.95 | 25115.93 | 25174.32 | 25224.11 | 25288.49 | 25347.28 | 25378.87 | 25400.65 | 25414.64
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 727b815a-1309-ed11-85f4-5cf370a20867 | e03d8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.44 | 2298.39 | 5.71 | 25080 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 |  |  | 6309573.44 | 4161794.64 |  |  | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 010081427802027710118705 2105881 17240328 | 11 | 2 | V834 | 2022-07-21 11:37 | 3289 | 280 | 6.577 | 5366.3527 | 3 | 12 | 63.07 | 50.29 | 33.12 | 14.14 | -0.44 | -10.01 | -21.39 | -28.17 | -31.35 | -20.32 | 6.3 | 58.12 | 166.15 | 393.77 | 859.79 | 1794.22 | 3370.64 | 5591.46 | 8265.68 | 11106.71 | 13810.93 | 16101.95 | 17902.58 | 19296 | 20391.22 | 21245.85 | 21934.47 | 22476.69 | 22922.52 | 23318.94 | 23618.36 | 23882.58 | 24132.61 | 24307.03 | 24425.05 | 24559.28 | 24677.5 | 24741.92 | 24827.35 | 24913.57 | 24973.99 | 25002.01 | 25043.84 | 25069.06 | 25080.08
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 787b815a-1309-ed11-85f4-5cf370a20867 | 1eff3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.43 | 2367.47 | 6.98 | 25496 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 |  |  | 6309573.44 | 4161794.64 |  |  | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 010081427802027710118705 2105881 17240328 | 4 | 2 | V834 | 2022-07-21 11:37 | 2183 | 247 | 10.5467 | 4277.56 | 3 | 12 | 55.09 | 41.55 | 26.6 | 14.25 | 0.91 | -10.64 | -19.19 | -21.33 | -24.88 | -19.63 | -0.77 | 54.68 | 162.13 | 390.99 | 874.64 | 1838.09 | 3465.75 | 5758.4 | 8506.45 | 11445.31 | 14229.96 | 16555.81 | 18362.87 | 19778.12 | 20856.77 | 21696.23 | 22379.08 | 22948.13 | 23395.79 | 23771.24 | 24087.89 | 24337.75 | 24537.4 | 24691.25 | 24846.91 | 24960.36 | 25086.01 | 25176.67 | 25271.32 | 25333.37 | 25412.43 | 25451.28 | 25486.13 | 25497.99 | 25495.84
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 7e7b815a-1309-ed11-85f4-5cf370a20867 | 823d8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.4 | 2435.14 | 6.91 | 25271 | 6.81 | 6.63 | 6456542.29 | 4258735.29 | 6.81 | 6.63 |  |  | 6456542.29 | 4258735.29 |  |  | 6.81 | 6.63 | 6456542.29 | 4258735.29 | 6.81 | 6.63 | 6456542.29 | 4258735.29 | 010081427802027710118705 2105881 17240328 | 5 | 2 | V834 | 2022-07-21 11:37 | 2203 | 258 | 11.3697 | 4297.6873 | 3 | 12 | 50.14 | 37.97 | 24.8 | 10.83 | 2.26 | -5.51 | -17.28 | -23.24 | -25.21 | -18.78 | -0.15 | 52.28 | 159.91 | 394.14 | 887.57 | 1878.8 | 3567.43 | 5921.66 | 8746.29 | 11752.72 | 14579.55 | 16923.38 | 18741.61 | 20100.64 | 21136.27 | 21936.5 | 22546.53 | 23042.96 | 23440.59 | 23763.62 | 24028.65 | 24277.68 | 24461.31 | 24636.94 | 24769.17 | 24874.6 | 24952.43 | 25036.86 | 25098.89 | 25159.12 | 25193.16 | 25233.99 | 25257.22 | 25268.65 | 25271.08
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 847b815a-1309-ed11-85f4-5cf370a20867 | c43c8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.18 | 2359.93 | 6.51 | 24908 | 6.88 | 6.7 | 7585775.75 | 5003577.68 | 6.88 | 6.7 |  |  | 7585775.75 | 5003577.68 |  |  | 6.88 | 6.7 | 7585775.75 | 5003577.68 | 6.88 | 6.7 | 7585775.75 | 5003577.68 | 010081427802027710118705 2105881 17240328 | 6 | 2 | V834 | 2022-07-21 11:37 | 2433 | 253 | 11.0485 | 4546.5564 | 3 | 12 | 49.2 | 37.35 | 25.7 | 13.05 | 3.6 | -7.65 | -16.5 | -25.54 | -29.19 | -23.84 | 4.91 | 55.46 | 183.01 | 448.76 | 1005.72 | 2092.27 | 3872.62 | 6258.77 | 9051.52 | 11963.27 | 14641.03 | 16819.18 | 18512.13 | 19791.28 | 20774.23 | 21533.78 | 22143.73 | 22642.69 | 23036.44 | 23374.59 | 23632.54 | 23861.89 | 24027.24 | 24201.6 | 24313.55 | 24423.9 | 24512.05 | 24601.6 | 24645.75 | 24697.5 | 24746.06 | 24807.81 | 24842.16 | 24882.71 | 24908.06
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | b3de19db-1309-ed11-85f4-5cf370a20867 | 685f6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.51 | 2123.17 | 6.78 | 22703 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 |  |  | 6025595.86 | 3974483.03 |  |  | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 010081427802027710118705 2105881 17240328 | 3 | 2 | V834 | 2022-07-21 11:41 | 2093 | 263 | 7.3515 | 3941.8836 | 3 | 12 | 38.96 | 27.41 | 19.86 | 9.51 | 0.96 | -6.59 | -10.34 | -19.7 | -19.65 | -16.6 | 0.65 | 41.9 | 135.75 | 333.8 | 748.84 | 1576.09 | 2990.94 | 5007.79 | 7443.64 | 10077.09 | 12615.33 | 14760.58 | 16449.63 | 17751.48 | 18730.13 | 19477.78 | 20070.63 | 20526.87 | 20907.52 | 21217.77 | 21466.62 | 21694.07 | 21870.32 | 22012.16 | 22166.41 | 22277.46 | 22366.31 | 22459.96 | 22523.21 | 22562.06 | 22610.7 | 22651.95 | 22673.8 | 22693.45 | 22702.5
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | b9de19db-1309-ed11-85f4-5cf370a20867 | 785e6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.61 | 2768.06 | 7.99 | 29669 | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 6.75 | 6.57 |  |  | 5623413.25 | 3709203.38 |  |  | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 010081427802027710118705 2105882 17240328 | 10 | 1 | V639 | 2022-07-21 11:41 | 1720 | 268 | 12.9721 | 4239.3891 | 3 | 12 | 42.04 | 30.27 | 20.29 | 10.52 | 1.55 | -2.22 | -12.19 | -21.17 | -26.14 | -19.11 | -2.08 | 50.55 | 164.57 | 402.6 | 903.03 | 1913.26 | 3686.88 | 6231.31 | 9376.94 | 12790.17 | 16090.2 | 18909.02 | 21127.45 | 22831.08 | 24134.51 | 25140.74 | 25944.76 | 26585.39 | 27110.02 | 27528.05 | 27891.08 | 28189.3 | 28459.13 | 28669.76 | 28883.79 | 29054.61 | 29186.44 | 29293.07 | 29379.1 | 29457.13 | 29526.95 | 29594.18 | 29641.41 | 29669.84 | 29669.07
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | bfde19db-1309-ed11-85f4-5cf370a20867 | 485e6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.29 | 2777.15 | 8.24 | 29533 | [20 empty] | 010081427802027710118705 2105882 17240328 | 11 | 1 | V639 | 2022-07-21 11:41 | 1560 | 266 | 14.7564 | 4080.1273 | 3 | 12 | 50.32 | 39.76 | 26 | 18.05 | 5.09 | -8.67 | -21.82 | -26.78 | -31.33 | -26.49 | 1.15 | 64.8 | 198.24 | 487.88 | 1107.33 | 2318.77 | 4334.01 | 7091.26 | 10361.5 | 13784.15 | 16974.79 | 19594.43 | 21619.48 | 23319.72 | 24516.76 | 25452.01 | 26211.45 | 26809.69 | 27145.74 | 27538.18 | 27862.03 | 28135.27 | 28381.11 | 28591.76 | 28965 | 29136.64 | 29263.09 | 29383.13 | 29454.97 | 29354.62 | 29398.86 | 29468.31 | 29503.75 | 29534.19 | 29532.84
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | c5de19db-1309-ed11-85f4-5cf370a20867 | 90fe3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.47 | 2893.71 | 8.04 | 30993 | [20 empty] | 010081427802027710118705 2105882 17240328 | 12 | 1 | V639 | 2022-07-21 11:41 | 1817 | 263 | 16.143 | 4407.3273 | 3 | 12 | 50.13 | 37.99 | 26.64 | 16.5 | 0.36 | -7.59 | -18.33 | -25.27 | -29.61 | -23.96 | 0.7 | 60.56 | 185.41 | 456.47 | 1032.33 | 2186.18 | 4153.84 | 6926.7 | 10279.36 | 13876.01 | 17307.67 | 20202.73 | 22476.98 | 24193.04 | 25519.9 | 26540.55 | 27353.61 | 27974.27 | 28510.12 | 28932.98 | 29283.24 | 29571.1 | 29826.55 | 30024.21 | 30218.47 | 30378.12 | 30491.78 | 30598.84 | 30692.89 | 30743.35 | 30791.61 | 30866.07 | 30921.92 | 30961.78 | 30992.84
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | cbde19db-1309-ed11-85f4-5cf370a20867 | 60fe3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.49 | 2361.31 | 5.95 | 25803 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 |  |  | 6025595.86 | 3974483.03 |  |  | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 010081427802027710118705 2105881 17240328 | 1 | 2 | V834 | 2022-07-21 11:41 | 3061 | 258 | 8.7067 | 5253.04 | 3 | 12 | 39.25 | 30.55 | 24.04 | 16.33 | 6.03 | -10.48 | -19.79 | -27.09 | -27.4 | -21.11 | 5.39 | 54.08 | 152.57 | 373.67 | 845.56 | 1785.25 | 3382.35 | 5644.84 | 8381.73 | 11322.83 | 14151.52 | 16549.21 | 18444.31 | 19904 | 21040.69 | 21921.79 | 22648.28 | 23217.57 | 23676.47 | 24048.36 | 24351.65 | 24589.15 | 24805.44 | 24989.13 | 25143.63 | 25274.12 | 25402.61 | 25492.91 | 25578.8 | 25661.09 | 25709.79 | 25731.88 | 25774.37 | 25796.87 | 25802.96
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | d1de19db-1309-ed11-85f4-5cf370a20867 | 985f6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.29 | 2373.91 | 6.32 | 24689 | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 6.84 | 6.66 |  |  | 6918309.71 | 4563317.08 |  |  | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 010081427802027710118705 2105881 17240328 | 2 | 2 | V834 | 2022-07-21 11:41 | 2647 | 271 | 10.5273 | 4659.0655 | 3 | 12 | 44.81 | 35.88 | 23.75 | 11.43 | 3.3 | -5.43 | -16.76 | -23.28 | -25.61 | -20.34 | -1.27 | 54.21 | 176.88 | 426.55 | 948.83 | 1980.3 | 3693.37 | 6050.44 | 8832.52 | 11759.19 | 14480.46 | 16707.33 | 18419.81 | 19707.08 | 20691.75 | 21429.03 | 22036.1 | 22487.37 | 22884.24 | 23172.92 | 23438.99 | 23664.46 | 23848.53 | 23983.41 | 24139.48 | 24248.75 | 24323.63 | 24429.1 | 24493.17 | 24528.44 | 24593.52 | 24647.39 | 24654.86 | 24676.53 | 24689.41
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | d7de19db-1309-ed11-85f4-5cf370a20867 | 954775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.27 | 2874.99 | 8.13 | 29909 | [20 empty] | 010081427802027710118705 2105882 17240328 | 8 | 1 | V639 | 2022-07-21 11:41 | 1689 | 271 | 15.3721 | 4204.9491 | 3 | 12 | 56.28 | 43.51 | 30.33 | 18.16 | 4.59 | -8.58 | -21.55 | -32.33 | -36.3 | -28.67 | 4.36 | 69.99 | 215.61 | 521.64 | 1166.07 | 2427.3 | 4521.52 | 7390.15 | 10761.38 | 14294.41 | 17551.04 | 20215.86 | 22247.89 | 23781.12 | 24944.75 | 25848.38 | 26570.6 | 27137.63 | 27593.06 | 27980.89 | 28300.32 | 28579.74 | 28832.17 | 29052 | 29225.83 | 29401.85 | 29540.28 | 29651.91 | 29743.54 | 29831.97 | 29862.79 | 29884.42 | 29905.85 | 29919.48 | 29909.11
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | ddde19db-1309-ed11-85f4-5cf370a20867 | db5e6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.06 | 2821.6 | 7.95 | 29515 | [20 empty] | 010081427802027710118705 2105882 17240328 | 9 | 1 | V639 | 2022-07-21 11:41 | 1825 | 274 | 18.3176 | 4254.1782 | 3 | 12 | 71.1 | 54.19 | 37.27 | 20.55 | 0.43 | -12.28 | -26 | -34.52 | -37.24 | -25.55 | 1.73 | 75.61 | 238.29 | 587.98 | 1318.06 | 2723.74 | 4945.02 | 7881.91 | 11260.99 | 14724.27 | 17849.35 | 20369.04 | 22293.72 | 23724.8 | 24819.88 | 25660.76 | 26338.25 | 26863.93 | 27327.41 | 27705.09 | 28016.78 | 28290.86 | 28513.14 | 28689.82 | 28837.71 | 28960.19 | 29050.47 | 29147.15 | 29249.64 | 29301.92 | 29365 | 29435.68 | 29481.77 | 29488.25 | 29514.93
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 65fdae7a-1409-ed11-85f4-5cf370a20867 | 364775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.27 | 2347.62 | 7.47 | 26286 | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 6.85 | 6.67 |  |  | 7079457.84 | 4669610.39 |  |  | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 010081427802027710118705 2105882 17240328 | 4 | 1 | V639 | 2022-07-21 11:45 | 1808 | 260 | 11.5855 | 4075.4291 | 3 | 12 | 41.59 | 32.8 | 23.41 | 14.03 | 4.84 | -6.74 | -17.13 | -24.91 | -29.5 | -21.88 | 1.13 | 56.75 | 183.16 | 443.17 | 985.79 | 2036.2 | 3769.62 | 6121.03 | 8900.85 | 11838.86 | 14607.08 | 16986.69 | 18808.71 | 20235.92 | 21344.13 | 22215.95 | 22841.96 | 23414.98 | 23894.19 | 24276.01 | 24605.02 | 24889.84 | 25141.45 | 25345.07 | 25523.88 | 25672.09 | 25800.51 | 25899.72 | 25982.54 | 26072.75 | 26132.97 | 26185.98 | 26233.6 | 26269.01 | 26286.03
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 6bfdae7a-1409-ed11-85f4-5cf370a20867 | a94675b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.02 | 2809.5 | 7.97 | 30622 | 6.92 | 6.74 | 8317637.71 | 5486313.83 | 6.92 | 6.74 |  |  | 8317637.71 | 5486313.83 |  |  | 6.92 | 6.74 | 8317637.71 | 5486313.83 | 6.92 | 6.74 | 8317637.71 | 5486313.83 | 010081427802027710118705 2105882 17240328 | 5 | 1 | V639 | 2022-07-21 11:45 | 1826 | 257 | 19.0012 | 4401.4909 | 3 | 12 | 70.11 | 50.31 | 35.11 | 17.7 | 3.5 | -11.9 | -21.5 | -32.9 | -36.3 | -31.9 | 1.1 | 77.09 | 249.09 | 615.89 | 1383.49 | 2826.49 | 5089.09 | 8047.89 | 11439.29 | 14906.48 | 18029.28 | 20575.88 | 22538.28 | 24048.28 | 25215.28 | 26166.28 | 26925.08 | 27552.48 | 28101.67 | 28572.47 | 28939.67 | 29291.67 | 29568.67 | 29771.47 | 29960.07 | 30108.07 | 30197.26 | 30313.26 | 30413.86 | 30466.86 | 30516.06 | 30566.06 | 30595.66 | 30612.26 | 30622.05
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 71fdae7a-1409-ed11-85f4-5cf370a20867 | 774675b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.14 | 2540.68 | 7.78 | 28700 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 |  |  | 7762471.17 | 5120125.98 |  |  | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 010081427802027710118705 2105882 17240328 | 6 | 1 | V639 | 2022-07-21 11:45 | 1746 | 262 | 16.9309 | 4242.9382 | 3 | 12 | 70.73 | 53.4 | 34.07 | 16.34 | 3.01 | -12.52 | -24.25 | -30.99 | -29.32 | -25.65 | 1.22 | 68.09 | 214.16 | 515.83 | 1168.5 | 2405.97 | 4377.84 | 7000.91 | 10063.57 | 13231.84 | 16159.11 | 18597.98 | 20504.65 | 21997.72 | 23174.39 | 24121.46 | 24863.73 | 25514.2 | 26027.27 | 26474.93 | 26823.8 | 27147.07 | 27389.34 | 27630.01 | 27821.88 | 27988.75 | 28148.02 | 28293.29 | 28393.56 | 28467.63 | 28549.89 | 28608.16 | 28645.63 | 28675.5 | 28700.37
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 77fdae7a-1409-ed11-85f4-5cf370a20867 | c64775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.03 | 2805.63 | 8.16 | 29973 | [20 empty] | 010081427802027710118705 2105882 17240328 | 7 | 1 | V639 | 2022-07-21 11:45 | 1646 | 261 | 19.9042 | 4183.5382 | 3 | 12 | 73.56 | 55.45 | 36.75 | 17.44 | 2.54 | -12.36 | -23.87 | -33.37 | -33.48 | -30.78 | 0.92 | 76.21 | 248.31 | 610.6 | 1365.1 | 2794.19 | 5037.09 | 7985.79 | 11358.88 | 14819.58 | 17920.07 | 20427.77 | 22349.86 | 23812.56 | 24922.46 | 25813.35 | 26550.05 | 27135.54 | 27641.04 | 28065.13 | 28407.63 | 28666.93 | 28898.82 | 29080.92 | 29233.81 | 29379.31 | 29506.4 | 29624.7 | 29706.6 | 29779.09 | 29847.39 | 29910.28 | 29935.78 | 29963.48 | 29972.97
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | f47fc1ae-1609-ed11-85f4-5cf370a20867 | 816b47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.14 | 2841.9 | 7.99 | 29636 | [20 empty] | 010081427802027710118705 2105882 17240328 | 1 | 1 | V639 | 2022-07-21 12:01 | 1679 | 254 | 18.6121 | 4231.3491 | 3 | 12 | 62.44 | 45.43 | 31.01 | 16 | 1.39 | -8.22 | -20.83 | -29.05 | -31.66 | -26.07 | -0.08 | 67.51 | 219.29 | 547.08 | 1236.47 | 2572.26 | 4740.04 | 7641.83 | 11008.22 | 14485.21 | 17656.8 | 20221.38 | 22181.77 | 23655.16 | 24786.95 | 25655.74 | 26369.32 | 26926.11 | 27385.1 | 27779.49 | 28119.48 | 28390.26 | 28614.65 | 28813.44 | 28970.03 | 29096.21 | 29218.6 | 29319.59 | 29407.18 | 29476.17 | 29551.95 | 29601.34 | 29633.13 | 29643.12 | 29635.91
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | fa7fc1ae-1609-ed11-85f4-5cf370a20867 | 226b47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.12 | 3139.19 | 8.24 | 32203 | [20 empty] | 010081427802027710118705 2105882 17240328 | 2 | 1 | V639 | 2022-07-21 12:01 | 1768 | 263 | 19.5394 | 4452.2145 | 3 | 12 | 69.25 | 53.31 | 38.37 | 22.43 | 0.89 | -11.85 | -26.59 | -36.53 | -41.87 | -28.01 | 0.85 | 82.31 | 256.37 | 636.43 | 1411.49 | 2904.36 | 5318.62 | 8545.68 | 12260.54 | 16096 | 19568.66 | 22343.12 | 24415.58 | 25967.44 | 27137.1 | 28055.96 | 28788.82 | 29387.08 | 29875.34 | 30258.6 | 30563.26 | 30833.52 | 31058.59 | 31259.25 | 31447.91 | 31614.97 | 31748.23 | 31868.29 | 31970.75 | 32041.21 | 32087.67 | 32145.93 | 32181.59 | 32192.45 | 32202.91
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | 0080c1ae-1609-ed11-85f4-5cf370a20867 | e16b47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.13 | 3191.32 | 8.32 | 32904 | [20 empty] | 010081427802027710118705 2105882 17240328 | 3 | 1 | V639 | 2022-07-21 12:01 | 1736 | 265 | 18.4218 | 4505.0364 | 3 | 12 | 80.34 | 60.72 | 38.3 | 19.28 | 5.85 | -11.77 | -27.99 | -35.81 | -40.63 | -34.65 | 5.12 | 82.3 | 256.08 | 622.26 | 1397.04 | 2903.81 | 5347.99 | 8604.57 | 12390.75 | 16275.13 | 19791.91 | 22620.28 | 24763.06 | 26352.24 | 27573.22 | 28518.8 | 29282.97 | 29904.35 | 30437.33 | 30877.71 | 31247.89 | 31560.07 | 31796.64 | 31980.42 | 32148.2 | 32282.78 | 32378.56 | 32502.93 | 32603.91 | 32671.69 | 32737.47 | 32803.05 | 32845.83 | 32878.4 | 32903.58
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | a79dad68-1f09-ed11-85f4-5cf370a20867 | b26dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | ad9dad68-1f09-ed11-85f4-5cf370a20867 | b66dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | b39dad68-1f09-ed11-85f4-5cf370a20867 | ba6dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | b99dad68-1f09-ed11-85f4-5cf370a20867 | be6dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | bf9dad68-1f09-ed11-85f4-5cf370a20867 | c26dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | c59dad68-1f09-ed11-85f4-5cf370a20867 | c76dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | cb9dad68-1f09-ed11-85f4-5cf370a20867 | 2f0bca23-2809-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 17.03 | 2397.7 | 7.16 | 25393 | 6.62 | 6.44 | 4168693.83 | 2749670.45 | 6.62 | 6.44 |  |  | 4168693.83 | 2749670.45 |  |  | 6.62 | 6.44 | 4168693.83 | 2749670.45 | 6.62 | 6.44 | 4168693.83 | 2749670.45 | 010081427802027710118705 2105883 17240328 | 11 | 4 | V1169 | 2022-07-21 13:03 | 2048 | 285 | 11.6691 | 4140.7564 | 3 | 13 | 40.77 | 32.51 | 24.24 | 19.77 | 6.1 | -6.77 | -12.84 | -18.51 | -28.58 | -32.65 | -22.12 | 4.41 | 66.95 | 212.88 | 523.01 | 1157.14 | 2362.87 | 4280.4 | 6784.33 | 9651.46 | 12575.59 | 15205.12 | 17309.05 | 18936.39 | 20151.52 | 21096.45 | 21855.38 | 22447.11 | 22910.64 | 23309.17 | 23655.7 | 23925.43 | 24162.16 | 24363.49 | 24538.83 | 24677.16 | 24813.89 | 24925.02 | 25024.15 | 25126.28 | 25187.41 | 25243.94 | 25316.87 | 25372.8 | 25392.73
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | d29dad68-1f09-ed11-85f4-5cf370a20867 | d10aca23-2809-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.59 | 2333.73 | 6.98 | 25183 | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 6.75 | 6.57 |  |  | 5623413.25 | 3709203.38 |  |  | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 010081427802027710118705 2105883 17240328 | 12 | 4 | V1169 | 2022-07-21 13:03 | 2088 | 278 | 10.3709 | 4218.2182 | 3 | 12 | 33.81 | 24.04 | 12.87 | 8.7 | 2.93 | -1.44 | -11.61 | -11.79 | -17.76 | -16.93 | -4.9 | 39.93 | 127.36 | 324.79 | 740.02 | 1584.25 | 3066.48 | 5220.51 | 7847.73 | 10709.56 | 13470.99 | 15827.82 | 17657.25 | 19070.88 | 20145.91 | 20988.34 | 21675.97 | 22224 | 22677.63 | 23067.45 | 23414.08 | 23687.91 | 23927.34 | 24151.17 | 24336.4 | 24494.83 | 24651.06 | 24781.49 | 24883.92 | 24988.95 | 25055.77 | 25103.2 | 25156.63 | 25179.46 | 25182.89
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 1f42d920-2009-ed11-85f4-5cf370a20867 | 41e5bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.64 | 2405.59 | 7.37 | 25353 | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 6.74 | 6.56 |  |  | 5495408.74 | 3624771.6 |  |  | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 010081427802027710118705 2105883 17240328 | 5 | 4 | V1169 | 2022-07-21 13:09 | 1954 | 288 | 13.0194 | 3986.2545 | 3 | 12 | 46.33 | 34.71 | 21.49 | 9.27 | 1.45 | -6.97 | -13.59 | -16.01 | -20.83 | -20.05 | -1.27 | 46.51 | 144.69 | 358.87 | 802.85 | 1682.04 | 3206.82 | 5420.8 | 8151.18 | 11128.96 | 13982.74 | 16428.92 | 18329.7 | 19750.48 | 20810.86 | 21667.84 | 22306.42 | 22847.2 | 23276.78 | 23625.56 | 23905.14 | 24167.32 | 24365.11 | 24559.89 | 24732.07 | 24839.25 | 24956.63 | 25029.41 | 25094.99 | 25128.77 | 25202.75 | 25247.93 | 25300.11 | 25322.69 | 25353.47
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 2642d920-2009-ed11-85f4-5cf370a20867 | b3e4bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.08 | 2303.83 | 7.28 | 24419 | 6.91 | 6.73 | 8128305.16 | 5361430.08 | 6.91 | 6.73 |  |  | 8128305.16 | 5361430.08 |  |  | 6.91 | 6.73 | 8128305.16 | 5361430.08 | 6.91 | 6.73 | 8128305.16 | 5361430.08 | 010081427802027710118705 2105883 17240328 | 6 | 4 | V1169 | 2022-07-21 13:09 | 1944 | 278 | 15.3091 | 3909.1018 | 3 | 12 | 67.79 | 49.68 | 32.77 | 16.06 | 3.75 | -12.96 | -22.87 | -27.17 | -33.28 | -25.99 | 2.9 | 66.79 | 203.48 | 501.77 | 1106.06 | 2259.35 | 4080.84 | 6481.53 | 9262.23 | 12115.72 | 14685.81 | 16762.5 | 18349.39 | 19546.48 | 20443.37 | 21166.46 | 21730.35 | 22191.84 | 22558.33 | 22881.23 | 23145.32 | 23386.81 | 23581.5 | 23752.39 | 23890.08 | 23992.37 | 24095.66 | 24177.15 | 24243.64 | 24316.13 | 24363.03 | 24379.12 | 24411.61 | 24427.7 | 24419.19
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 2c42d920-2009-ed11-85f4-5cf370a20867 | bde6bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.63 | 2270.38 | 7.03 | 24694 | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 6.74 | 6.56 |  |  | 5495408.74 | 3624771.6 |  |  | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 010081427802027710118705 2105883 17240328 | 7 | 4 | V1169 | 2022-07-21 13:09 | 2123 | 298 | 10.1176 | 4103.1782 | 3 | 12 | 46.3 | 33.39 | 21.07 | 9.55 | 3.63 | -4.08 | -13.6 | -22.12 | -25.04 | -17.75 | 0.13 | 48.21 | 145.89 | 347.38 | 759.46 | 1594.74 | 3044.62 | 5133.31 | 7716.39 | 10532.47 | 13243.95 | 15574.64 | 17409.12 | 18799.8 | 19879.88 | 20742.16 | 21412.25 | 21961.33 | 22414.61 | 22797.69 | 23091.18 | 23355.06 | 23595.14 | 23811.42 | 23955.71 | 24099.39 | 24230.27 | 24326.15 | 24388.04 | 24481.12 | 24559.6 | 24609.08 | 24653.17 | 24690.05 | 24693.93
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 3242d920-2009-ed11-85f4-5cf370a20867 | 5ee6bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.32 | 2201.98 | 6.92 | 24142 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 |  |  | 6760829.75 | 4459443.3 |  |  | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 010081427802027710118705 2105883 17240328 | 8 | 4 | V1169 | 2022-07-21 13:09 | 2146 | 296 | 13.8606 | 4100.0255 | 3 | 12 | 57.71 | 41.65 | 25.59 | 13.33 | -0.93 | -6.99 | -18.25 | -19.31 | -26.17 | -20.23 | -1.09 | 54.05 | 165.99 | 410.93 | 903.07 | 1864.8 | 3461.54 | 5648.48 | 8245.22 | 10996.76 | 13561.9 | 15694.64 | 17367.98 | 18634.32 | 19603.06 | 20390.2 | 21009.14 | 21503.08 | 21923.22 | 22284.96 | 22578.1 | 22830.24 | 23060.57 | 23255.11 | 23435.25 | 23583.39 | 23722.33 | 23807.07 | 23886.41 | 23959.95 | 24009.69 | 24045.23 | 24102.37 | 24140.71 | 24141.65
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 3842d920-2009-ed11-85f4-5cf370a20867 | d2e5bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.73 | 2195.79 | 7.21 | 24578 | 6.71 | 6.53 | 5128613.84 | 3382833.69 | 6.71 | 6.53 |  |  | 5128613.84 | 3382833.69 |  |  | 6.71 | 6.53 | 5128613.84 | 3382833.69 | 6.71 | 6.53 | 5128613.84 | 3382833.69 | 010081427802027710118705 2105883 17240328 | 9 | 4 | V1169 | 2022-07-21 13:09 | 1973 | 289 | 7.6679 | 3971.2109 | 3 | 12 | 43.52 | 33.65 | 20.59 | 8.52 | 0.05 | -6.02 | -15.49 | -15.95 | -16.42 | -13.09 | -0.96 | 38.77 | 119.51 | 294.84 | 660.17 | 1408.1 | 2741.84 | 4700.77 | 7148.1 | 9868.23 | 12540.96 | 14879.7 | 16767.03 | 18246.16 | 19375.49 | 20283.22 | 21015.96 | 21584.49 | 22045.02 | 22451.15 | 22771.08 | 23032.42 | 23271.15 | 23477.88 | 23662.61 | 23827.95 | 23961.88 | 24079.21 | 24184.94 | 24257.27 | 24337.21 | 24421.54 | 24492.47 | 24539.2 | 24577.53
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 3e42d920-2009-ed11-85f4-5cf370a20867 | a1e5bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.5 | 2247.9 | 6.88 | 24232 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 |  |  | 6025595.86 | 3974483.03 |  |  | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 010081427802027710118705 2105883 17240328 | 10 | 4 | V1169 | 2022-07-21 13:09 | 2142 | 280 | 10.1648 | 4127.8836 | 3 | 12 | 35.35 | 26.39 | 20.62 | 11.06 | 2.09 | -4.07 | -11.64 | -23.6 | -24.57 | -19.53 | 1.1 | 48.54 | 154.17 | 365.21 | 807.44 | 1685.08 | 3199.11 | 5327.15 | 7920.58 | 10706.82 | 13358.65 | 15587.09 | 17329.12 | 18640.76 | 19665.4 | 20445.23 | 21073.07 | 21585.3 | 22013.94 | 22351.37 | 22657.21 | 22912.04 | 23135.28 | 23301.71 | 23469.95 | 23610.18 | 23733.22 | 23851.65 | 23956.09 | 24018.32 | 24091.56 | 24157.59 | 24181.43 | 24212.06 | 24231.5
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 4442d920-2009-ed11-85f4-5cf370a20867 | 813d7504-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:09 |  |  |  |  | 0
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | b7477ca2-3609-ed11-85f4-5cf370a20867 | f3c2a3da-3609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:50 |  |  |  |  | 0
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | f5c2a3da-3609-ed11-85f4-5cf370a20867 | 8eb7eb13-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:51 |  |  |  |  | 0
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 90b7eb13-3709-ed11-85f4-5cf370a20867 | b8bbce4d-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:53 |  |  |  |  | 0
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | f2bbce4d-3709-ed11-85f4-5cf370a20867 | 0c592f87-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:54 |  |  |  |  | 0
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 0e592f87-3709-ed11-85f4-5cf370a20867 | 4f3e36c2-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | Aborted | TargetAborted | 7162 (Warning, Quantity not sufficient), 1000 (Error, User Aborted Test) |  |  | 0 | [25 empty] | 2022-07-21 15:56 |  |  |  |  | 0
SecondDerivative
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 0d5c9e96-1209-ed11-85f4-5cf370a20867 | 8dc854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.3 | 2650.56 | 8.31 | 29213 | [20 empty] | 010081427802027710118705 2105880 17240328 | 8 | 3 | V790 | 2022-07-21 11:32 | 1436 | 266 | 7.9055 | 3994.9491 | 3 | 12 | 0 | -7.55 | -17.45 | -8.2 | -0.1 | -11 | 1.05 | 40.15 | 66.8 | 99.45 | 187 | 378.9 | 765.95 | 1408.35 | 2153.25 | 2624.15 | 2506 | 1761.65 | 585.95 | -675.8 | -1614.25 | -1975.9 | -1835.4 | -1461.8 | -1071.2 | -770.85 | -585.9 | -458.2 | -354.4 | -298.35 | -246.2 | -156 | -95.15 | -105.6 | -130.75 | -123.35 | -103.75 | -96.65 | -74.85 | -34.3 | -21.95 | -35.75 | -44.35 | -34.65 | -8.4
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 145c9e96-1209-ed11-85f4-5cf370a20867 | 2fc854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.24 | 2705.51 | 8.3 | 29527 | [20 empty] | 010081427802027710118705 2105880 17240328 | 9 | 3 | V790 | 2022-07-21 11:32 | 1469 | 268 | 12.2582 | 4050.6436 | 3 | 12 | 0 | -0.25 | -0.95 | -1.3 | 0.2 | 5.25 | 13.95 | 25.7 | 48.6 | 100.2 | 201.9 | 402.95 | 806.35 | 1475.05 | 2234.4 | 2687.6 | 2530.2 | 1740 | 508.05 | -799.75 | -1725.85 | -2030.05 | -1845.25 | -1444.3 | -1037.85 | -751.05 | -573.45 | -436.75 | -348.45 | -289.15 | -207.45 | -153.55 | -149.45 | -127.15 | -101.5 | -110.9 | -101.6 | -63.3 | -58.65 | -86.25 | -92.75 | -76.9 | -66.85 | -49.6 | -21.1
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 1b5c9e96-1209-ed11-85f4-5cf370a20867 | 78ca54cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.2 | 2626.87 | 8.51 | 28852 | [20 empty] | 010081427802027710118705 2105880 17240328 | 10 | 3 | V790 | 2022-07-21 11:32 | 1288 | 265 | 13.2618 | 3841.6764 | 3 | 12 | 0 | 0.7 | 3.75 | 7.7 | 7.85 | 4.95 | 8.65 | 26.5 | 55.8 | 102.2 | 198.65 | 404.1 | 806.2 | 1455.05 | 2186.5 | 2614.25 | 2434.25 | 1637.8 | 439.95 | -795.45 | -1655.9 | -1932.4 | -1745.7 | -1379.1 | -1034.5 | -759.3 | -556.75 | -433.4 | -331.6 | -224.6 | -184.1 | -200.45 | -182.7 | -133.25 | -106.65 | -101.25 | -81.5 | -50.15 | -61.7 | -101.8 | -94.05 | -64 | -65.35 | -64.65 | -44.15
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 225c9e96-1209-ed11-85f4-5cf370a20867 | 17ca54cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.22 | 2575.04 | 8.24 | 28421 | [20 empty] | 010081427802027710118705 2105880 17240328 | 11 | 3 | V790 | 2022-07-21 11:32 | 1419 | 259 | 12.4945 | 3931.0509 | 3 | 12 | 0 | 0.05 | 2.05 | 7.35 | 11.7 | 12 | 12.05 | 19.1 | 43.8 | 96.85 | 195.85 | 396.35 | 792.75 | 1427.45 | 2138.8 | 2560.65 | 2398.15 | 1635.65 | 472.7 | -744.1 | -1610.45 | -1909.05 | -1737.95 | -1367.1 | -1020.7 | -765.05 | -571.4 | -426 | -320.35 | -240.8 | -197.35 | -179.15 | -158.3 | -148.7 | -140.15 | -96.65 | -65.6 | -79.65 | -74.25 | -47.65 | -54.5 | -63.2 | -46.05 | -27.1 | -9.65
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 295c9e96-1209-ed11-85f4-5cf370a20867 | b5c954cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.25 | 2811.07 | 8.27 | 29297 | [20 empty] | 010081427802027710118705 2105880 17240328 | 12 | 3 | V790 | 2022-07-21 11:32 | 1453 | 244 | 11.0388 | 4033.9891 | 3 | 12 | 0 | 5.75 | 4.7 | -10.75 | -5.85 | 9.15 | 2.4 | 13.5 | 62.95 | 116.7 | 199.15 | 400.1 | 809.9 | 1481.55 | 2280.65 | 2790 | 2619.25 | 1767.05 | 500.1 | -841.95 | -1822.2 | -2138.85 | -1911.75 | -1481.15 | -1084.05 | -789.05 | -568.55 | -416.5 | -338.8 | -275.8 | -208.95 | -190.05 | -178.5 | -118.3 | -86.1 | -102.5 | -79.45 | -45.2 | -69.65 | -86.05 | -48.95 | -23.05 | -31.05 | -29.15 | -8.5
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 305c9e96-1209-ed11-85f4-5cf370a20867 | 1255c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.65 | 2107.83 | 6.31 | 23255 | 6.73 | 6.55 | 5370317.96 | 3542261.73 | 6.73 | 6.55 |  |  | 5370317.96 | 3542261.73 |  |  | 6.73 | 6.55 | 5370317.96 | 3542261.73 | 6.73 | 6.55 | 5370317.96 | 3542261.73 | 010081427802027710118705 2105879 17240328 | 1 | 4 | V1169 | 2022-07-21 11:32 | 2573 | 296 | 10.5939 | 4394.4255 | 3 | 12 | 0 | 3.55 | 10.2 | 9.8 | 4.3 | 4.7 | 8.05 | 13.1 | 29.15 | 58.55 | 113.05 | 236.5 | 493.35 | 939 | 1522.6 | 1999.65 | 2077.55 | 1649.6 | 814.65 | -191 | -1045.45 | -1497.55 | -1514.55 | -1268.25 | -975.05 | -732.55 | -547.2 | -418.8 | -324.85 | -252.1 | -201.25 | -143.5 | -92.55 | -84.75 | -88.6 | -85.15 | -85.95 | -66.75 | -31.65 | -21.7 | -39.85 | -65.8 | -81 | -68.3 | -41.8
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 375c9e96-1209-ed11-85f4-5cf370a20867 | b354c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.42 | 1950.48 | 6.9 | 23336 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 |  |  | 6309573.44 | 4161794.64 |  |  | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 010081427802027710118705 2105879 17240328 | 2 | 4 | V1169 | 2022-07-21 11:32 | 2058 | 284 | 10.0558 | 3971.7818 | 3 | 12 | 0 | 1.4 | 3.95 | 3.3 | 0 | 1.1 | 8.8 | 19.95 | 37.15 | 68.9 | 130.5 | 263.55 | 539.8 | 1000.5 | 1543.95 | 1914.45 | 1883.95 | 1409.25 | 609 | -290.8 | -1006.75 | -1338.5 | -1305.45 | -1095.8 | -862.5 | -655.35 | -507 | -416.25 | -335.8 | -258.4 | -192.6 | -134.1 | -109.9 | -114.2 | -101.15 | -79.05 | -73.3 | -70.15 | -57.1 | -49.45 | -62.8 | -74.95 | -65.25 | -39.95 | -10.15
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 3e5c9e96-1209-ed11-85f4-5cf370a20867 | 7255c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.3 | 2023.88 | 6.93 | 22748 | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 6.84 | 6.66 |  |  | 6918309.71 | 4563317.08 |  |  | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 010081427802027710118705 2105879 17240328 | 3 | 4 | V1169 | 2022-07-21 11:32 | 1979 | 281 | 10.2691 | 3858.8618 | 3 | 12 | 0 | 2.4 | 7.75 | 11.25 | 12 | 10.4 | 7.2 | 11.75 | 29.5 | 61.95 | 131.55 | 283 | 576.8 | 1058.4 | 1630.8 | 2002.6 | 1911.5 | 1349.3 | 483.4 | -463.85 | -1185.1 | -1444.8 | -1331.45 | -1094.1 | -846.9 | -624.75 | -472.2 | -361.25 | -251.45 | -185.05 | -171.9 | -158.05 | -138.15 | -128.15 | -109.05 | -73.75 | -41.6 | -35.45 | -52.6 | -58.1 | -49.95 | -54.25 | -61.9 | -43.5 | -7.85
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 6042e1f8-1209-ed11-85f4-5cf370a20867 | 233d8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.34 | 2623.16 | 6.33 | 27960 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 |  |  | 6760829.75 | 4459443.3 |  |  | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 010081427802027710118705 2105881 17240328 | 12 | 2 | V834 | 2022-07-21 11:34 | 2839 | 265 | 11.2448 | 5279.6436 | 3 | 12 | 0 | 0.05 | 1.95 | 8.4 | 15.7 | 17.45 | 17.1 | 24.3 | 44.65 | 83.7 | 166.5 | 349.1 | 714.85 | 1330.15 | 2080.75 | 2587.6 | 2494.05 | 1772.6 | 632.85 | -625.7 | -1598.5 | -1972.3 | -1815.35 | -1446.5 | -1088.1 | -789.8 | -560.75 | -421.55 | -343.6 | -268.5 | -208.7 | -187.35 | -155.85 | -113.6 | -106.55 | -103.35 | -73.35 | -51.65 | -52.85 | -51.55 | -42.65 | -47.45 | -65.8 | -64.95 | -42.25
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 6642e1f8-1209-ed11-85f4-5cf370a20867 | 7e8c0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.35 | 2535.4 | 8.04 | 28768 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 |  |  | 6606934.48 | 4357933.98 |  |  | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 010081427802027710118705 2105880 17240328 | 1 | 3 | V790 | 2022-07-21 11:34 | 1595 | 256 | 12.2133 | 4091.92 | 3 | 12 | 0 | 1.5 | 6.05 | 8.75 | 5.75 | 4.55 | 10.2 | 21 | 42.7 | 89.15 | 181.55 | 364.25 | 720.8 | 1316.1 | 2028.05 | 2501.4 | 2417.4 | 1736.1 | 633.6 | -584.45 | -1516.2 | -1880.9 | -1738 | -1376.5 | -1026.95 | -751.3 | -549 | -421.6 | -336 | -259.25 | -201.8 | -174.1 | -161.65 | -146.8 | -129.3 | -113.25 | -88.5 | -62.3 | -56.7 | -58.85 | -53.05 | -57.45 | -78.8 | -74.7 | -34.7
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 6c42e1f8-1209-ed11-85f4-5cf370a20867 | ef8b0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.32 | 2588.44 | 7.64 | 29845 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 |  |  | 6760829.75 | 4459443.3 |  |  | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 010081427802027710118705 2105880 17240328 | 2 | 3 | V790 | 2022-07-21 11:34 | 1882 | 264 | 9.5576 | 4509.8982 | 3 | 12 | 0 | -1.25 | 0.25 | 8.9 | 12.15 | 6.4 | 11.2 | 28.05 | 47.55 | 88.75 | 183 | 370.9 | 741.4 | 1363.7 | 2093.65 | 2559.8 | 2455.25 | 1754.05 | 643.3 | -565.3 | -1483.1 | -1848.8 | -1731.85 | -1408.85 | -1083.3 | -814.15 | -601.6 | -439.8 | -330.5 | -271.75 | -239.65 | -214.6 | -190.8 | -162.7 | -138.65 | -116.5 | -83.5 | -62 | -61.45 | -51.2 | -43.05 | -66.45 | -94.75 | -81.3 | -37.25
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 7242e1f8-1209-ed11-85f4-5cf370a20867 | bf8b0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.37 | 2773.69 | 8.52 | 30635 | [20 empty] | 010081427802027710118705 2105880 17240328 | 3 | 3 | V790 | 2022-07-21 11:34 | 1410 | 242 | 14.6582 | 4079.0836 | 3 | 12 | 0 | -0.4 | -0.9 | -0.15 | 2.25 | 6.9 | 14.35 | 25.05 | 45.05 | 88.35 | 180.2 | 371.1 | 752.6 | 1401.2 | 2192.7 | 2731.95 | 2648.65 | 1897.3 | 686.45 | -651.85 | -1681.45 | -2080.75 | -1917.4 | -1517.55 | -1121.5 | -798.4 | -559.55 | -406.25 | -315.1 | -254.55 | -217.4 | -204.3 | -184.3 | -140.45 | -110.5 | -103.6 | -90.75 | -83.35 | -85.9 | -71.05 | -54.25 | -59.1 | -74.2 | -65.2 | -25.55
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 7842e1f8-1209-ed11-85f4-5cf370a20867 | 608b0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.46 | 2532.36 | 8.41 | 30118 | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 6.79 | 6.61 |  |  | 6165950.02 | 4067060.63 |  |  | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 010081427802027710118705 2105880 17240328 | 4 | 3 | V790 | 2022-07-21 11:34 | 1354 | 259 | 12.6812 | 4069.7309 | 3 | 12 | 0 | 0.55 | 2.1 | 4.2 | 6.8 | 9.95 | 12.8 | 17.1 | 33.8 | 80 | 171.8 | 346.4 | 691.1 | 1274.8 | 1980.75 | 2478 | 2455.95 | 1844.65 | 805 | -365.2 | -1309.1 | -1763.8 | -1734.1 | -1455.75 | -1145.4 | -866.7 | -633 | -469.65 | -362.65 | -285.65 | -237.6 | -197.85 | -149.05 | -119.4 | -117.6 | -116.8 | -119 | -123.6 | -99.75 | -62 | -52.5 | -60.85 | -72.85 | -70.8 | -37.9
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 7e42e1f8-1209-ed11-85f4-5cf370a20867 | ebaadb1d-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified | 2610 (Informational, Delays in sample processing) | 16.12 | 1603.89 | 6.52 | 16922 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 |  |  | 7762471.17 | 5120125.98 |  |  | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 010081427802027710118705 2105880 17240328 | 5 | 3 | V790 | 2022-07-21 11:34 | 1372 | 268 | 2.2497 | 3104.4473 | 3 | 12 | 0 | -0.25 | -3.5 | -11.65 | -17 | -12.7 | -1.25 | 18.25 | 44.5 | 75.75 | 135.1 | 261.95 | 497.2 | 877.2 | 1333.15 | 1600.7 | 1438 | 912.65 | 228.7 | -506.15 | -1075.25 | -1211.8 | -1023.8 | -799.95 | -591.2 | -398.1 | -298.8 | -272.7 | -255.2 | -231.25 | -191.45 | -120.7 | -40.55 | -22.1 | -57.5 | -58.35 | -72.85 | -161.3 | -153.45 | 33.1 | 181.85 | 180.3 | 112.5 | 57.5 | 59.6
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 8442e1f8-1209-ed11-85f4-5cf370a20867 | 57c954cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.29 | 2625.53 | 8.37 | 30075 | [20 empty] | 010081427802027710118705 2105880 17240328 | 6 | 3 | V790 | 2022-07-21 11:34 | 1388 | 245 | 11.3673 | 4079.7455 | 3 | 12 | 0 | 4.9 | 3.55 | -12.95 | -13.5 | 1.3 | 3.1 | 16.85 | 60.9 | 111.55 | 192.85 | 383.9 | 767.2 | 1392.95 | 2129.75 | 2600.4 | 2475.7 | 1748.3 | 622.75 | -607.5 | -1544.55 | -1903.75 | -1769.25 | -1427.2 | -1072.35 | -785.25 | -582.9 | -443.5 | -363.05 | -314.5 | -243.95 | -180 | -158.1 | -133.7 | -113.6 | -127.15 | -119.9 | -79.4 | -60.4 | -56.1 | -40.1 | -38.95 | -70 | -80.35 | -44.8
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 8a42e1f8-1209-ed11-85f4-5cf370a20867 | edc854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.13 | 2178.33 | 8.18 | 26676 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 |  |  | 7762471.17 | 5120125.98 |  |  | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 010081427802027710118705 2105880 17240328 | 7 | 3 | V790 | 2022-07-21 11:34 | 1317 | 270 | 11.3709 | 3719.5782 | 3 | 12 | 0 | 0.95 | 1.45 | -1 | -2.05 | 2.05 | 9.75 | 23.7 | 50.3 | 96.75 | 186.3 | 377.2 | 744.3 | 1296.65 | 1872.2 | 2174.35 | 1995.9 | 1351.45 | 406.25 | -565.9 | -1239.6 | -1468.9 | -1380.25 | -1160.15 | -913.5 | -703.45 | -547.1 | -421.55 | -323.25 | -255.35 | -211.45 | -186.15 | -159.25 | -125.55 | -111.5 | -106.25 | -75.35 | -45.6 | -53.45 | -64.65 | -56.7 | -69.45 | -104.3 | -97.8 | -49.65
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 5a7b815a-1309-ed11-85f4-5cf370a20867 | 223c8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.28 | 2631.98 | 6.57 | 27149 | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 6.85 | 6.67 |  |  | 7079457.84 | 4669610.39 |  |  | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 010081427802027710118705 2105881 17240328 | 7 | 2 | V834 | 2022-07-21 11:37 | 2538 | 268 | 11.423 | 4907.5673 | 3 | 12 | 0 | 0.6 | 3.65 | 7.2 | 9.55 | 15.05 | 21 | 25.4 | 42.95 | 90.6 | 184.15 | 371.5 | 751.8 | 1389.05 | 2135.15 | 2608.25 | 2474.55 | 1700.05 | 493.55 | -796.25 | -1739.05 | -2061.55 | -1849.9 | -1413.7 | -1016.85 | -744.8 | -553.55 | -409.05 | -313.8 | -241.35 | -168 | -121.05 | -120.1 | -121.1 | -93.3 | -72.7 | -74.25 | -70.15 | -53.35 | -38.65 | -32.55 | -27.45 | -21.7 | -11.75 | 6.7
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 607b815a-1309-ed11-85f4-5cf370a20867 | 50597999-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.37 | 2314.8 | 6.2 | 25184 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 |  |  | 6606934.48 | 4357933.98 |  |  | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 010081427802027710118705 2105881 17240328 | 8 | 2 | V834 | 2022-07-21 11:37 | 2749 | 292 | 8.6594 | 4878.6545 | 3 | 12 | 0 | -0.45 | 0.95 | 5.8 | 9.05 | 10.15 | 14.35 | 23.05 | 39.5 | 75 | 150.5 | 306.85 | 620.5 | 1160.5 | 1826 | 2279.3 | 2210.9 | 1593.55 | 592.9 | -527.2 | -1385.35 | -1696.7 | -1544.55 | -1233.3 | -940.3 | -697.1 | -538.65 | -447.85 | -344.3 | -242.65 | -199.55 | -171.5 | -135.7 | -135.65 | -133.3 | -86.3 | -63 | -74.55 | -49.4 | -14.55 | -18.7 | -22.65 | -21.75 | -30.25 | -23.4
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 667b815a-1309-ed11-85f4-5cf370a20867 | 9c3e8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.45 | 2392.68 | 5.45 | 25142 | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 6.79 | 6.61 |  |  | 6165950.02 | 4067060.63 |  |  | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 010081427802027710118705 2105881 17240328 | 9 | 2 | V834 | 2022-07-21 11:37 | 3532 | 272 | 7.1382 | 5690.0436 | 3 | 12 | 0 | 1.95 | 6.1 | 6.8 | 4.35 | 5.65 | 11.55 | 21.45 | 39.8 | 77.15 | 150.9 | 299.2 | 608.5 | 1157 | 1844.5 | 2339.8 | 2314 | 1697.85 | 661.8 | -506.8 | -1435.2 | -1829.75 | -1720.8 | -1376.85 | -1014.05 | -724.8 | -532.6 | -393.8 | -278.8 | -218.85 | -195.8 | -156.9 | -126.2 | -122.65 | -109.75 | -84.95 | -62.95 | -49.05 | -44.85 | -23.7 | 7.6 | 1.3 | -33.55 | -40.95 | -15.65
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 6c7b815a-1309-ed11-85f4-5cf370a20867 | 3e3e8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.38 | 2436.69 | 5.53 | 25415 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 |  |  | 6606934.48 | 4357933.98 |  |  | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 010081427802027710118705 2105881 17240328 | 10 | 2 | V834 | 2022-07-21 11:37 | 3437 | 275 | 7.8133 | 5650.56 | 3 | 12 | 0 | -2.15 | -4.85 | -2.05 | 4.3 | 10.3 | 16.2 | 25.45 | 45.5 | 80.8 | 151 | 312.55 | 645.45 | 1209.15 | 1906.9 | 2396.4 | 2330.15 | 1660.25 | 584.5 | -601.65 | -1517.45 | -1874.75 | -1725 | -1354.85 | -985.25 | -691.9 | -491.35 | -366.3 | -293.05 | -258.3 | -221.35 | -161.1 | -119.3 | -104.85 | -81.6 | -64.45 | -73.65 | -71 | -33.9 | -6.5 | -25.6 | -54.5 | -56.55 | -39.2 | -17.65
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 727b815a-1309-ed11-85f4-5cf370a20867 | e03d8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.44 | 2298.39 | 5.71 | 25080 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 |  |  | 6309573.44 | 4161794.64 |  |  | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 010081427802027710118705 2105881 17240328 | 11 | 2 | V834 | 2022-07-21 11:37 | 3289 | 280 | 6.577 | 5366.3527 | 3 | 12 | 0 | -0.9 | 1.2 | 8.25 | 10.8 | 8.9 | 13.5 | 27.65 | 47.95 | 77.75 | 143.05 | 292.6 | 597.4 | 1120.15 | 1774 | 2248.75 | 2220.8 | 1617.6 | 624.6 | -460.85 | -1307.1 | -1664.75 | -1562.85 | -1259.2 | -963.5 | -734.65 | -554.5 | -406.15 | -312.9 | -257.75 | -195.25 | -170.45 | -188.75 | -151.55 | -80.45 | -70.45 | -71.45 | -31.95 | -25.15 | -61.6 | -64.55 | -38.3 | -37.8 | -43.25 | -26.3
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 787b815a-1309-ed11-85f4-5cf370a20867 | 1eff3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.43 | 2367.47 | 6.98 | 25496 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 |  |  | 6309573.44 | 4161794.64 |  |  | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 010081427802027710118705 2105881 17240328 | 4 | 2 | V834 | 2022-07-21 11:37 | 2183 | 247 | 10.5467 | 4277.56 | 3 | 12 | 0 | 0.7 | 2 | 3.3 | 6.95 | 12.1 | 13.85 | 17.25 | 36.15 | 78.45 | 153.05 | 303.1 | 618 | 1162.45 | 1835.3 | 2318.65 | 2284.75 | 1666.4 | 638.95 | -515.15 | -1411.45 | -1751.5 | -1617.6 | -1304.75 | -978.4 | -702.4 | -526.55 | -421.95 | -333.65 | -270.1 | -238.7 | -202.15 | -144.35 | -94.8 | -69.5 | -57.45 | -55 | -54.3 | -46.7 | -47.3 | -61.2 | -68.45 | -64.65 | -49.2 | -25.35
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 7e7b815a-1309-ed11-85f4-5cf370a20867 | 823d8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.4 | 2435.14 | 6.91 | 25271 | 6.81 | 6.63 | 6456542.29 | 4258735.29 | 6.81 | 6.63 |  |  | 6456542.29 | 4258735.29 |  |  | 6.81 | 6.63 | 6456542.29 | 4258735.29 | 6.81 | 6.63 | 6456542.29 | 4258735.29 | 010081427802027710118705 2105881 17240328 | 5 | 2 | V834 | 2022-07-21 11:37 | 2203 | 258 | 11.3697 | 4297.6873 | 3 | 12 | 0 | 0.7 | 4.1 | 6.85 | 3.85 | 2.95 | 11 | 22.25 | 38.7 | 75.3 | 151.85 | 311.05 | 637.2 | 1201.35 | 1901.35 | 2391.9 | 2336.1 | 1682.45 | 609.15 | -584.4 | -1499.05 | -1856.35 | -1717.4 | -1358.65 | -1009.9 | -739.85 | -536.9 | -398.45 | -301.1 | -218.65 | -168.05 | -147.45 | -131.5 | -123.15 | -111.55 | -78.05 | -47.25 | -42.25 | -47.95 | -44.3 | -37 | -40.2 | -46.05 | -36.6 | -17.65
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 847b815a-1309-ed11-85f4-5cf370a20867 | c43c8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.18 | 2359.93 | 6.51 | 24908 | 6.88 | 6.7 | 7585775.75 | 5003577.68 | 6.88 | 6.7 |  |  | 7585775.75 | 5003577.68 |  |  | 6.88 | 6.7 | 7585775.75 | 5003577.68 | 6.88 | 6.7 | 7585775.75 | 5003577.68 | 010081427802027710118705 2105881 17240328 | 6 | 2 | V834 | 2022-07-21 11:37 | 2433 | 253 | 11.0485 | 4546.5564 | 3 | 12 | 0 | 0.35 | 1.6 | 2.65 | 2.6 | 3.75 | 9.3 | 23.35 | 47.7 | 86.5 | 169.9 | 354.15 | 713.35 | 1297.3 | 1965.4 | 2350.55 | 2171.25 | 1449.65 | 377.65 | -751.25 | -1542.3 | -1773.75 | -1582.85 | -1240.2 | -912.1 | -659.15 | -491.05 | -381.85 | -303.7 | -246.8 | -198.75 | -151.25 | -120.45 | -107.75 | -84.4 | -60.45 | -64 | -65.45 | -33.55 | -2.2 | -7.35 | -26.3 | -39.1 | -41.65 | -26.55
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | b3de19db-1309-ed11-85f4-5cf370a20867 | 685f6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.51 | 2123.17 | 6.78 | 22703 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 |  |  | 6025595.86 | 3974483.03 |  |  | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 010081427802027710118705 2105881 17240328 | 3 | 2 | V834 | 2022-07-21 11:41 | 2093 | 263 | 7.3515 | 3941.8836 | 3 | 12 | 0 | 0.05 | 0.55 | 2.85 | 5 | 3.9 | 5.8 | 16 | 32.7 | 63.8 | 129.6 | 264.9 | 534.95 | 1001.95 | 1599.45 | 2054.6 | 2062.05 | 1551.2 | 672.45 | -347.5 | -1188.3 | -1569.1 | -1516.2 | -1255.75 | -955.45 | -699.5 | -510.65 | -374.45 | -282.1 | -215.6 | -170.9 | -144.85 | -106.55 | -77.35 | -82 | -76.2 | -62.3 | -68.3 | -57.85 | -29.55 | -27.35 | -38.9 | -39.25 | -31.15 | -17.9
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | b9de19db-1309-ed11-85f4-5cf370a20867 | 785e6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.61 | 2768.06 | 7.99 | 29669 | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 6.75 | 6.57 |  |  | 5623413.25 | 3709203.38 |  |  | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 010081427802027710118705 2105882 17240328 | 10 | 1 | V639 | 2022-07-21 11:41 | 1720 | 268 | 12.9721 | 4239.3891 | 3 | 12 | 0 | 0.75 | 3.25 | 5.5 | 2.7 | -1.9 | 3.6 | 19.95 | 41.15 | 78.95 | 158.8 | 319.8 | 646.15 | 1233.6 | 2014.2 | 2641.6 | 2715.15 | 2102.75 | 961.8 | -383.2 | -1497.2 | -2024.95 | -1967.4 | -1612.95 | -1217.75 | -892.95 | -664.25 | -508.3 | -387.95 | -290 | -222.05 | -179.5 | -140.8 | -114.15 | -121.5 | -131.2 | -110.2 | -75.9 | -47.75 | -33.05 | -38.8 | -60.75 | -81.1 | -73.7 | -44.25
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | bfde19db-1309-ed11-85f4-5cf370a20867 | 485e6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.29 | 2777.15 | 8.24 | 29533 | [20 empty] | 010081427802027710118705 2105882 17240328 | 11 | 1 | V639 | 2022-07-21 11:41 | 1560 | 266 | 14.7564 | 4080.1273 | 3 | 12 | 0 | 0.85 | 0.45 | -3.15 | -2.1 | 7.15 | 15.45 | 24.05 | 48.35 | 97.15 | 188 | 385.05 | 790.9 | 1461.3 | 2246.15 | 2750.4 | 2615.9 | 1806.85 | 552.35 | -787 | -1726.2 | -1971.5 | -1793.2 | -1534 | -1193.2 | -878.05 | -731.15 | -559.2 | -299.15 | -169.5 | -166.6 | -103.4 | 19.35 | -11.05 | -195.4 | -247.2 | -196.45 | -226.05 | -157.2 | 49.35 | 104.1 | -2.6 | -70.35 | -67.9 | -44.35
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | c5de19db-1309-ed11-85f4-5cf370a20867 | 90fe3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.47 | 2893.71 | 8.04 | 30993 | [20 empty] | 010081427802027710118705 2105882 17240328 | 12 | 1 | V639 | 2022-07-21 11:41 | 1817 | 263 | 16.143 | 4407.3273 | 3 | 12 | 0 | -0.7 | -2.05 | -0.1 | 4.75 | 6.95 | 10.05 | 21.75 | 46.35 | 90.6 | 175.85 | 359.85 | 742.4 | 1401.1 | 2223.45 | 2825.05 | 2806.9 | 2081.1 | 847.95 | -550.35 | -1669.25 | -2164.55 | -2058 | -1655.3 | -1243.75 | -928.5 | -685.8 | -502.65 | -390.6 | -313.6 | -241.3 | -187.8 | -146.75 | -117.1 | -117.4 | -114.3 | -88.45 | -81.75 | -75.7 | -28.6 | 3.6 | -20.45 | -56.65 | -57.45 | -26.8
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | cbde19db-1309-ed11-85f4-5cf370a20867 | 60fe3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.49 | 2361.31 | 5.95 | 25803 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 |  |  | 6025595.86 | 3974483.03 |  |  | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 010081427802027710118705 2105881 17240328 | 1 | 2 | V834 | 2022-07-21 11:41 | 3061 | 258 | 8.7067 | 5253.04 | 3 | 12 | 0 | -0.7 | -4.55 | -8.95 | -4.5 | 7.7 | 17.3 | 25.95 | 43.15 | 73.3 | 135.6 | 287.25 | 607 | 1143.4 | 1806.9 | 2300.6 | 2297.9 | 1717.45 | 726.9 | -402.45 | -1319.3 | -1729.15 | -1650.35 | -1339.8 | -1009.15 | -753.05 | -586.45 | -465.6 | -365.25 | -288.6 | -218.4 | -154.65 | -121.75 | -106.2 | -84.9 | -73.6 | -70.6 | -57.3 | -59.6 | -76.6 | -59 | -28.5 | -36.35 | -46.9 | -29.2
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | d1de19db-1309-ed11-85f4-5cf370a20867 | 985f6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.29 | 2373.91 | 6.32 | 24689 | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 6.84 | 6.66 |  |  | 6918309.71 | 4563317.08 |  |  | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 010081427802027710118705 2105881 17240328 | 2 | 2 | V834 | 2022-07-21 11:41 | 2647 | 271 | 10.5273 | 4659.0655 | 3 | 12 | 0 | 0.15 | 2.2 | 4.05 | 1.85 | 2.4 | 10.55 | 21.5 | 39.7 | 82.55 | 169.25 | 335.8 | 666.7 | 1232.3 | 1910.75 | 2350 | 2236.05 | 1545.15 | 473.9 | -680.25 | -1531.55 | -1818.2 | -1640.25 | -1287.65 | -949.15 | -687.75 | -507.2 | -383.8 | -288.4 | -206.65 | -164.25 | -152.25 | -115.85 | -83.15 | -88.05 | -72.85 | -48.8 | -61.45 | -49.95 | -17.25 | -37.05 | -61.35 | -42.55 | -22.3 | -13.3
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | d7de19db-1309-ed11-85f4-5cf370a20867 | 954775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.27 | 2874.99 | 8.13 | 29909 | [20 empty] | 010081427802027710118705 2105882 17240328 | 8 | 1 | V639 | 2022-07-21 11:41 | 1689 | 271 | 15.3721 | 4204.9491 | 3 | 12 | 0 | 0.05 | -0.25 | -0.75 | 0.2 | 4.25 | 13.3 | 30.4 | 58.3 | 104 | 197.3 | 403.95 | 821.35 | 1513.05 | 2330.3 | 2850.95 | 2691.85 | 1829.25 | 514.55 | -877.15 | -1881.15 | -2201.3 | -1963.55 | -1516.75 | -1104.5 | -807.6 | -607.7 | -452.65 | -330.2 | -244.85 | -180.45 | -141.7 | -131.35 | -115.55 | -95.15 | -97.5 | -98.05 | -83.8 | -91.4 | -103.85 | -74.35 | -37.7 | -39.45 | -46.35 | -33.2
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | ddde19db-1309-ed11-85f4-5cf370a20867 | db5e6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.06 | 2821.6 | 7.95 | 29515 | [20 empty] | 010081427802027710118705 2105882 17240328 | 9 | 1 | V639 | 2022-07-21 11:41 | 1825 | 274 | 18.3176 | 4254.1782 | 3 | 12 | 0 | -0.75 | -1.3 | 2.25 | 8.05 | 11.7 | 18.05 | 31.95 | 55.95 | 108.05 | 224.65 | 465.8 | 930.25 | 1659.3 | 2435.05 | 2820.25 | 2521.15 | 1582.45 | 257.85 | -1066.7 | -1943.65 | -2160.05 | -1886.3 | -1441.4 | -1042.05 | -750.65 | -538.3 | -388.25 | -304.25 | -250.9 | -207.05 | -184.6 | -164.75 | -135.2 | -108.9 | -73.05 | -35.7 | -40.45 | -55.2 | -31.1 | -24.85 | -61.15 | -73.15 | -37.95 | 4.1
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 65fdae7a-1409-ed11-85f4-5cf370a20867 | 364775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.27 | 2347.62 | 7.47 | 26286 | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 6.85 | 6.67 |  |  | 7079457.84 | 4669610.39 |  |  | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 010081427802027710118705 2105882 17240328 | 4 | 1 | V639 | 2022-07-21 11:45 | 1808 | 260 | 11.5855 | 4075.4291 | 3 | 12 | 0 | -0.1 | -0.7 | -1.95 | -1.55 | 2.85 | 10.75 | 23.75 | 45.7 | 86.4 | 171.55 | 346.9 | 688.9 | 1253.75 | 1915.2 | 2328.05 | 2204.9 | 1547.1 | 553.25 | -517.4 | -1369.45 | -1742 | -1625.05 | -1310.55 | -1031.35 | -761.3 | -502.55 | -365.75 | -317 | -252.55 | -188.65 | -161.4 | -148.7 | -129.05 | -109.8 | -99.8 | -85.8 | -58.85 | -44.4 | -50.25 | -47.4 | -39.6 | -51.05 | -55.45 | -33.2
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 6bfdae7a-1409-ed11-85f4-5cf370a20867 | a94675b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.02 | 2809.5 | 7.97 | 30622 | 6.92 | 6.74 | 8317637.71 | 5486313.83 | 6.92 | 6.74 |  |  | 8317637.71 | 5486313.83 |  |  | 6.92 | 6.74 | 8317637.71 | 5486313.83 | 6.92 | 6.74 | 8317637.71 | 5486313.83 | 010081427802027710118705 2105882 17240328 | 5 | 1 | V639 | 2022-07-21 11:45 | 1826 | 257 | 19.0012 | 4401.4909 | 3 | 12 | 0 | 0.85 | 2.45 | 4 | 6.2 | 8.5 | 12.75 | 26.6 | 58.6 | 119.7 | 239.75 | 489.15 | 968.6 | 1700.3 | 2456.25 | 2809.35 | 2484.3 | 1539.7 | 231.9 | -1054.9 | -1887.05 | -2076.75 | -1803.85 | -1377 | -1005.15 | -741 | -541.1 | -401.05 | -336.05 | -287.3 | -239.55 | -231.35 | -214.45 | -169.55 | -145.75 | -109.4 | -56.05 | -59.55 | -83.7 | -61.45 | -38.15 | -45.45 | -53.7 | -42.7 | -18.95
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 71fdae7a-1409-ed11-85f4-5cf370a20867 | 774675b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.14 | 2540.68 | 7.78 | 28700 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 |  |  | 7762471.17 | 5120125.98 |  |  | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 010081427802027710118705 2105882 17240328 | 6 | 1 | V639 | 2022-07-21 11:45 | 1746 | 262 | 16.9309 | 4242.9382 | 3 | 12 | 0 | 1.4 | 4.85 | 6.45 | 6.55 | 11.7 | 19.65 | 26.35 | 45.85 | 98.7 | 201.8 | 407.95 | 819.8 | 1470.3 | 2164.3 | 2535.25 | 2323.4 | 1539.55 | 397.7 | -767.45 | -1582.8 | -1839.65 | -1648.5 | -1302.05 | -991.95 | -742.05 | -553.65 | -439.2 | -363.65 | -293.8 | -234.95 | -185.6 | -144.85 | -121.95 | -97.8 | -73.85 | -86.4 | -107.55 | -85.6 | -54.35 | -57.05 | -62 | -53.65 | -41.2 | -18.6
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 77fdae7a-1409-ed11-85f4-5cf370a20867 | c64775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.03 | 2805.63 | 8.16 | 29973 | [20 empty] | 010081427802027710118705 2105882 17240328 | 7 | 1 | V639 | 2022-07-21 11:45 | 1646 | 261 | 19.9042 | 4183.5382 | 3 | 12 | 0 | 0.65 | 3.35 | 6.7 | 8.15 | 10.55 | 16.5 | 27 | 54 | 116.45 | 239.45 | 484.3 | 953.85 | 1680.75 | 2441.9 | 2805.25 | 2487.45 | 1543.4 | 223.85 | -1085.8 | -1927.95 | -2108.85 | -1830.45 | -1392.55 | -988.15 | -709.7 | -536.1 | -411.55 | -342.3 | -313.85 | -266.6 | -201.8 | -153.95 | -112.45 | -73.35 | -60.2 | -72.2 | -78.45 | -58.25 | -39.55 | -51.5 | -64.9 | -61 | -49.15 | -29.55
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | f47fc1ae-1609-ed11-85f4-5cf370a20867 | 816b47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.14 | 2841.9 | 7.99 | 29636 | [20 empty] | 010081427802027710118705 2105882 17240328 | 1 | 1 | V639 | 2022-07-21 12:01 | 1679 | 254 | 18.6121 | 4231.3491 | 3 | 12 | 0 | 0.45 | 2.2 | 4.9 | 5.85 | 6.4 | 12.5 | 25.35 | 49.8 | 102.85 | 213.05 | 439.4 | 881.8 | 1593.85 | 2390.7 | 2834.95 | 2586.85 | 1673.15 | 356.85 | -989.4 | -1914.2 | -2162.05 | -1893.8 | -1452.25 | -1052.25 | -762.45 | -570.1 | -420.5 | -305.6 | -251.45 | -227.15 | -188.55 | -151.9 | -131.85 | -103.35 | -71.4 | -61.55 | -59.35 | -46.45 | -42.55 | -60.5 | -77.2 | -76.95 | -57.7 | -29.8
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | fa7fc1ae-1609-ed11-85f4-5cf370a20867 | 226b47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.12 | 3139.19 | 8.24 | 32203 | [20 empty] | 010081427802027710118705 2105882 17240328 | 2 | 1 | V639 | 2022-07-21 12:01 | 1768 | 263 | 19.5394 | 4452.2145 | 3 | 12 | 0 | -1.65 | -4.15 | -0.85 | 6.55 | 10.35 | 16.8 | 34.15 | 62.75 | 118.6 | 242.8 | 500 | 989.2 | 1769.8 | 2647.8 | 3133.4 | 2838 | 1804.75 | 334.55 | -1175.35 | -2212.1 | -2458.05 | -2102.35 | -1566.7 | -1108.2 | -787.45 | -586.85 | -470.75 | -388.35 | -296.1 | -207.75 | -149.25 | -107.8 | -83.75 | -87.6 | -92.9 | -84.6 | -86.05 | -89.95 | -66.25 | -43.1 | -53 | -64.75 | -46.05 | -14.2
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | 0080c1ae-1609-ed11-85f4-5cf370a20867 | e16b47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.13 | 3191.32 | 8.32 | 32904 | [20 empty] | 010081427802027710118705 2105882 17240328 | 3 | 1 | V639 | 2022-07-21 12:01 | 1736 | 265 | 18.4218 | 4505.0364 | 3 | 12 | 0 | 2.4 | 7.4 | 7.25 | 3.65 | 8.45 | 18.65 | 32.5 | 64.45 | 123.8 | 237.3 | 485.3 | 986.15 | 1795.8 | 2695.2 | 3185.05 | 2893.15 | 1847.45 | 340.3 | -1180.6 | -2199.65 | -2454.35 | -2125.15 | -1598.05 | -1136.8 | -804.45 | -576 | -430.9 | -347.7 | -296.9 | -269.15 | -245.75 | -193.6 | -138.6 | -110.65 | -72.65 | -35.9 | -49.9 | -68.1 | -47.95 | -35.95 | -46.95 | -56.45 | -48.8 | -22.7
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | a79dad68-1f09-ed11-85f4-5cf370a20867 | b26dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | ad9dad68-1f09-ed11-85f4-5cf370a20867 | b66dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | b39dad68-1f09-ed11-85f4-5cf370a20867 | ba6dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | b99dad68-1f09-ed11-85f4-5cf370a20867 | be6dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | bf9dad68-1f09-ed11-85f4-5cf370a20867 | c26dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | c59dad68-1f09-ed11-85f4-5cf370a20867 | c76dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | cb9dad68-1f09-ed11-85f4-5cf370a20867 | 2f0bca23-2809-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 17.03 | 2397.7 | 7.16 | 25393 | 6.62 | 6.44 | 4168693.83 | 2749670.45 | 6.62 | 6.44 |  |  | 4168693.83 | 2749670.45 |  |  | 6.62 | 6.44 | 4168693.83 | 2749670.45 | 6.62 | 6.44 | 4168693.83 | 2749670.45 | 010081427802027710118705 2105883 17240328 | 11 | 4 | V1169 | 2022-07-21 13:03 | 2048 | 285 | 11.6691 | 4140.7564 | 3 | 13 | 0 | -0.4 | -4.25 | -7.5 | -0.15 | 8 | 4.7 | 5.15 | 23.3 | 51.8 | 96.95 | 198.1 | 410.3 | 805.6 | 1420.6 | 2080.85 | 2397.5 | 2111.75 | 1292.5 | 169.3 | -938.85 | -1648.4 | -1793.9 | -1543 | -1164.45 | -851.5 | -648.8 | -484 | -338.7 | -260.1 | -225.3 | -181.2 | -140.55 | -120.05 | -98.35 | -74.5 | -60.5 | -51.25 | -56.05 | -62 | -35 | -16.25 | -54.15 | -82.7 | -57.15
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | d29dad68-1f09-ed11-85f4-5cf370a20867 | d10aca23-2809-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.59 | 2333.73 | 6.98 | 25183 | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 6.75 | 6.57 |  |  | 5623413.25 | 3709203.38 |  |  | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 010081427802027710118705 2105883 17240328 | 12 | 4 | V1169 | 2022-07-21 13:03 | 2088 | 278 | 10.3709 | 4218.2182 | 3 | 12 | 0 | 2.75 | 6.8 | 4.2 | -1.05 | 0.85 | 5.45 | 9.45 | 26.3 | 65.6 | 132.2 | 263.95 | 540.85 | 1035.85 | 1694.75 | 2230.55 | 2284.8 | 1750.6 | 791.5 | -334.15 | -1280 | -1721.25 | -1649 | -1326.85 | -981.15 | -712.65 | -532.7 | -394.5 | -286.4 | -233.6 | -206.65 | -160.45 | -121.3 | -109.45 | -91.35 | -72.55 | -74.95 | -75.3 | -73.6 | -79.6 | -68.3 | -55.4 | -67.6 | -65.1 | -35.35
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 1f42d920-2009-ed11-85f4-5cf370a20867 | 41e5bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.64 | 2405.59 | 7.37 | 25353 | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 6.74 | 6.56 |  |  | 5495408.74 | 3624771.6 |  |  | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 010081427802027710118705 2105883 17240328 | 5 | 4 | V1169 | 2022-07-21 13:09 | 1954 | 288 | 13.0194 | 3986.2545 | 3 | 12 | 0 | 1.2 | 4.7 | 7.05 | 6.6 | 6.8 | 6.95 | 11.15 | 32.3 | 73.6 | 142.3 | 282.55 | 570.25 | 1069.9 | 1729.5 | 2279.35 | 2366.45 | 1841.6 | 833.9 | -353.45 | -1338.95 | -1826.8 | -1778.55 | -1414.95 | -1028.35 | -746.7 | -547.35 | -417.3 | -330.4 | -245.9 | -180.65 | -139.95 | -112.1 | -115.45 | -121.8 | -101.75 | -88.1 | -74.6 | -33.05 | -0.35 | -11.35 | -35.8 | -47.35 | -37.45 | -8.1
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 2642d920-2009-ed11-85f4-5cf370a20867 | b3e4bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.08 | 2303.83 | 7.28 | 24419 | 6.91 | 6.73 | 8128305.16 | 5361430.08 | 6.91 | 6.73 |  |  | 8128305.16 | 5361430.08 |  |  | 6.91 | 6.73 | 8128305.16 | 5361430.08 | 6.91 | 6.73 | 8128305.16 | 5361430.08 | 010081427802027710118705 2105883 17240328 | 6 | 4 | V1169 | 2022-07-21 13:09 | 1944 | 278 | 15.3091 | 3909.1018 | 3 | 12 | 0 | 1.5 | 4.15 | 4.4 | 6.05 | 12.05 | 15.65 | 23.4 | 50.05 | 98.55 | 190.65 | 387.7 | 767.2 | 1354.85 | 1982.9 | 2302 | 2072.45 | 1318.2 | 227.25 | -873.15 | -1592.55 | -1763.8 | -1541.5 | -1176.15 | -843.15 | -613.8 | -462.15 | -343 | -250.45 | -189.85 | -156.35 | -139.6 | -130.2 | -120.45 | -96.7 | -67.75 | -54.05 | -44.45 | -44.75 | -62.85 | -58 | -35.15 | -42.2 | -52.6 | -36.55
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 2c42d920-2009-ed11-85f4-5cf370a20867 | bde6bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.63 | 2270.38 | 7.03 | 24694 | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 6.74 | 6.56 |  |  | 5495408.74 | 3624771.6 |  |  | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 010081427802027710118705 2105883 17240328 | 7 | 4 | V1169 | 2022-07-21 13:09 | 2123 | 298 | 10.1176 | 4103.1782 | 3 | 12 | 0 | 1.95 | 6.45 | 7.1 | 1.55 | -0.7 | 7.4 | 21.8 | 39.3 | 70.1 | 134 | 263.7 | 529.25 | 1008.95 | 1645.65 | 2160.15 | 2231.25 | 1740.9 | 803.8 | -305.15 | -1224.7 | -1672.7 | -1622.15 | -1299.6 | -970.7 | -728.7 | -541.5 | -405.35 | -328.35 | -264.1 | -181.3 | -129.5 | -139.05 | -141.55 | -100.4 | -81.3 | -91.65 | -61.65 | -17.8 | -29.55 | -57.5 | -56.2 | -59.95 | -65.05 | -46.55
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 3242d920-2009-ed11-85f4-5cf370a20867 | 5ee6bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.32 | 2201.98 | 6.92 | 24142 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 |  |  | 6760829.75 | 4459443.3 |  |  | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 010081427802027710118705 2105883 17240328 | 8 | 4 | V1169 | 2022-07-21 13:09 | 2146 | 296 | 13.8606 | 4100.0255 | 3 | 12 | 0 | 1.4 | 4.8 | 7.7 | 8.6 | 8.65 | 9.55 | 16.3 | 38.15 | 81.35 | 160.45 | 319.3 | 629.9 | 1146.35 | 1767.4 | 2176.55 | 2086.2 | 1471.85 | 495.05 | -569.6 | -1357.45 | -1649.5 | -1521 | -1191.9 | -871.2 | -651.75 | -489.5 | -353.3 | -271.5 | -226.2 | -175.45 | -130.25 | -108.15 | -97.15 | -94.1 | -104.95 | -105.35 | -76.35 | -55 | -51.65 | -29.65 | -15.9 | -51.25 | -76.9 | -55.15
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 3842d920-2009-ed11-85f4-5cf370a20867 | d2e5bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.73 | 2195.79 | 7.21 | 24578 | 6.71 | 6.53 | 5128613.84 | 3382833.69 | 6.71 | 6.53 |  |  | 5128613.84 | 3382833.69 |  |  | 6.71 | 6.53 | 5128613.84 | 3382833.69 | 6.71 | 6.53 | 5128613.84 | 3382833.69 | 010081427802027710118705 2105883 17240328 | 9 | 4 | V1169 | 2022-07-21 13:09 | 1973 | 289 | 7.6679 | 3971.2109 | 3 | 12 | 0 | 0.6 | 3.85 | 7.15 | 6.3 | 7.2 | 11.65 | 14.15 | 23.65 | 54.85 | 115.8 | 233.4 | 475.45 | 918.9 | 1529.65 | 2056.1 | 2175.85 | 1765.2 | 923 | -123.05 | -1033.65 | -1518.6 | -1552.4 | -1311.1 | -1000.3 | -762.75 | -598.3 | -445.9 | -331.5 | -274.95 | -218.3 | -151.55 | -112.1 | -97.9 | -94.65 | -91.5 | -81.1 | -73.35 | -57.15 | -23.85 | -13.6 | -40.1 | -68.5 | -63.15 | -28.05
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 3e42d920-2009-ed11-85f4-5cf370a20867 | a1e5bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.5 | 2247.9 | 6.88 | 24232 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 |  |  | 6025595.86 | 3974483.03 |  |  | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 010081427802027710118705 2105883 17240328 | 10 | 4 | V1169 | 2022-07-21 13:09 | 2142 | 280 | 10.1648 | 4127.8836 | 3 | 12 | 0 | -0.95 | -1.85 | 0.35 | 1.35 | -0.8 | 5.2 | 21.65 | 41.2 | 73.5 | 141.85 | 281.2 | 563.05 | 1061.25 | 1700.55 | 2178.15 | 2179.85 | 1630.4 | 667.85 | -437.45 | -1315.3 | -1688.75 | -1586.05 | -1265.3 | -942.7 | -681.05 | -493 | -378.1 | -293.65 | -218.4 | -176.7 | -162.65 | -134.75 | -94.3 | -72.75 | -61.45 | -55.65 | -69.35 | -69.25 | -46.95 | -50.85 | -63.75 | -52.4 | -35.35 | -19.7
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 4442d920-2009-ed11-85f4-5cf370a20867 | 813d7504-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:09 |  |  |  |  | 0
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | b7477ca2-3609-ed11-85f4-5cf370a20867 | f3c2a3da-3609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:50 |  |  |  |  | 0
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | f5c2a3da-3609-ed11-85f4-5cf370a20867 | 8eb7eb13-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:51 |  |  |  |  | 0
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 90b7eb13-3709-ed11-85f4-5cf370a20867 | b8bbce4d-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:53 |  |  |  |  | 0
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | f2bbce4d-3709-ed11-85f4-5cf370a20867 | 0c592f87-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:54 |  |  |  |  | 0
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 0e592f87-3709-ed11-85f4-5cf370a20867 | 4f3e36c2-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | Aborted | TargetAborted | 7162 (Warning, Quantity not sufficient), 1000 (Error, User Aborted Test) |  |  | 0 | [25 empty] | 2022-07-21 15:56 |  |  |  |  | 0
Modulated
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 0d5c9e96-1209-ed11-85f4-5cf370a20867 | 8dc854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.3 | 2650.56 | 8.31 | 29213 | [20 empty] | 010081427802027710118705 2105880 17240328 | 8 | 3 | V790 | 2022-07-21 11:32 | 1436 | 266 | 7.9055 | 3994.9491 | 3 | 12 | 4043 | 4024 | 4015 | 4022 | 4148 | 3986 | 4014 | 4026 | 4025 | 4029 | 4052 | 4104 | 4209 | 4410 | 4756 | 5543 | 7164 | 10158 | 14158 | 17999 | 21370 | 23720 | 25748 | 27126 | 28235 | 29116 | 29754 | 30412 | 30998 | 31409 | 31609 | 31979 | 32215 | 32431 | 32640 | 32900 | 33024 | 33159 | 33105 | 33309 | 33392 | 33360 | 33538 | 33464 | 33605
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 145c9e96-1209-ed11-85f4-5cf370a20867 | 2fc854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.24 | 2705.51 | 8.3 | 29527 | [20 empty] | 010081427802027710118705 2105880 17240328 | 9 | 3 | V790 | 2022-07-21 11:32 | 1469 | 268 | 12.2582 | 4050.6436 | 3 | 12 | 4109 | 4124 | 4116 | 4115 | 4112 | 4114 | 4123 | 4116 | 4116 | 4129 | 4168 | 4207 | 4307 | 4507 | 4925 | 5724 | 7445 | 10564 | 14652 | 18604 | 21933 | 24363 | 26242 | 27543 | 28722 | 29599 | 30262 | 30895 | 31477 | 31771 | 32232 | 32522 | 32667 | 33019 | 33286 | 33411 | 33544 | 33682 | 33832 | 33883 | 34051 | 34091 | 34143 | 34066 | 34145
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 1b5c9e96-1209-ed11-85f4-5cf370a20867 | 78ca54cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.2 | 2626.87 | 8.51 | 28852 | [20 empty] | 010081427802027710118705 2105880 17240328 | 10 | 3 | V790 | 2022-07-21 11:32 | 1288 | 265 | 13.2618 | 3841.6764 | 3 | 12 | 3928 | 3907 | 3908 | 3907 | 3920 | 3901 | 3919 | 3908 | 3938 | 3919 | 3950 | 4021 | 4125 | 4325 | 4718 | 5541 | 7259 | 10329 | 14304 | 18151 | 21282 | 23663 | 25398 | 26848 | 27936 | 28812 | 29554 | 30022 | 30571 | 30986 | 31481 | 31521 | 32008 | 32258 | 32452 | 32569 | 32615 | 32875 | 33065 | 33082 | 33144 | 33309 | 33188 | 33351 | 33305
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 225c9e96-1209-ed11-85f4-5cf370a20867 | 17ca54cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.22 | 2575.04 | 8.24 | 28421 | [20 empty] | 010081427802027710118705 2105880 17240328 | 11 | 3 | V790 | 2022-07-21 11:32 | 1419 | 259 | 12.4945 | 3931.0509 | 3 | 12 | 4014 | 4007 | 3996 | 3990 | 3998 | 3986 | 3996 | 4002 | 4013 | 4029 | 4038 | 4072 | 4189 | 4396 | 4782 | 5584 | 7258 | 10236 | 14166 | 17917 | 21050 | 23420 | 25158 | 26637 | 27635 | 28541 | 29200 | 29848 | 30272 | 30687 | 31025 | 31336 | 31651 | 31943 | 32028 | 32211 | 32400 | 32419 | 32603 | 32736 | 32748 | 32749 | 32869 | 32928 | 32925
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 295c9e96-1209-ed11-85f4-5cf370a20867 | b5c954cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.25 | 2811.07 | 8.27 | 29297 | [20 empty] | 010081427802027710118705 2105880 17240328 | 12 | 3 | V790 | 2022-07-21 11:32 | 1453 | 244 | 11.0388 | 4033.9891 | 3 | 12 | 4100 | 4092 | 4088 | 4064 | 4078 | 4194 | 4055 | 4061 | 4090 | 4123 | 4119 | 4172 | 4271 | 4496 | 4873 | 5736 | 7382 | 10533 | 14811 | 19038 | 22182 | 24725 | 26490 | 27958 | 29033 | 29848 | 30468 | 31082 | 31537 | 31910 | 32357 | 32493 | 32742 | 32959 | 33262 | 33182 | 33362 | 33545 | 33574 | 33654 | 33764 | 33751 | 33742 | 33767 | 33876
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 305c9e96-1209-ed11-85f4-5cf370a20867 | 1255c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.65 | 2107.83 | 6.31 | 23255 | 6.73 | 6.55 | 5370317.96 | 3542261.73 | 6.73 | 6.55 |  |  | 5370317.96 | 3542261.73 |  |  | 6.73 | 6.55 | 5370317.96 | 3542261.73 | 6.73 | 6.55 | 5370317.96 | 3542261.73 | 010081427802027710118705 2105879 17240328 | 1 | 4 | V1169 | 2022-07-21 11:32 | 2573 | 296 | 10.5939 | 4394.4255 | 3 | 12 | 4467 | 4426 | 4441 | 4444 | 4447 | 4462 | 4459 | 4445 | 4469 | 4481 | 4506 | 4515 | 4598 | 4696 | 4943 | 5445 | 6473 | 8471 | 11432 | 14636 | 17446 | 19684 | 21295 | 22648 | 23638 | 24381 | 25039 | 25545 | 25915 | 26353 | 26551 | 26794 | 27045 | 27160 | 27311 | 27597 | 27615 | 27790 | 27799 | 27870 | 27938 | 28089 | 28131 | 28193 | 28102
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 375c9e96-1209-ed11-85f4-5cf370a20867 | b354c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.42 | 1950.48 | 6.9 | 23336 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 |  |  | 6309573.44 | 4161794.64 |  |  | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 010081427802027710118705 2105879 17240328 | 2 | 4 | V1169 | 2022-07-21 11:32 | 2058 | 284 | 10.0558 | 3971.7818 | 3 | 12 | 4028 | 4027 | 4009 | 4017 | 4030 | 4028 | 4046 | 4015 | 4022 | 4043 | 4070 | 4117 | 4164 | 4304 | 4558 | 5124 | 6278 | 8394 | 11237 | 14177 | 16687 | 18747 | 20313 | 21638 | 22566 | 23391 | 24115 | 24592 | 25125 | 25534 | 25772 | 26075 | 26301 | 26515 | 26809 | 26994 | 27044 | 27181 | 27312 | 27514 | 27593 | 27574 | 27671 | 27750 | 27793
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 3e5c9e96-1209-ed11-85f4-5cf370a20867 | 7255c0e4-1c09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.3 | 2023.88 | 6.93 | 22748 | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 6.84 | 6.66 |  |  | 6918309.71 | 4563317.08 |  |  | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 010081427802027710118705 2105879 17240328 | 3 | 4 | V1169 | 2022-07-21 11:32 | 1979 | 281 | 10.2691 | 3858.8618 | 3 | 12 | 3934 | 3912 | 3912 | 3899 | 3907 | 3911 | 3921 | 3912 | 3944 | 3934 | 3952 | 3987 | 4056 | 4178 | 4465 | 5049 | 6260 | 8487 | 11528 | 14537 | 16963 | 18870 | 20440 | 21511 | 22439 | 23239 | 23777 | 24244 | 24644 | 25038 | 25286 | 25579 | 25920 | 26007 | 26240 | 26321 | 26465 | 26598 | 26677 | 26772 | 26959 | 26915 | 27015 | 27016 | 27105
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 6042e1f8-1209-ed11-85f4-5cf370a20867 | 233d8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.34 | 2623.16 | 6.33 | 27960 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 |  |  | 6760829.75 | 4459443.3 |  |  | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 010081427802027710118705 2105881 17240328 | 12 | 2 | V834 | 2022-07-21 11:34 | 2839 | 265 | 11.2448 | 5279.6436 | 3 | 12 | 5375 | 5371 | 5342 | 5323 | 5342 | 5325 | 5337 | 5330 | 5344 | 5366 | 5384 | 5426 | 5530 | 5670 | 6041 | 6763 | 8262 | 11089 | 15033 | 18933 | 22118 | 24533 | 26450 | 27814 | 28853 | 29716 | 30369 | 30983 | 31350 | 31803 | 32170 | 32399 | 32595 | 32876 | 33003 | 33112 | 33344 | 33349 | 33462 | 33522 | 33601 | 33751 | 33667 | 33758 | 33767
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 6642e1f8-1209-ed11-85f4-5cf370a20867 | 7e8c0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.35 | 2535.4 | 8.04 | 28768 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 |  |  | 6606934.48 | 4357933.98 |  |  | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 010081427802027710118705 2105880 17240328 | 1 | 3 | V790 | 2022-07-21 11:34 | 1595 | 256 | 12.2133 | 4091.92 | 3 | 12 | 4172 | 4146 | 4149 | 4156 | 4155 | 4158 | 4157 | 4168 | 4157 | 4192 | 4193 | 4252 | 4329 | 4520 | 4896 | 5628 | 7128 | 9905 | 13707 | 17505 | 20670 | 23110 | 24952 | 26372 | 27426 | 28335 | 29108 | 29752 | 30204 | 30703 | 31073 | 31422 | 31752 | 31995 | 32264 | 32452 | 32577 | 32791 | 32864 | 33001 | 33160 | 33219 | 33323 | 33346 | 33459
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 6c42e1f8-1209-ed11-85f4-5cf370a20867 | ef8b0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.32 | 2588.44 | 7.64 | 29845 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 |  |  | 6760829.75 | 4459443.3 |  |  | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 010081427802027710118705 2105880 17240328 | 2 | 3 | V790 | 2022-07-21 11:34 | 1882 | 264 | 9.5576 | 4509.8982 | 3 | 12 | 4594 | 4587 | 4547 | 4563 | 4569 | 4545 | 4555 | 4560 | 4562 | 4553 | 4604 | 4639 | 4729 | 4907 | 5284 | 6030 | 7598 | 10494 | 14358 | 18201 | 21406 | 23897 | 25878 | 27428 | 28560 | 29568 | 30308 | 30996 | 31546 | 31971 | 32447 | 32818 | 33206 | 33395 | 33637 | 33787 | 34012 | 34217 | 34247 | 34360 | 34528 | 34604 | 34732 | 34792 | 34801
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 7242e1f8-1209-ed11-85f4-5cf370a20867 | bf8b0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.37 | 2773.69 | 8.52 | 30635 | [20 empty] | 010081427802027710118705 2105880 17240328 | 3 | 3 | V790 | 2022-07-21 11:34 | 1410 | 242 | 14.6582 | 4079.0836 | 3 | 12 | 4144 | 4152 | 4139 | 4159 | 4160 | 4155 | 4164 | 4152 | 4178 | 4195 | 4223 | 4266 | 4351 | 4538 | 4927 | 5683 | 7263 | 10235 | 14392 | 18557 | 21943 | 24553 | 26497 | 28019 | 29105 | 30070 | 30791 | 31420 | 31948 | 32447 | 32905 | 33276 | 33564 | 33869 | 34122 | 34268 | 34510 | 34707 | 34835 | 34960 | 35101 | 35115 | 35276 | 35346 | 35415
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 7842e1f8-1209-ed11-85f4-5cf370a20867 | 608b0127-2009-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.46 | 2532.36 | 8.41 | 30118 | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 6.79 | 6.61 |  |  | 6165950.02 | 4067060.63 |  |  | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 010081427802027710118705 2105880 17240328 | 4 | 3 | V790 | 2022-07-21 11:34 | 1354 | 259 | 12.6812 | 4069.7309 | 3 | 12 | 4133 | 4139 | 4135 | 4131 | 4128 | 4133 | 4143 | 4146 | 4157 | 4167 | 4197 | 4218 | 4302 | 4486 | 4849 | 5530 | 6984 | 9667 | 13396 | 17217 | 20527 | 23057 | 25178 | 26825 | 28109 | 29049 | 29895 | 30520 | 31157 | 31655 | 32133 | 32419 | 32817 | 33103 | 33356 | 33636 | 33858 | 34050 | 34193 | 34356 | 34406 | 34486 | 34657 | 34751 | 34794
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 7e42e1f8-1209-ed11-85f4-5cf370a20867 | ebaadb1d-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified | 2610 (Informational, Delays in sample processing) | 16.12 | 1603.89 | 6.52 | 16922 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 |  |  | 7762471.17 | 5120125.98 |  |  | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 010081427802027710118705 2105880 17240328 | 5 | 3 | V790 | 2022-07-21 11:34 | 1372 | 268 | 2.2497 | 3104.4473 | 3 | 12 | 3105 | 3103 | 3124 | 3126 | 3132 | 3128 | 3123 | 3094 | 3091 | 3093 | 3084 | 3168 | 3199 | 3336 | 3582 | 4130 | 5141 | 6950 | 9485 | 11869 | 13452 | 14885 | 16099 | 16778 | 17310 | 17890 | 18230 | 18622 | 19012 | 19183 | 19360 | 19480 | 19449 | 19720 | 19563 | 19849 | 19986 | 19853 | 19880 | 19984 | 19914 | 19530 | 19591 | 20062 | 20328
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 8442e1f8-1209-ed11-85f4-5cf370a20867 | 57c954cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.29 | 2625.53 | 8.37 | 30075 | [20 empty] | 010081427802027710118705 2105880 17240328 | 6 | 3 | V790 | 2022-07-21 11:34 | 1388 | 245 | 11.3673 | 4079.7455 | 3 | 12 | 4128 | 4127 | 4124 | 4130 | 4124 | 4235 | 4116 | 4113 | 4130 | 4171 | 4172 | 4229 | 4320 | 4516 | 4914 | 5718 | 7289 | 10243 | 14193 | 18133 | 21287 | 23841 | 25807 | 27324 | 28492 | 29410 | 30185 | 30819 | 31510 | 31879 | 32351 | 32682 | 32920 | 33181 | 33578 | 33667 | 33878 | 34050 | 34125 | 34253 | 34375 | 34480 | 34552 | 34592 | 34729
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 8a42e1f8-1209-ed11-85f4-5cf370a20867 | edc854cd-1e09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.13 | 2178.33 | 8.18 | 26676 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 |  |  | 7762471.17 | 5120125.98 |  |  | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 010081427802027710118705 2105880 17240328 | 7 | 3 | V790 | 2022-07-21 11:34 | 1317 | 270 | 11.3709 | 3719.5782 | 3 | 12 | 3779 | 3769 | 3785 | 3770 | 3790 | 3782 | 3787 | 3772 | 3783 | 3792 | 3822 | 3863 | 3968 | 4148 | 4514 | 5312 | 6887 | 9537 | 12837 | 15993 | 18684 | 20840 | 22557 | 23866 | 25049 | 25937 | 26662 | 27199 | 27796 | 28164 | 28666 | 28911 | 29225 | 29488 | 29729 | 29886 | 30109 | 30238 | 30369 | 30451 | 30673 | 30796 | 30791 | 30888 | 30952
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | 5a7b815a-1309-ed11-85f4-5cf370a20867 | 223c8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.28 | 2631.98 | 6.57 | 27149 | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 6.85 | 6.67 |  |  | 7079457.84 | 4669610.39 |  |  | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 010081427802027710118705 2105881 17240328 | 7 | 2 | V834 | 2022-07-21 11:37 | 2538 | 268 | 11.423 | 4907.5673 | 3 | 12 | 4997 | 4971 | 4993 | 4972 | 4954 | 4956 | 4956 | 4973 | 4970 | 4995 | 5012 | 5047 | 5152 | 5343 | 5721 | 6458 | 8061 | 11004 | 14983 | 18892 | 22022 | 24368 | 26070 | 27323 | 28292 | 28978 | 29654 | 30188 | 30575 | 30803 | 31095 | 31354 | 31569 | 31775 | 31876 | 32095 | 32030 | 32266 | 32329 | 32379 | 32446 | 32361 | 32533 | 32564 | 32585
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | 607b815a-1309-ed11-85f4-5cf370a20867 | 50597999-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.37 | 2314.8 | 6.2 | 25184 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 |  |  | 6606934.48 | 4357933.98 |  |  | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 010081427802027710118705 2105881 17240328 | 8 | 2 | V834 | 2022-07-21 11:37 | 2749 | 292 | 8.6594 | 4878.6545 | 3 | 12 | 4950 | 4950 | 4924 | 4929 | 4918 | 4913 | 4925 | 4917 | 4925 | 4932 | 4956 | 5001 | 5073 | 5219 | 5533 | 6158 | 7462 | 9936 | 13419 | 16856 | 19686 | 21900 | 23546 | 24796 | 25744 | 26623 | 27401 | 27795 | 28296 | 28641 | 28928 | 29167 | 29473 | 29645 | 29729 | 29934 | 30041 | 30026 | 30150 | 30296 | 30231 | 30299 | 30369 | 30417 | 30492
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | 667b815a-1309-ed11-85f4-5cf370a20867 | 9c3e8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.45 | 2392.68 | 5.45 | 25142 | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 6.79 | 6.61 |  |  | 6165950.02 | 4067060.63 |  |  | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 6.79 | 6.61 | 6165950.02 | 4067060.63 | 010081427802027710118705 2105881 17240328 | 9 | 2 | V834 | 2022-07-21 11:37 | 3532 | 272 | 7.1382 | 5690.0436 | 3 | 12 | 5760 | 5743 | 5739 | 5723 | 5719 | 5728 | 5736 | 5717 | 5715 | 5723 | 5748 | 5810 | 5846 | 6020 | 6305 | 6910 | 8215 | 10685 | 14208 | 17831 | 20807 | 23050 | 24798 | 25995 | 26995 | 27665 | 28301 | 28715 | 29200 | 29505 | 29674 | 30014 | 30234 | 30394 | 30471 | 30658 | 30650 | 30867 | 30830 | 30892 | 30982 | 30984 | 31019 | 31156 | 31197
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | 6c7b815a-1309-ed11-85f4-5cf370a20867 | 3e3e8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.38 | 2436.69 | 5.53 | 25415 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 |  |  | 6606934.48 | 4357933.98 |  |  | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 6.82 | 6.64 | 6606934.48 | 4357933.98 | 010081427802027710118705 2105881 17240328 | 10 | 2 | V834 | 2022-07-21 11:37 | 3437 | 275 | 7.8133 | 5650.56 | 3 | 12 | 5704 | 5729 | 5696 | 5710 | 5676 | 5687 | 5681 | 5688 | 5677 | 5702 | 5704 | 5764 | 5859 | 5987 | 6289 | 6951 | 8316 | 10893 | 14529 | 18165 | 21138 | 23344 | 25052 | 26230 | 27204 | 27859 | 28515 | 28937 | 29396 | 29775 | 29980 | 30341 | 30416 | 30567 | 30724 | 30921 | 31003 | 31052 | 31138 | 31203 | 31252 | 31291 | 31413 | 31471 | 31400
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | 727b815a-1309-ed11-85f4-5cf370a20867 | e03d8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.44 | 2298.39 | 5.71 | 25080 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 |  |  | 6309573.44 | 4161794.64 |  |  | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 010081427802027710118705 2105881 17240328 | 11 | 2 | V834 | 2022-07-21 11:37 | 3289 | 280 | 6.577 | 5366.3527 | 3 | 12 | 5452 | 5420 | 5404 | 5421 | 5399 | 5390 | 5380 | 5389 | 5397 | 5398 | 5407 | 5468 | 5555 | 5689 | 5971 | 6578 | 7831 | 10260 | 13604 | 17108 | 19982 | 22069 | 23814 | 25092 | 26144 | 26982 | 27578 | 28120 | 28568 | 28888 | 29244 | 29593 | 29650 | 29922 | 30171 | 30149 | 30216 | 30354 | 30546 | 30526 | 30609 | 30680 | 30689 | 30719 | 30768
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | 787b815a-1309-ed11-85f4-5cf370a20867 | 1eff3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.43 | 2367.47 | 6.98 | 25496 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 |  |  | 6309573.44 | 4161794.64 |  |  | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 6.8 | 6.62 | 6309573.44 | 4161794.64 | 010081427802027710118705 2105881 17240328 | 4 | 2 | V834 | 2022-07-21 11:37 | 2183 | 247 | 10.5467 | 4277.56 | 3 | 12 | 4343 | 4344 | 4343 | 4328 | 4321 | 4334 | 4330 | 4338 | 4338 | 4363 | 4369 | 4409 | 4485 | 4668 | 4953 | 5566 | 6880 | 9355 | 12859 | 16469 | 19359 | 21627 | 23331 | 24541 | 25557 | 26488 | 27073 | 27581 | 28008 | 28455 | 28779 | 29003 | 29217 | 29310 | 29506 | 29601 | 29834 | 29837 | 29991 | 30012 | 30127 | 30197 | 30285 | 30238 | 30239
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 7e7b815a-1309-ed11-85f4-5cf370a20867 | 823d8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.4 | 2435.14 | 6.91 | 25271 | 6.81 | 6.63 | 6456542.29 | 4258735.29 | 6.81 | 6.63 |  |  | 6456542.29 | 4258735.29 |  |  | 6.81 | 6.63 | 6456542.29 | 4258735.29 | 6.81 | 6.63 | 6456542.29 | 4258735.29 | 010081427802027710118705 2105881 17240328 | 5 | 2 | V834 | 2022-07-21 11:37 | 2203 | 258 | 11.3697 | 4297.6873 | 3 | 12 | 4364 | 4351 | 4353 | 4360 | 4355 | 4351 | 4365 | 4371 | 4358 | 4382 | 4398 | 4454 | 4521 | 4677 | 4977 | 5626 | 6978 | 9534 | 13177 | 16805 | 19806 | 22067 | 23725 | 24953 | 25953 | 26658 | 27302 | 27783 | 28060 | 28492 | 28703 | 28974 | 29165 | 29362 | 29467 | 29638 | 29692 | 29749 | 29808 | 29946 | 30005 | 30050 | 29976 | 30069 | 30119
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | 847b815a-1309-ed11-85f4-5cf370a20867 | c43c8093-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.18 | 2359.93 | 6.51 | 24908 | 6.88 | 6.7 | 7585775.75 | 5003577.68 | 6.88 | 6.7 |  |  | 7585775.75 | 5003577.68 |  |  | 6.88 | 6.7 | 7585775.75 | 5003577.68 | 6.88 | 6.7 | 7585775.75 | 5003577.68 | 010081427802027710118705 2105881 17240328 | 6 | 2 | V834 | 2022-07-21 11:37 | 2433 | 253 | 11.0485 | 4546.5564 | 3 | 12 | 4607 | 4604 | 4609 | 4603 | 4604 | 4599 | 4612 | 4608 | 4614 | 4614 | 4636 | 4694 | 4807 | 4922 | 5307 | 6020 | 7534 | 10295 | 13879 | 17293 | 20039 | 22148 | 23739 | 24825 | 25813 | 26490 | 27118 | 27592 | 27930 | 28363 | 28514 | 28864 | 28937 | 29132 | 29245 | 29441 | 29479 | 29544 | 29628 | 29748 | 29717 | 29793 | 29842 | 29992 | 29975
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | b3de19db-1309-ed11-85f4-5cf370a20867 | 685f6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.51 | 2123.17 | 6.78 | 22703 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 |  |  | 6025595.86 | 3974483.03 |  |  | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 010081427802027710118705 2105881 17240328 | 3 | 2 | V834 | 2022-07-21 11:41 | 2093 | 263 | 7.3515 | 3941.8836 | 3 | 12 | 3995 | 3979 | 3977 | 3974 | 3994 | 3980 | 3973 | 3976 | 3992 | 3984 | 4017 | 4025 | 4099 | 4235 | 4490 | 5044 | 6137 | 8272 | 11346 | 14611 | 17260 | 19341 | 21000 | 22109 | 23093 | 23806 | 24271 | 24775 | 25110 | 25411 | 25746 | 25859 | 26056 | 26284 | 26329 | 26492 | 26667 | 26648 | 26765 | 26834 | 26845 | 26898 | 26928 | 27008 | 26980
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | b9de19db-1309-ed11-85f4-5cf370a20867 | 785e6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.61 | 2768.06 | 7.99 | 29669 | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 6.75 | 6.57 |  |  | 5623413.25 | 3709203.38 |  |  | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 010081427802027710118705 2105882 17240328 | 10 | 1 | V639 | 2022-07-21 11:41 | 1720 | 268 | 12.9721 | 4239.3891 | 3 | 12 | 4298 | 4287 | 4291 | 4304 | 4313 | 4314 | 4307 | 4337 | 4319 | 4333 | 4354 | 4407 | 4487 | 4647 | 4968 | 5609 | 6974 | 9603 | 13580 | 17755 | 21402 | 24105 | 26168 | 27739 | 28912 | 29985 | 30687 | 31264 | 31824 | 32180 | 32673 | 32842 | 33144 | 33380 | 33594 | 33791 | 33977 | 34063 | 34104 | 34192 | 34286 | 34432 | 34477 | 34505 | 34493
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | bfde19db-1309-ed11-85f4-5cf370a20867 | 485e6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.29 | 2777.15 | 8.24 | 29533 | [20 empty] | 010081427802027710118705 2105882 17240328 | 11 | 1 | V639 | 2022-07-21 11:41 | 1560 | 266 | 14.7564 | 4080.1273 | 3 | 12 | 4140 | 4148 | 4158 | 4161 | 4145 | 4174 | 4157 | 4163 | 4169 | 4194 | 4225 | 4255 | 4375 | 4561 | 4935 | 5747 | 7426 | 10506 | 14711 | 18795 | 22172 | 24613 | 26533 | 27883 | 28994 | 30747 | 30672 | 31283 | 31754 | 32059 | 32501 | 32708 | 32976 | 33194 | 33362 | 33628 | 34648 | 33908 | 33900 | 34036 | 34061 | 34220 | 34203 | 34321 | 34287
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | c5de19db-1309-ed11-85f4-5cf370a20867 | 90fe3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.47 | 2893.71 | 8.04 | 30993 | [20 empty] | 010081427802027710118705 2105882 17240328 | 12 | 1 | V639 | 2022-07-21 11:41 | 1817 | 263 | 16.143 | 4407.3273 | 3 | 12 | 4471 | 4497 | 4458 | 4491 | 4495 | 4501 | 4497 | 4499 | 4518 | 4541 | 4560 | 4606 | 4703 | 4898 | 5246 | 5996 | 7566 | 10553 | 14817 | 19191 | 22840 | 25630 | 27792 | 29373 | 30643 | 31501 | 32345 | 32976 | 33519 | 33827 | 34261 | 34540 | 34808 | 35039 | 35185 | 35330 | 35592 | 35687 | 35688 | 35801 | 35881 | 35925 | 36009 | 36141 | 36161
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | cbde19db-1309-ed11-85f4-5cf370a20867 | 60fe3abb-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.49 | 2361.31 | 5.95 | 25803 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 |  |  | 6025595.86 | 3974483.03 |  |  | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 010081427802027710118705 2105881 17240328 | 1 | 2 | V834 | 2022-07-21 11:41 | 3061 | 258 | 8.7067 | 5253.04 | 3 | 12 | 5299 | 5300 | 5308 | 5299 | 5310 | 5304 | 5292 | 5269 | 5296 | 5317 | 5346 | 5367 | 5445 | 5583 | 5853 | 6495 | 7770 | 10187 | 13612 | 17209 | 20223 | 22519 | 24374 | 25644 | 26728 | 27565 | 28246 | 28823 | 29320 | 29618 | 29903 | 30149 | 30383 | 30551 | 30743 | 30865 | 30965 | 31079 | 31237 | 31238 | 31338 | 31420 | 31366 | 31391 | 31494
BKV 10m4 U R6 |  | BKV URINE 10 | L999 | 0.0.40 | d1de19db-1309-ed11-85f4-5cf370a20867 | 985f6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.29 | 2373.91 | 6.32 | 24689 | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 6.84 | 6.66 |  |  | 6918309.71 | 4563317.08 |  |  | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 6.84 | 6.66 | 6918309.71 | 4563317.08 | 010081427802027710118705 2105881 17240328 | 2 | 2 | V834 | 2022-07-21 11:41 | 2647 | 271 | 10.5273 | 4659.0655 | 3 | 12 | 4716 | 4707 | 4717 | 4724 | 4708 | 4707 | 4719 | 4726 | 4720 | 4728 | 4748 | 4798 | 4874 | 5050 | 5394 | 6049 | 7462 | 10084 | 13668 | 17232 | 20012 | 22148 | 23743 | 24855 | 25847 | 26501 | 27124 | 27482 | 27943 | 28156 | 28538 | 28620 | 28865 | 29123 | 29129 | 29265 | 29453 | 29464 | 29550 | 29709 | 29638 | 29682 | 29842 | 29872 | 29799
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | d7de19db-1309-ed11-85f4-5cf370a20867 | 954775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.27 | 2874.99 | 8.13 | 29909 | [20 empty] | 010081427802027710118705 2105882 17240328 | 8 | 1 | V639 | 2022-07-21 11:41 | 1689 | 271 | 15.3721 | 4204.9491 | 3 | 12 | 4278 | 4278 | 4271 | 4291 | 4289 | 4294 | 4287 | 4282 | 4303 | 4312 | 4351 | 4402 | 4524 | 4708 | 5117 | 5958 | 7701 | 10907 | 15256 | 19537 | 22891 | 25443 | 27267 | 28657 | 29774 | 30634 | 31338 | 31862 | 32345 | 32686 | 32988 | 33354 | 33536 | 33819 | 34025 | 34164 | 34300 | 34493 | 34588 | 34660 | 34699 | 34819 | 34724 | 34773 | 34844
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | ddde19db-1309-ed11-85f4-5cf370a20867 | db5e6bb4-2209-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.06 | 2821.6 | 7.95 | 29515 | [20 empty] | 010081427802027710118705 2105882 17240328 | 9 | 1 | V639 | 2022-07-21 11:41 | 1825 | 274 | 18.3176 | 4254.1782 | 3 | 12 | 4338 | 4363 | 4341 | 4345 | 4345 | 4346 | 4354 | 4369 | 4368 | 4394 | 4424 | 4504 | 4597 | 4829 | 5299 | 6264 | 8246 | 11717 | 16027 | 20075 | 23251 | 25654 | 27434 | 28717 | 29790 | 30498 | 31221 | 31730 | 32196 | 32510 | 32907 | 33201 | 33380 | 33658 | 33713 | 33882 | 34032 | 34084 | 34201 | 34288 | 34486 | 34385 | 34491 | 34646 | 34610
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 65fdae7a-1409-ed11-85f4-5cf370a20867 | 364775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.27 | 2347.62 | 7.47 | 26286 | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 6.85 | 6.67 |  |  | 7079457.84 | 4669610.39 |  |  | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 6.85 | 6.67 | 7079457.84 | 4669610.39 | 010081427802027710118705 2105882 17240328 | 4 | 1 | V639 | 2022-07-21 11:45 | 1808 | 260 | 11.5855 | 4075.4291 | 3 | 12 | 4128 | 4123 | 4136 | 4142 | 4139 | 4139 | 4135 | 4136 | 4148 | 4158 | 4174 | 4231 | 4309 | 4484 | 4848 | 5532 | 7002 | 9619 | 13209 | 16663 | 19489 | 21750 | 23518 | 25165 | 25831 | 26683 | 27349 | 27935 | 28353 | 28754 | 29137 | 29316 | 29638 | 29835 | 30070 | 30213 | 30268 | 30437 | 30535 | 30624 | 30685 | 30777 | 30796 | 30858 | 30920
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 6bfdae7a-1409-ed11-85f4-5cf370a20867 | a94675b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.02 | 2809.5 | 7.97 | 30622 | 6.92 | 6.74 | 8317637.71 | 5486313.83 | 6.92 | 6.74 |  |  | 8317637.71 | 5486313.83 |  |  | 6.92 | 6.74 | 8317637.71 | 5486313.83 | 6.92 | 6.74 | 8317637.71 | 5486313.83 | 010081427802027710118705 2105882 17240328 | 5 | 1 | V639 | 2022-07-21 11:45 | 1826 | 257 | 19.0012 | 4401.4909 | 3 | 12 | 4487 | 4499 | 4493 | 4483 | 4506 | 4495 | 4523 | 4511 | 4530 | 4544 | 4573 | 4640 | 4771 | 5005 | 5499 | 6502 | 8573 | 12081 | 16413 | 20388 | 23554 | 26004 | 27790 | 29241 | 30295 | 31199 | 31934 | 32640 | 33130 | 33527 | 34040 | 34383 | 34571 | 34985 | 35007 | 35149 | 35421 | 35406 | 35526 | 35682 | 35747 | 35781 | 35747 | 35871 | 35925
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 71fdae7a-1409-ed11-85f4-5cf370a20867 | 774675b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.14 | 2540.68 | 7.78 | 28700 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 |  |  | 7762471.17 | 5120125.98 |  |  | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 6.89 | 6.71 | 7762471.17 | 5120125.98 | 010081427802027710118705 2105882 17240328 | 6 | 1 | V639 | 2022-07-21 11:45 | 1746 | 262 | 16.9309 | 4242.9382 | 3 | 12 | 4331 | 4325 | 4335 | 4329 | 4319 | 4327 | 4343 | 4342 | 4355 | 4370 | 4420 | 4446 | 4561 | 4774 | 5185 | 6013 | 7794 | 10833 | 14718 | 18385 | 21411 | 23720 | 25554 | 26997 | 28003 | 28961 | 29688 | 30374 | 30793 | 31340 | 31611 | 32011 | 32203 | 32494 | 32636 | 32899 | 33055 | 33122 | 33375 | 33447 | 33485 | 33510 | 33618 | 33751 | 33719
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 77fdae7a-1409-ed11-85f4-5cf370a20867 | c64775b5-2309-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.03 | 2805.63 | 8.16 | 29973 | [20 empty] | 010081427802027710118705 2105882 17240328 | 7 | 1 | V639 | 2022-07-21 11:45 | 1646 | 261 | 19.9042 | 4183.5382 | 3 | 12 | 4272 | 4292 | 4277 | 4281 | 4278 | 4275 | 4317 | 4302 | 4323 | 4330 | 4374 | 4430 | 4560 | 4799 | 5290 | 6285 | 8302 | 11805 | 16113 | 20133 | 23250 | 25705 | 27407 | 28751 | 29843 | 30663 | 31354 | 31961 | 32534 | 32870 | 33290 | 33574 | 33773 | 33930 | 34129 | 34300 | 34438 | 34600 | 34665 | 34820 | 34809 | 34900 | 35041 | 35079 | 35047
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | f47fc1ae-1609-ed11-85f4-5cf370a20867 | 816b47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.14 | 2841.9 | 7.99 | 29636 | [20 empty] | 010081427802027710118705 2105882 17240328 | 1 | 1 | V639 | 2022-07-21 12:01 | 1679 | 254 | 18.6121 | 4231.3491 | 3 | 12 | 4309 | 4326 | 4309 | 4317 | 4330 | 4327 | 4346 | 4354 | 4347 | 4382 | 4407 | 4467 | 4577 | 4778 | 5234 | 6139 | 8007 | 11349 | 15710 | 19836 | 23064 | 25485 | 27300 | 28626 | 29731 | 30524 | 31237 | 31737 | 32287 | 32608 | 32912 | 33302 | 33530 | 33734 | 33823 | 33999 | 34178 | 34254 | 34439 | 34421 | 34530 | 34616 | 34726 | 34779 | 34673
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | fa7fc1ae-1609-ed11-85f4-5cf370a20867 | 226b47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.12 | 3139.19 | 8.24 | 32203 | [20 empty] | 010081427802027710118705 2105882 17240328 | 2 | 1 | V639 | 2022-07-21 12:01 | 1768 | 263 | 19.5394 | 4452.2145 | 3 | 12 | 4536 | 4548 | 4549 | 4554 | 4559 | 4554 | 4538 | 4583 | 4578 | 4607 | 4625 | 4705 | 4825 | 5083 | 5575 | 6623 | 8678 | 12387 | 17252 | 21808 | 25295 | 27953 | 29848 | 31222 | 32268 | 33152 | 33899 | 34540 | 34984 | 35357 | 35691 | 35913 | 36161 | 36433 | 36580 | 36792 | 36954 | 37094 | 37197 | 37278 | 37402 | 37404 | 37424 | 37586 | 37554
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | 0080c1ae-1609-ed11-85f4-5cf370a20867 | e16b47c6-2609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Indeterminate | TargetIndeterminate | 1012 (Error, Outlier Removal Failed) | 16.13 | 3191.32 | 8.32 | 32904 | [20 empty] | 010081427802027710118705 2105882 17240328 | 3 | 1 | V639 | 2022-07-21 12:01 | 1736 | 265 | 18.4218 | 4505.0364 | 3 | 12 | 4614 | 4585 | 4592 | 4608 | 4594 | 4611 | 4610 | 4596 | 4619 | 4647 | 4679 | 4732 | 4887 | 5097 | 5608 | 6602 | 8698 | 12513 | 17410 | 21983 | 25625 | 28212 | 30189 | 31644 | 32789 | 33663 | 34409 | 35009 | 35557 | 35988 | 36420 | 36703 | 36952 | 37210 | 37263 | 37431 | 37634 | 37717 | 37781 | 37977 | 38028 | 38065 | 38138 | 38201 | 38283
BKV 10m4 U R1 |  | BKV URINE 10 | L999 | 0.0.40 | a79dad68-1f09-ed11-85f4-5cf370a20867 | b26dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R2 |  | BKV URINE 10 | L999 | 0.0.40 | ad9dad68-1f09-ed11-85f4-5cf370a20867 | b66dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | b39dad68-1f09-ed11-85f4-5cf370a20867 | ba6dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | b99dad68-1f09-ed11-85f4-5cf370a20867 | be6dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R5 |  | BKV URINE 10 | L999 | 0.0.40 | bf9dad68-1f09-ed11-85f4-5cf370a20867 | c26dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | c59dad68-1f09-ed11-85f4-5cf370a20867 | c76dfdc6-1f09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:03 |  |  |  |  | 0
BKV 10m4 U R10 |  | BKV URINE 10 | L999 | 0.0.40 | cb9dad68-1f09-ed11-85f4-5cf370a20867 | 2f0bca23-2809-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 17.03 | 2397.7 | 7.16 | 25393 | 6.62 | 6.44 | 4168693.83 | 2749670.45 | 6.62 | 6.44 |  |  | 4168693.83 | 2749670.45 |  |  | 6.62 | 6.44 | 4168693.83 | 2749670.45 | 6.62 | 6.44 | 4168693.83 | 2749670.45 | 010081427802027710118705 2105883 17240328 | 11 | 4 | V1169 | 2022-07-21 13:03 | 2048 | 285 | 11.6691 | 4140.7564 | 3 | 13 | 4175 | 4235 | 4206 | 4192 | 4192 | 4211 | 4225 | 4200 | 4220 | 4222 | 4219 | 4263 | 4311 | 4411 | 4593 | 5007 | 5872 | 7540 | 10498 | 14239 | 17585 | 20266 | 22219 | 23704 | 24817 | 25780 | 26400 | 27002 | 27557 | 27834 | 28156 | 28451 | 28793 | 28964 | 29076 | 29221 | 29386 | 29543 | 29706 | 29690 | 29775 | 29955 | 29907 | 30047 | 30113
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | d29dad68-1f09-ed11-85f4-5cf370a20867 | d10aca23-2809-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.59 | 2333.73 | 6.98 | 25183 | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 6.75 | 6.57 |  |  | 5623413.25 | 3709203.38 |  |  | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 6.75 | 6.57 | 5623413.25 | 3709203.38 | 010081427802027710118705 2105883 17240328 | 12 | 4 | V1169 | 2022-07-21 13:03 | 2088 | 278 | 10.3709 | 4218.2182 | 3 | 12 | 4261 | 4262 | 4267 | 4264 | 4257 | 4292 | 4285 | 4297 | 4265 | 4308 | 4314 | 4341 | 4409 | 4541 | 4797 | 5353 | 6469 | 8682 | 12004 | 15619 | 18541 | 20830 | 22541 | 23840 | 24818 | 25661 | 26257 | 26805 | 27330 | 27610 | 27981 | 28258 | 28590 | 28751 | 28859 | 29152 | 29236 | 29434 | 29584 | 29563 | 29716 | 29813 | 29820 | 29873 | 29882
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 1f42d920-2009-ed11-85f4-5cf370a20867 | 41e5bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.64 | 2405.59 | 7.37 | 25353 | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 6.74 | 6.56 |  |  | 5495408.74 | 3624771.6 |  |  | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 010081427802027710118705 2105883 17240328 | 5 | 4 | V1169 | 2022-07-21 13:09 | 1954 | 288 | 13.0194 | 3986.2545 | 3 | 12 | 4044 | 4045 | 4051 | 4051 | 4043 | 4048 | 4071 | 4074 | 4083 | 4096 | 4089 | 4140 | 4233 | 4387 | 4652 | 5225 | 6425 | 8694 | 12076 | 15787 | 18942 | 21379 | 23028 | 24372 | 25356 | 26111 | 26746 | 27378 | 27630 | 28125 | 28324 | 28555 | 28841 | 29006 | 29179 | 29363 | 29481 | 29442 | 29658 | 29608 | 29756 | 29715 | 29877 | 29949 | 29934
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 2642d920-2009-ed11-85f4-5cf370a20867 | b3e4bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.08 | 2303.83 | 7.28 | 24419 | 6.91 | 6.73 | 8128305.16 | 5361430.08 | 6.91 | 6.73 |  |  | 8128305.16 | 5361430.08 |  |  | 6.91 | 6.73 | 8128305.16 | 5361430.08 | 6.91 | 6.73 | 8128305.16 | 5361430.08 | 010081427802027710118705 2105883 17240328 | 6 | 4 | V1169 | 2022-07-21 13:09 | 1944 | 278 | 15.3091 | 3909.1018 | 3 | 12 | 3993 | 3988 | 3994 | 3979 | 3985 | 3986 | 4003 | 3987 | 4006 | 4040 | 4032 | 4116 | 4208 | 4402 | 4800 | 5600 | 7214 | 10051 | 13586 | 16880 | 19580 | 21558 | 22978 | 24046 | 24891 | 25642 | 26119 | 26670 | 26942 | 27275 | 27551 | 27810 | 28067 | 28226 | 28325 | 28482 | 28575 | 28655 | 28819 | 28809 | 28891 | 29014 | 28966 | 28976 | 29048
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 2c42d920-2009-ed11-85f4-5cf370a20867 | bde6bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.63 | 2270.38 | 7.03 | 24694 | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 6.74 | 6.56 |  |  | 5495408.74 | 3624771.6 |  |  | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 6.74 | 6.56 | 5495408.74 | 3624771.6 | 010081427802027710118705 2105883 17240328 | 7 | 4 | V1169 | 2022-07-21 13:09 | 2123 | 298 | 10.1176 | 4103.1782 | 3 | 12 | 4167 | 4143 | 4154 | 4153 | 4156 | 4160 | 4164 | 4166 | 4156 | 4164 | 4196 | 4251 | 4306 | 4447 | 4703 | 5254 | 6362 | 8533 | 11747 | 15197 | 18220 | 20493 | 22141 | 23451 | 24420 | 25224 | 25944 | 26503 | 26852 | 27216 | 27541 | 27910 | 28021 | 28222 | 28467 | 28673 | 28682 | 28790 | 28927 | 28997 | 29033 | 29198 | 29233 | 29225 | 29268
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 3242d920-2009-ed11-85f4-5cf370a20867 | 5ee6bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.32 | 2201.98 | 6.92 | 24142 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 |  |  | 6760829.75 | 4459443.3 |  |  | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 6.83 | 6.65 | 6760829.75 | 4459443.3 | 010081427802027710118705 2105883 17240328 | 8 | 4 | V1169 | 2022-07-21 13:09 | 2146 | 296 | 13.8606 | 4100.0255 | 3 | 12 | 4173 | 4177 | 4162 | 4162 | 4162 | 4181 | 4175 | 4201 | 4175 | 4226 | 4216 | 4274 | 4366 | 4520 | 4855 | 5510 | 6804 | 9244 | 12573 | 15859 | 18563 | 20631 | 22139 | 23306 | 24295 | 24964 | 25544 | 26144 | 26470 | 26834 | 27134 | 27422 | 27679 | 27800 | 28055 | 28176 | 28392 | 28489 | 28564 | 28548 | 28642 | 28829 | 28807 | 28811 | 28903
BKV 10m4 U R11 |  | BKV URINE 10 | L999 | 0.0.40 | 3842d920-2009-ed11-85f4-5cf370a20867 | d2e5bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.73 | 2195.79 | 7.21 | 24578 | 6.71 | 6.53 | 5128613.84 | 3382833.69 | 6.71 | 6.53 |  |  | 5128613.84 | 3382833.69 |  |  | 6.71 | 6.53 | 5128613.84 | 3382833.69 | 6.71 | 6.53 | 5128613.84 | 3382833.69 | 010081427802027710118705 2105883 17240328 | 9 | 4 | V1169 | 2022-07-21 13:09 | 1973 | 289 | 7.6679 | 3971.2109 | 3 | 12 | 4028 | 4017 | 4011 | 4017 | 4001 | 4006 | 4013 | 4019 | 4008 | 4037 | 4042 | 4068 | 4118 | 4245 | 4479 | 4957 | 5933 | 7896 | 10952 | 14312 | 17232 | 19572 | 21298 | 22684 | 23787 | 24666 | 25257 | 25875 | 26386 | 26668 | 27007 | 27326 | 27513 | 27731 | 27900 | 28079 | 28288 | 28378 | 28439 | 28525 | 28646 | 28688 | 28816 | 28899 | 28918
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 3e42d920-2009-ed11-85f4-5cf370a20867 | a1e5bc8c-2a09-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 | 0 | BKV | Urine | Positive | TargetAmplified |  | 16.5 | 2247.9 | 6.88 | 24232 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 |  |  | 6025595.86 | 3974483.03 |  |  | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 6.78 | 6.6 | 6025595.86 | 3974483.03 | 010081427802027710118705 2105883 17240328 | 10 | 4 | V1169 | 2022-07-21 13:09 | 2142 | 280 | 10.1648 | 4127.8836 | 3 | 12 | 4174 | 4170 | 4175 | 4180 | 4196 | 4177 | 4176 | 4195 | 4193 | 4187 | 4223 | 4252 | 4349 | 4481 | 4766 | 5329 | 6514 | 8788 | 12102 | 15457 | 18347 | 20496 | 22098 | 23295 | 24218 | 24956 | 25670 | 26048 | 26485 | 26830 | 27150 | 27408 | 27628 | 27810 | 27997 | 28033 | 28300 | 28380 | 28476 | 28640 | 28606 | 28662 | 28797 | 28857 | 28810
BKV 10m4 U R12 |  | BKV URINE 10 | L999 | 0.0.40 | 4442d920-2009-ed11-85f4-5cf370a20867 | 813d7504-2109-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 13:09 |  |  |  |  | 0
BKV 10m4 U R3 |  | BKV URINE 10 | L999 | 0.0.40 | b7477ca2-3609-ed11-85f4-5cf370a20867 | f3c2a3da-3609-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:50 |  |  |  |  | 0
BKV 10m4 U R4 |  | BKV URINE 10 | L999 | 0.0.40 | f5c2a3da-3609-ed11-85f4-5cf370a20867 | 8eb7eb13-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:51 |  |  |  |  | 0
BKV 10m4 U R7 |  | BKV URINE 10 | L999 | 0.0.40 | 90b7eb13-3709-ed11-85f4-5cf370a20867 | b8bbce4d-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:53 |  |  |  |  | 0
BKV 10m4 U R8 |  | BKV URINE 10 | L999 | 0.0.40 | f2bbce4d-3709-ed11-85f4-5cf370a20867 | 0c592f87-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | NoResult | NoResult | 7162 (Warning, Quantity not sufficient), 1002 (Error, Insufficient Raw Readings) |  |  | 0 | [25 empty] | 2022-07-21 15:54 |  |  |  |  | 0
BKV 10m4 U R9 |  | BKV URINE 10 | L999 | 0.0.40 | 0e592f87-3709-ed11-85f4-5cf370a20867 | 4f3e36c2-3709-ed11-85f4-5cf370a20867 |  | 1 | 1 | 1 | PP1 |  | BKV | Urine | Aborted | TargetAborted | 7162 (Warning, Quantity not sufficient), 1000 (Error, User Aborted Test) |  |  | 0 | [25 empty] | 2022-07-21 15:56 |  |  |  |  | 0

## Orange_585_610
Sample ID | Patient ID | Assay Name | Result Code | Assay Version | Test Guid | Target Result Guid | Test Comment | Replicate Number | Extraction Number | Elution Number | Primer/Probe | Calibration Coefficient | Target Name | Target Setting Specimen Type | Replicate Result | Localized Result | Flags | Ct | Max Peak Height | EPR | End Point Fluorescence | Log Conc. | Log Conc. with Unit Conv. | Conc. | Conc. with Unit Conv. | Log Conc. with Dilution Factor | Log Conc. with Dilution Factor Unit Conv. | Log Conc. Precision | Log Conc. Conv. Precision | Conc. with Dilution Factor | Conc. with Dilution Factor Unit Conv. | Conc. Precision | Conc. Conv. Precision | Log Conc. with Quant Adj. | Log Conc. with Quant Adj. Unit Conv. | Conc. with Quant Adj. | Conc. with Quant Adj. Unit Conv. | Log Intermed. Conc. | Log Intermed. Conc. with Unit Conv. | Intermed. Conc. | Intermed. Conc. with Unit Conv. | Cartridge Barcode | Cartridge Lane | XPCR Module Index | XPCR Module Serial | Start Time | Blank Reading | Dark Reading | Baseline Slope | Baseline Y Intercept | Baseline First Cycle | Baseline Last Cycle | Readings 1 | Readings 2 | Readings 3 | Readings 4 | Readings 5 | Readings 6 | Readings 7 | Readings 8 | Readings 9 | Readings 10 | Readings 11 | Readings 12 | Readings 13 | Readings 14 | Readings 15 | Readings 16 | Readings 17 | Readings 18 | Readings 19 | Readings 20 | Readings 21 | Readings 22 | Readings 23 | Readings 24 | Readings 25 | Readings 26 | Readings 27 | Readings 28 | Readings 29 | Readings 30 | Readings 31 | Readings 32 | Readings 33 | Readings 34 | Readings 35 | Readings 36 | Readings 37 | Readings 38 | Readings 39 | Readings 40 | Readings 41 | Readings 42 | Readings 43 | Readings 44 | Readings 45 | Readings 46 | Readings 47 | Readings 48 | Readings 49 | Readings 50 | Readings 51 | Readings 52 | Readings 53 | Readings 54 | Readings 55 | Readings 56 | Readings 57 | Readings 58 | Readings 59 | Readings 60 | Readings 61 | Readings 62 | Readings 63 | Readings 64 | Readings 65 | Readings 66 | Readings 67 | Readings 68 | Readings 69 | Readings 70 | Readings 71 | Readings 72 | Readings 73 | Readings 74 | Readings 75 | Readings 76 | Readings 77 | Readings 78 | Readings 79 | Readings 80 | Readings 81 | Readings 82 | Readings 83 | Readings 84 | Readings 85 | Readings 86 | Readings 87 | Readings 88 | Readings 89 | Readings 90 | Readings 91 | Readings 92 | Readings 93 | Readings 94 | Readings 95 | Readings 96 | Readings 97 | Readings 98 | Readings 99 | Readings 100 | Readings 101 | Readings 102 | Readings 103 | Readings 104 | Readings 105 | Readings 106 | Readings 107 | Readings 108 | Readings 109 | Readings 110 | Readings 111 | Readings 112 | Readings 113 | Readings 114 | Readings 115 | Readings 116 | Readings 117 | Readings 118 | Readings 119 | Readings 120 | Readings 121 | Readings 122 | Readings 123 | Readings 124 | Readings 125 | Readings 126 | Readings 127 | Readings 128 | Readings 129 | Readings 130 | Readings 131 | Readings 132 | Readings 133 | Readings 134 | Readings 135 | Readings 136 | Readings 137 | Readings 138 | Readings 139 | Readings 140 | Readings 141 | Readings 142 | Readings 143 | Readings 144 | Readings 145 | Readings 146 | Readings 147 | Readings 148 | Readings 149 | Readings 150 | Readings 151 | Readings 152 | Readings 153 | Readings 154 | Readings 155 | Readings 156 | Readings 157 | Readings 158 | Readings 159 | Readings 160 | Readings 161 | Readings 162 | Readings 163 | Readings 164 | Readings 165 | Readings 166 | Readings 167 | Readings 168 | Readings 169 | Readings 170 | Readings 171 | Readings 172 | Readings 173 | Readings 174 | Readings 175 | Readings 176 | Readings 177 | Readings 178 | Readings 179 | Readings 180 | Readings 181 | Readings 182 | Readings 183 | Readings 184 | Readings 185 | Readings 186 | Readings 187 | Readings 188 | Readings 189 | Readings 190 | Readings 191 | Readings 192 | Readings 193 | Readings 194 | Readings 195 | Readings 196 | Readings 197 | Readings 198 | Readings 199 | Readings 200 | Readings 201 | Readings 202 | Readings 203 | Readings 204 | Readings 205 | Readings 206 | Readings 207 | Readings 208 | Readings 209 | Readings 210 | Readings 211 | Readings 212 | Readings 213 | Readings 214 | Readings 215 | Readings 216 | Readings 217 | Readings 218 | Readings 219 | Readings 220 | Readings 221 | Readings 222 | Readings 223 | Readings 224 | Readings 225 | Readings 226 | Readings 227 | Readings 228 | Readings 229 | Readings 230 | Readings 231 | Readings 232 | Readings 233 | Readings 234 | Readings 235 | Readings 236 | Readings 237 | Readings 238 | Readings 239 | Readings 240 | Readings 241 | Readings 242 | Readings 243 | Readings 244 | Readings 245 | Readings 246 | Readings 247 | Readings 248 | Readings 249 | Readings 250 | Readings 251 | Readings 252 | Readings 253 | Readings 254 | Readings 255

## Far_Red_680_715
Sample ID | Patient ID | Assay Name | Result Code | Assay Version | Test Guid | Target Result Guid | Test Comment | Replicate Number | Extraction Number | Elution Number | Primer/Probe | Calibration Coefficient | Target Name | Target Setting Specimen Type | Replicate Result | Localized Result | Flags | Ct | Max Peak Height | EPR | End Point Fluorescence | Log Conc. | Log Conc. with Unit Conv. | Conc. | Conc. with Unit Conv. | Log Conc. with Dilution Factor | Log Conc. with Dilution Factor Unit Conv. | Log Conc. Precision | Log Conc. Conv. Precision | Conc. with Dilution Factor | Conc. with Dilution Factor Unit Conv. | Conc. Precision | Conc. Conv. Precision | Log Conc. with Quant Adj. | Log Conc. with Quant Adj. Unit Conv. | Conc. with Quant Adj. | Conc. with Quant Adj. Unit Conv. | Log Intermed. Conc. | Log Intermed. Conc. with Unit Conv. | Intermed. Conc. | Intermed. Conc. with Unit Conv. | Cartridge Barcode | Cartridge Lane | XPCR Module Index | XPCR Module Serial | Start Time | Blank Reading | Dark Reading | Baseline Slope | Baseline Y Intercept | Baseline First Cycle | Baseline Last Cycle | Readings 1 | Readings 2 | Readings 3 | Readings 4 | Readings 5 | Readings 6 | Readings 7 | Readings 8 | Readings 9 | Readings 10 | Readings 11 | Readings 12 | Readings 13 | Readings 14 | Readings 15 | Readings 16 | Readings 17 | Readings 18 | Readings 19 | Readings 20 | Readings 21 | Readings 22 | Readings 23 | Readings 24 | Readings 25 | Readings 26 | Readings 27 | Readings 28 | Readings 29 | Readings 30 | Readings 31 | Readings 32 | Readings 33 | Readings 34 | Readings 35 | Readings 36 | Readings 37 | Readings 38 | Readings 39 | Readings 40 | Readings 41 | Readings 42 | Readings 43 | Readings 44 | Readings 45 | Readings 46 | Readings 47 | Readings 48 | Readings 49 | Readings 50 | Readings 51 | Readings 52 | Readings 53 | Readings 54 | Readings 55 | Readings 56 | Readings 57 | Readings 58 | Readings 59 | Readings 60 | Readings 61 | Readings 62 | Readings 63 | Readings 64 | Readings 65 | Readings 66 | Readings 67 | Readings 68 | Readings 69 | Readings 70 | Readings 71 | Readings 72 | Readings 73 | Readings 74 | Readings 75 | Readings 76 | Readings 77 | Readings 78 | Readings 79 | Readings 80 | Readings 81 | Readings 82 | Readings 83 | Readings 84 | Readings 85 | Readings 86 | Readings 87 | Readings 88 | Readings 89 | Readings 90 | Readings 91 | Readings 92 | Readings 93 | Readings 94 | Readings 95 | Readings 96 | Readings 97 | Readings 98 | Readings 99 | Readings 100 | Readings 101 | Readings 102 | Readings 103 | Readings 104 | Readings 105 | Readings 106 | Readings 107 | Readings 108 | Readings 109 | Readings 110 | Readings 111 | Readings 112 | Readings 113 | Readings 114 | Readings 115 | Readings 116 | Readings 117 | Readings 118 | Readings 119 | Readings 120 | Readings 121 | Readings 122 | Readings 123 | Readings 124 | Readings 125 | Readings 126 | Readings 127 | Readings 128 | Readings 129 | Readings 130 | Readings 131 | Readings 132 | Readings 133 | Readings 134 | Readings 135 | Readings 136 | Readings 137 | Readings 138 | Readings 139 | Readings 140 | Readings 141 | Readings 142 | Readings 143 | Readings 144 | Readings 145 | Readings 146 | Readings 147 | Readings 148 | Readings 149 | Readings 150 | Readings 151 | Readings 152 | Readings 153 | Readings 154 | Readings 155 | Readings 156 | Readings 157 | Readings 158 | Readings 159 | Readings 160 | Readings 161 | Readings 162 | Readings 163 | Readings 164 | Readings 165 | Readings 166 | Readings 167 | Readings 168 | Readings 169 | Readings 170 | Readings 171 | Readings 172 | Readings 173 | Readings 174 | Readings 175 | Readings 176 | Readings 177 | Readings 178 | Readings 179 | Readings 180 | Readings 181 | Readings 182 | Readings 183 | Readings 184 | Readings 185 | Readings 186 | Readings 187 | Readings 188 | Readings 189 | Readings 190 | Readings 191 | Readings 192 | Readings 193 | Readings 194 | Readings 195 | Readings 196 | Readings 197 | Readings 198 | Readings 199 | Readings 200 | Readings 201 | Readings 202 | Readings 203 | Readings 204 | Readings 205 | Readings 206 | Readings 207 | Readings 208 | Readings 209 | Readings 210 | Readings 211 | Readings 212 | Readings 213 | Readings 214 | Readings 215 | Readings 216 | Readings 217 | Readings 218 | Readings 219 | Readings 220 | Readings 221 | Readings 222 | Readings 223 | Readings 224 | Readings 225 | Readings 226 | Readings 227 | Readings 228 | Readings 229 | Readings 230 | Readings 231 | Readings 232 | Readings 233 | Readings 234 | Readings 235 | Readings 236 | Readings 237 | Readings 238 | Readings 239 | Readings 240 | Readings 241 | Readings 242 | Readings 243 | Readings 244 | Readings 245 | Readings 246 | Readings 247 | Readings 248 | Readings 249 | Readings 250 | Readings 251 | Readings 252 | Readings 253 | Readings 254 | Readings 255

## User Defined Controls
Control Guid | Target Guid | Target Name | Target Result | Control Name | Specimen Types | Assay Name | Last Success | Is Active | Require Lot Frequency | Require Run Frequency | Run Frequency Days | Expected Concentration | Concentration Lower Limit | Concentration Upper Limit

## Standard Curves
Standard Curve GUID | Lot GUID | Lot Number | Standard Curve Name | Assay Name | Software Version Created On | Assay Version Created Against | Instrument Serial Number Created On | Valid Date Time | Entry Type | Created By User | Master Mix Lot Number | Primer Probe Lot Number | Target Name | Slope | Y Intercept | R2 | Target Name | Slope | Y Intercept | R2 | Target Name | Slope | Y Intercept | R2

## ADF BKV URINE 10,0.0.40
Key | Value
Id | f394bea5-0509-ed11-85f4-5cf370a20867
Assay Type | LDT
Result Type | Quantitative
Chemistry Type | DNA
Master Mix Type | 019
Assay Display Name | VIR urine BKV 10
Assay Name | BKV URINE 10
Assay Version | 0.0.40
Release Date | 2022-07-21
Classification | LDT
Barcode
Number of Extractions | 1
Number of PCR Mixes | 1
RP Version | V2
Result Processing Algorithm | Ct and EP
Ct Calling Algorithm | Second Derivative
Starting Fluorescence Start Cycle | 5
Starting Fluorescence End Cycle | 8
End Point Fluorescence Start Cycle | 45
End Point Fluorescence End Cycle | 45
Open-Life (hours) | 8
Fill Check Reporter | Green (470/510)
Fill Check Threshold | 500
Enable Quantitation Precision | FALSE
Green Reporter Crosstalk Into Yellow Reporter (%) | 0
Yellow Reporter Crosstalk Into Orange Reporter (%) | 0
Orange Reporter Crosstalk Into Red Reporter (%) | 0
Red Reporter Crosstalk Into Far Red Reporter (%) | 0
Specimen Urine - Specimen Type | Urine
Specimen Urine - Specimen Stability (hrs.) | 24
Specimen Urine - Specimen Aspirate Volume (µL) | 400
Specimen Urine - Specimen Mix Volume (µL) | 400
Specimen Urine - Specimen Mix Count | 3
Specimen Urine - Buffer | Lysis Buffer 2
Specimen Urine - Use Override Buffer | FALSE
Specimen Urine - Enable Calibration Factor | FALSE
Specimen Urine - Treat first rerun as repeat | FALSE
Specimen Urine - Use different rerun parameters | FALSE
Specimen Urine - Extraction Plate Heater Temp (°C) | 38
Specimen Urine - Lysis Duration (sec) | 600
XPCR Stage Hold - Stage Type | Hold
XPCR Stage Hold - Duration (seconds) | 60
XPCR Stage Hold - Temperature (°C) | 95
XPCR Stage Cycle - Stage Type | Cycle
XPCR Stage Cycle - Cycle Count | 45
XPCR Stage Cycle, Step Denature - Duration (sec) | 6
XPCR Stage Cycle, Step Denature - Temperature (°C) | 95
XPCR Stage Cycle, Step Denature - Detect? | FALSE
XPCR Stage Cycle, Step Anneal - Duration (sec) | 19
XPCR Stage Cycle, Step Anneal - Temperature (°C) | 60
XPCR Stage Cycle, Step Anneal - Detect? | TRUE
XPCR Primer Probe PP1, Target SPC1 - Name | SPC1
XPCR Primer Probe PP1, Target SPC1 - Reporter | Yellow (530/555)
XPCR Primer Probe PP1, Target SPC1 - Sample Process Control? | TRUE
XPCR Target SPC1 Setting, SpecimenType Urine - Minimum Peak Height | 50
XPCR Target SPC1 Setting, SpecimenType Urine - Peak Minimum Cycle | 28
XPCR Target SPC1 Setting, SpecimenType Urine - Peak Maximum Cycle | 35
XPCR Target SPC1 Setting, SpecimenType Urine - Minimum End Point Fluorescence | 1000
XPCR Target SPC1 Setting, SpecimenType Urine - Max Outliers | 10
XPCR Target SPC1 Setting, SpecimenType Urine - Enable Derivative Data Filtering | TRUE
XPCR Target SPC1 Setting, SpecimenType Urine - Derivative Data Filtering Trailing Window Size | 3
XPCR Target SPC1 Setting, SpecimenType Urine - Derivative Data Filtering Leading Window Size | 3
XPCR Target SPC1 Setting, SpecimenType Urine - Derivative Data Filtering Z-Score Threshold | 6
XPCR Target SPC1 Setting, SpecimenType Urine - Derivative Data Filtering Step Height Threshold | 200
XPCR Target SPC1 Setting, SpecimenType Urine - Baseline Start Cycle | 5
XPCR Target SPC1 Setting, SpecimenType Urine - Fixed Baseline Start | 2
XPCR Target SPC1 Setting, SpecimenType Urine - Fixed Baseline End | 4
XPCR Target SPC1 Setting, SpecimenType Urine - Baseline Lookback Cycles | 8
XPCR Target SPC1 Setting, SpecimenType Urine - Baseline Minimum Window Cycles | 5
XPCR Target SPC1 Setting, SpecimenType Urine - Baseline Slope Lower Threshold | -250
XPCR Target SPC1 Setting, SpecimenType Urine - Baseline Slope Upper Threshold | 250
XPCR Target SPC1 Setting, SpecimenType Urine - Maximum Starting Fluorescence | 50000
XPCR Target SPC1 Setting, SpecimenType Urine - Ct Below Fixed Baseline Start Result | Indeterminate
XPCR Target SPC1 Setting, SpecimenType Urine - EPR Threshold | 1.5
XPCR Target SPC1 Setting, SpecimenType Urine - Overall EPR Threshold | 1.03
XPCR Target SPC1 Setting, SpecimenType Urine - Overall EPR Threshold Check Failed Result | Indeterminate
XPCR Target SPC1 Setting, SpecimenType Urine - EPR Check Ct Threshold | 35
XPCR Target SPC1 Setting, SpecimenType Urine - EPR Check Failed Result | Indeterminate
XPCR Primer Probe PP1, Target SPC1, Localization 1 - Assigned Result | Target Amplified
XPCR Primer Probe PP1, Target SPC1, Localization 1 - Locale | en
XPCR Primer Probe PP1, Target SPC1, Localization 1 - Result String | Valid
XPCR Primer Probe PP1, Target SPC1, Localization 2 - Assigned Result | Target Not Amplified
XPCR Primer Probe PP1, Target SPC1, Localization 2 - Locale | en
XPCR Primer Probe PP1, Target SPC1, Localization 2 - Result String | Not Detected
XPCR Primer Probe PP1, Target SPC1, Localization 3 - Assigned Result | Target Indeterminate
XPCR Primer Probe PP1, Target SPC1, Localization 3 - Locale | en
XPCR Primer Probe PP1, Target SPC1, Localization 3 - Result String | Indeterminate
XPCR Primer Probe PP1, Target BKV - Name | BKV
XPCR Primer Probe PP1, Target BKV - Reporter | Green (470/510)
XPCR Primer Probe PP1, Target BKV - Sample Process Control? | FALSE
XPCR Target BKV Setting, SpecimenType Urine - Minimum Peak Height | 40
XPCR Target BKV Setting, SpecimenType Urine - Peak Minimum Cycle | 10
XPCR Target BKV Setting, SpecimenType Urine - Peak Maximum Cycle | 45
XPCR Target BKV Setting, SpecimenType Urine - Minimum End Point Fluorescence | 300
XPCR Target BKV Setting, SpecimenType Urine - Max Outliers | 10
XPCR Target BKV Setting, SpecimenType Urine - Enable Derivative Data Filtering | TRUE
XPCR Target BKV Setting, SpecimenType Urine - Derivative Data Filtering Trailing Window Size | 3
XPCR Target BKV Setting, SpecimenType Urine - Derivative Data Filtering Leading Window Size | 3
XPCR Target BKV Setting, SpecimenType Urine - Derivative Data Filtering Z-Score Threshold | 6
XPCR Target BKV Setting, SpecimenType Urine - Derivative Data Filtering Step Height Threshold | 200
XPCR Target BKV Setting, SpecimenType Urine - Baseline Start Cycle | 3
XPCR Target BKV Setting, SpecimenType Urine - Fixed Baseline Start | 2
XPCR Target BKV Setting, SpecimenType Urine - Fixed Baseline End | 4
XPCR Target BKV Setting, SpecimenType Urine - Baseline Lookback Cycles | 4
XPCR Target BKV Setting, SpecimenType Urine - Baseline Minimum Window Cycles | 3
XPCR Target BKV Setting, SpecimenType Urine - Baseline Slope Lower Threshold | -250
XPCR Target BKV Setting, SpecimenType Urine - Baseline Slope Upper Threshold | 250
XPCR Target BKV Setting, SpecimenType Urine - Maximum Starting Fluorescence | 50000
XPCR Target BKV Setting, SpecimenType Urine - Ct Below Fixed Baseline Start Result | Indeterminate
XPCR Target BKV Setting, SpecimenType Urine - Inhibition Check | 45
XPCR Target BKV Setting, SpecimenType Urine - EPR Threshold | 2
XPCR Target BKV Setting, SpecimenType Urine - Overall EPR Threshold | 1.03
XPCR Target BKV Setting, SpecimenType Urine - Overall EPR Threshold Check Failed Result | Indeterminate
XPCR Target BKV Setting, SpecimenType Urine - EPR Check Ct Threshold | 8
XPCR Target BKV Setting, SpecimenType Urine - EPR Check Failed Result | Indeterminate
XPCR Target BKV Setting, SpecimenType Urine - Lower Limit of Quantification (log Conc.) | 1
XPCR Target BKV Setting, SpecimenType Urine - Upper Limit of Quantification (log Conc.) | 10
XPCR Target BKV Setting, SpecimenType Urine - Concentration Display Unit | c/mL
XPCR Target BKV Setting, SpecimenType Urine - Enable Conversion Factor | TRUE
XPCR Target BKV Setting, SpecimenType Urine - Concentration Conversion Factor | 0.6596
XPCR Target BKV Setting, SpecimenType Urine - Concentration Conversion Factor Display Unit | IU/mL
XPCR Target BKV Setting, SpecimenType Urine - Quant Check Ct Threshold | 40
XPCR Target BKV Setting, SpecimenType Urine - Enable Quant Adjustment | FALSE
XPCR Primer Probe PP1, Target BKV, Localization 1 - Assigned Result | Target Amplified
XPCR Primer Probe PP1, Target BKV, Localization 1 - Locale | en
XPCR Primer Probe PP1, Target BKV, Localization 1 - Result String | Detected
XPCR Primer Probe PP1, Target BKV, Localization 2 - Assigned Result | Target Not Amplified
XPCR Primer Probe PP1, Target BKV, Localization 2 - Locale | en
XPCR Primer Probe PP1, Target BKV, Localization 2 - Result String | Not Detected
XPCR Primer Probe PP1, Target BKV, Localization 3 - Assigned Result | Target Indeterminate
XPCR Primer Probe PP1, Target BKV, Localization 3 - Locale | en
XPCR Primer Probe PP1, Target BKV, Localization 3 - Result String | Indeterminate
XPCR Primer Probe PP1, Target BKV, Localization 4 - Assigned Result | Target Unresolved
XPCR Primer Probe PP1, Target BKV, Localization 4 - Locale | en
XPCR Primer Probe PP1, Target BKV, Localization 4 - Result String | Unresolved
Result Code L999 - Result Name | L999
Result Code L999 - Result Type | Quantitative
Result Code L999 - Target Name BKV - Is Selected | TRUE
Result Code L999 - Target Name BKV - Reporter | Green (470/510)
Assay Standard Curve - Name | StdCurveUrine
Assay Standard Curve Target BKV, Reporter Green_470_510 - Slope | -3.2886
Assay Standard Curve Target BKV, Reporter Green_470_510 - YIntercept | 38.7933
Assay Standard Curve Target BKV, Reporter Green_470_510 - R2 | 0.9958